## Complaint Analysis - 2025-06-18
Complaint No | Created Date | Close Date | Customer Number | Branch | Division | Complaint Status | Comp Type | Equipment Type | Equipment Name | Problem Desc | Work Detail | Resolution Comment | Warranty | Trip Type | Market Vertical | Critical | Critical Type | Manufacturing Div Name
AF4X24163252 | 2025-02-04 | 2025-02-28 | Cust 456 | CSD HO | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | Xenon Lamp Power supply for Multiparameter(AWA-XLAMPPS) | Xenon Power Supply in newly supplied CX2000 failed | Replacment Material Shipped Vide AWB 772198417740. | Replacment Material Shipped Vide AWB 772198417740. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD23035923 | 2023-06-06 | 2023-06-07 | Cust 457 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | Transmitter not powering up | Checked transmitters and observed moisture. Suggested to check mA and revert. In case of mA not proper, need to send the transmitters to Pune for repair. Refer attached mail. | Checked transmitters and observed moisture. Suggested to check mA and revert. In case of mA not proper, need to send the transmitters to Pune for repair. Refer attached mail. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFAD23035923 | 2023-06-06 | 2023-06-07 | Cust 457 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | Transmitter not powering up | Checked transmitters and observed moisture. Suggested to check mA and revert. In case of mA not proper, need to send the transmitters to Pune for repair. Refer attached mail. | Checked transmitters and observed moisture. Suggested to check mA and revert. In case of mA not proper, need to send the transmitters to Pune for repair. Refer attached mail. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFAD23035923 | 2023-06-06 | 2023-06-07 | Cust 457 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | Transmitter not powering up | Checked transmitters and observed moisture. Suggested to check mA and revert. In case of mA not proper, need to send the transmitters to Pune for repair. Refer attached mail. | Checked transmitters and observed moisture. Suggested to check mA and revert. In case of mA not proper, need to send the transmitters to Pune for repair. Refer attached mail. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFAD23035923 | 2023-06-06 | 2023-06-07 | Cust 457 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Transmitter not powering up | Checked transmitters and observed moisture. Suggested to check mA and revert. In case of mA not proper, need to send the transmitters to Pune for repair. Refer attached mail. | Checked transmitters and observed moisture. Suggested to check mA and revert. In case of mA not proper, need to send the transmitters to Pune for repair. Refer attached mail. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFAD23037264 | 2023-06-09 | 2023-06-13 | Cust 458 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Aqua4Trans | Aqua4Trans(1) | 4-20mA Communication | guided customer over the phone for 4-20 mA wiring. Refer attached mail. | guided customer over the phone for 4-20 mA wiring. Refer attached mail. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFAD23038365 | 2023-06-12 | 2023-06-15 | Cust 459 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | RSPCB data uploaded issue. | Data is live on RSPCB server, hence we are closing this complaint. | Data is live on RSPCB server, hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD23039836 | 2023-06-14 | 2023-06-16 | Cust 460 | AHMEDABAD | Steam System Division | CLOSED | QUALITY PROBLEM | TRAP - SOFT | FORBES MARSHALL SOFT53 SINGLE ORIFICE FLOAT TRAP - SIZE 40NB(040SOFT53) | Product found with pin holes in the flange | As a interim solution we have called back the product to Chakan - NCR has been raised on the concern | As a interim solution we have called back the product to Chakan - NCR has been raised on the concern | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFAD23039917 | 2023-06-14 | 2023-06-16 | Cust 461 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4 trans-A4TW033000(1) | Commissioning | Guided over phone for commissioning, Found working ok. Refer attached mail. | Guided over phone for commissioning, Found working ok. Refer attached mail. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD23040838 | 2023-06-16 | 2023-06-21 | Cust 462 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software installation and CPCB + RSPCB connectivity needs to be done. | As of now we are closing this complaint, we will reopen it after getting system access from customers end. | As of now we are closing this complaint, we will reopen it after getting system access from customers end. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD23044439 | 2023-06-23 | 2023-06-25 | Cust 463 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Aqua4Trans | Aqua4Trans(1) | CCDS IN AQUA4TRANS CONTROLLER NOT WORKING | CHANGE NEW ONE AQUA4TRANS CONTROLLER OF ANOTHER CCDS. DONE PROGRAMING AND SET 50 PPM. | CHANGE NEW ONE AQUA4TRANS CONTROLLER OF ANOTHER CCDS. DONE PROGRAMING AND SET 50 PPM. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFAD230447710 | 2023-06-23 | 2023-06-24 | Cust 464 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera streming issue on CPCB and RSPCB server. | now the camera streaming is live on both servers | now the camera streaming is live on both servers | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD230464311 | 2023-06-26 | 2023-06-28 | Cust 465 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | CPCB data uploaded issue. | We took access of the system and found a network issue at the customer end and after that, they resolved the same and now the data is live on the server. | Data is live hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD230469912 | 2023-06-27 | 2023-06-28 | Cust 458 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Chlorine sensor | CHLORINE SENSOR(1) | Commissioning of Aqua4Trans with chlorine and pH Sensor . The instrument was supplied by OEM Pure Oxidane to M/s Aquatics | Site visit made for Commissioning, Refer to the attached Mail. | Site visit made for commissioning. Refer attached mail. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD230469912 | 2023-06-27 | 2023-06-28 | Cust 458 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4Trans(1) | Commissioning of Aqua4Trans with Turbidity sensor. The instrument was supplied by OEM Pure Oxidane to M/s Aquatics | Site visit made for Commissioning, Refer to the attached Mail. | Site visit made for commissioning. Refer attached mail. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD230470113 | 2023-06-27 | 2023-06-29 | Cust 462 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation needs to be done. | We took the access to the system and successfully did the software reinstallation + local + CPCB + RSPCB configuration. now data is live on all the server | We took the access to the system and successfully did the software reinstallation + local + CPCB + RSPCB configuration. now data is live on all the server | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD230473514 | 2023-06-28 | 2023-06-28 | Cust 466 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | RSPCB configuration needs to be done. | we took the access to the system and successfully did the RSPCB configuration and now data is live on the RSPCB server. | we took the access to the system and successfully did the RSPCB configuration and now data is live on the RSPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD230501315 | 2023-07-03 | 2023-07-04 | Cust 464 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera streaming issue. | The issue has been resolved and camera streaming is live. | The issue has been resolved and camera streaming is live. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD230509716 | 2023-07-04 | 2023-08-21 | Cust 467 | AHMEDABAD | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | Aqua4Trans | Aqua4 trans-A4TW033000(1) | Display Blank | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210055856 LR Dt - 03-AUG-2023 | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210055856 LR Dt - 03-AUG-2023 | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD230514517 | 2023-07-05 | 2023-07-06 | Cust 458 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4Trans(1) | Commissioning of Aqua4trans with TDS measurement at Dahej Site | Guided the site person for termination and found working ok. Refer attached Mail. | Guided the site person for termination and found working ok. Refer attached Mail. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD230521018 | 2023-07-06 | 2023-07-10 | Cust 468 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Unit not powering up | Checked the meter over the phone and suggested sending the same to Pun e for further checking. Refer attached mail. | Checked the meter over the phone and suggested sending the same to Pun e for further checking. Refer attached mail. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFAD230525019 | 2023-07-07 | 2023-07-10 | Cust 469 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation needs to be done. | data is live on all the servers, hence we are closing this complaint. | data is live on all the servers, hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD230540820 | 2023-07-10 | 2023-07-11 | Cust 458 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Transmitter Aquacon | Aquacon 2000(AQCN 2000 p) | Commissioning and calibration of TDS Transmitter | Guided over the phone for wiring and calibration, Commissioned the analyser successfully, Refer to attached mail. | Guided over phone and commissioned the analysers, Refer attached mail. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD230540820 | 2023-07-10 | 2023-07-11 | Cust 458 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4Trans(1) | Commissioning and calibration of Conductivity Transmitter | Guided over the phone for calibration and commissioned the analysers, Refer to attached mail. | Guided over phone and commissioned the analysers, Refer attached mail. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD230543321 | 2023-07-11 | 2023-07-14 | Cust 470 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB + RSPCB data upload issue. | As of now we are closing this complaint, we will reopen it after the hardware issue resolves from customers end. | As of now we are closing this complaint, we will reopen it after the hardware issue resolves from customers end. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD230561922 | 2023-07-14 | 2023-07-17 | Cust 471 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | RSPCB data upload issue. | We have update the output.jar file and now data availability for all the devices is increase above 85%. | We have update the output.jar file and now data availability for all the devices is increase above 85%. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD230573623 | 2023-07-17 | 2023-07-18 | Cust 472 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Turbidity Input card of Aqua4Trans(1) | Turbidity meter calibration | Guided over phone for calibration, Meter working and under observation, Refer attached mail. | Guided over phone for calibration, Meter working and under observation, Refer attached mail. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD230583124 | 2023-07-18 | 2023-07-20 | Cust 470 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue. | We took the access to the system and found that the RS485 converter is faulty, hence here is issue from hardware end, as of now we are closing this complaint. | We took the access to the system and found that the RS485 converter is faulty, hence here is issue from hardware end, as of now we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD230583325 | 2023-07-18 | 2023-07-20 | Cust 470 | AHMEDABAD | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 1st AMC | AMC visit done by Nimish at Bithuja Site for Multiparamater. Refer attached MOM | AMC visit made by Nimish, Refer attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFAD230588226 | 2023-07-19 | 2023-07-21 | Cust 473 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | RSPCB data uploaded issue. | As of now, we are closing this complaint and will reopen this case once you are ready with the remort access of the system | As of now, we are closing this complaint and will reopen this case once you are ready with the remort access of the system | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD230597527 | 2023-07-21 | 2023-07-24 | Cust 458 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4Trans(1) | Commissioning of chlorine Analyser | Site visit made by Chirag Shah for Commissioning of Chlorine Analyser, Refer attached mail for the activities carried out, | Site visit made for commissioning. Refer attached mail. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD230608728 | 2023-07-24 | 2023-07-27 | Cust 471 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | RSPCB data availability issue. | We took the access to the system and update the RTDAS and now data availability is above 90% for both devices. | We took the access to the system and update the RTDAS and now data availability is above 90% for both devices. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD230608829 | 2023-07-24 | 2023-07-25 | Cust 470 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue. | Data is live on all the servers, hence we are closing this complaint. | Data is live on all the servers, hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD230623430 | 2023-07-26 | 2023-07-27 | Cust 464 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | RSPCB data uploaded issue. | We took access of the system and made some changes so now the data is live on the server. | Data is live hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD230628531 | 2023-07-27 | 2023-07-28 | Cust 472 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Aqua4Trans | Turbidity Input card of Aqua4Trans(1) | Turbidity Analyser Troubleshooting | Site visit made for the troubleshooting of Turbidity analyser, Observed air ingress in flow through the chamber during backwash, installed air eliminator and found working ok. Refer to the attached MOM. | Site visit made for Troubleshooting. Refer attached MOM | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD230628632 | 2023-07-27 | 2023-07-28 | Cust 474 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | PC hang issue, need to change the PC and reinstall the software. | We took the access to the system and successfully did the software reinstallation + Local +CPCB + RSPCB configuration, and now data is live on all the servers. | We took the access to the system and successfully did the software reinstallation + Local +CPCB + RSPCB configuration, and now data is live on all the servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD230632333 | 2023-07-28 | 2023-09-27 | Cust 471 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | RSPCB data upload issue. | Data is not communication continuously on mod scan hence we are not getting continuous data but we have updated the jar file for delayed data and data is live on the RSPCB server. | Data is not communication continuously on mod scan hence we are not getting continuous data but we have updated the jar file for delayed data and data is live on the RSPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD230635334 | 2023-07-28 | 2023-07-31 | Cust 475 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | DO Transmitters | optex series(optex series) | DO reading issue | Guided over the phone and resolved the issue, Refer attached Mail. | Guided over the phone and resolved the issue, Refer attached Mail. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFAD230651635 | 2023-07-30 | 2023-07-31 | Cust 476 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Display of Smartpro(1) | Commissioning of the new analyzers | Did the commissioning remotely. | Did the commissioning remotely. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFAD230661336 | 2023-07-31 | 2023-08-01 | Cust 477 | AHMEDABAD | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 3rd AMC Visit | 3rd AMC visit made by Nimish, Refer attached MOM. | AMC Visit done. Refer attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFAD230667137 | 2023-08-01 | 2023-08-04 | Cust 464 | AHMEDABAD | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Emergency Visit | Site visit made for emergency visit as per AMC. Refer attached MOM. | Site visit made for emergency visit of AMC. Refer attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFAD230672938 | 2023-08-02 | 2023-08-04 | Cust 478 | AHMEDABAD | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 4th AMC Visit | Site visit made for preventive Maintenance under AMC. Refer attached MOM for the activities carried out. | AMC Visit done. Refer attached MOM. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFAD230676339 | 2023-08-03 | 2023-08-04 | Cust 464 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | RSPCB data uploaded issue, | Data is live on RSPCB server, hence we are closing this complaint. | Data is live on RSPCB server, hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD230685640 | 2023-08-05 | 2023-08-07 | Cust 479 | AHMEDABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | PH Sensor wasn't sensing changes of Value of PH of water and checked other PH Meter's. | Done Calibration of PH Meter and One PH Sensor wasn't sensing and Checked 2 other PH meters. | Done Calibration of PH Meter and One PH Sensor wasn't sensing and Checked 2 other PH meters. ( GRAINSPAN ) | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFAD230685640 | 2023-08-05 | 2023-08-07 | Cust 479 | AHMEDABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | PH Sensor wasn't sensing changes of Value of PH of water and checked other PH Meter's. | Done Calibration of PH Meter and One PH Sensor wasn't sensing and Checked 2 other PH meters. | Done Calibration of PH Meter and One PH Sensor wasn't sensing and Checked 2 other PH meters. ( GRAINSPAN ) | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFAD230685640 | 2023-08-05 | 2023-08-07 | Cust 479 | AHMEDABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | PH Sensor wasn't sensing changes of Value of PH of water and checked other PH Meter's. | Done Calibration of PH Meter and One PH Sensor wasn't sensing and Checked 2 other PH meters. | Done Calibration of PH Meter and One PH Sensor wasn't sensing and Checked 2 other PH meters. ( GRAINSPAN ) | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFAD230705841 | 2023-08-08 | 2023-08-10 | Cust 461 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Electrode | ORP Electrode- Polilyte RX 120(1) | ORP Sensor contaminated | Suspected issue of ORP sensor due to installation of sensor in line, Refer attached Mail. | Suspected issue of ORP sensor due to installation of sensor in line, Refer attached Mail. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD230714542 | 2023-08-10 | 2023-08-14 | Cust 466 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | RSPCB data upload issue need to update enviroconnect software patch | We took the access to your system and successfully updated the enviroconnect software patch till 55 and output.jar file, currently data is live on RSPCB server. | We took the access to your system and successfully updated the enviroconnect software patch till 55 and output.jar file, currently data is live on RSPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD230724743 | 2023-08-14 | 2023-08-17 | Cust 479 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | DO Sensor not sensing properly. Offer submitted. | DO Sensor not working submitted offer. | DO Sensor not working submitted offer. | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFAD230727144 | 2023-08-14 | 2023-08-16 | Cust 480 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | Issue in Conductivity reading | Guided over the phone and checked the transmitter, concluded that transmitter needs to send to Pune for repair. Refer attached mail. | Guided over the phone and resolved the issue in one transmitter, another transmitter needs to send to Pune for repair. Refer attached mail. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFAD230727144 | 2023-08-14 | 2023-08-16 | Cust 480 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | Issue in Conductivity reading | Guided over the phone and resolved the issue by cleaning of sensor and calibration with ref. solution. Refer attached mail. | Guided over the phone and resolved the issue in one transmitter, another transmitter needs to send to Pune for repair. Refer attached mail. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFAD230734245 | 2023-08-15 | 2023-08-16 | Cust 480 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | Commissioning and troubleshooting | Guided over the phone for proper wiring, refer attached Mail. | Guided over the phone for proper wiring, refer attached Mail. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD230736546 | 2023-08-16 | 2023-08-17 | Cust 461 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Reading at higher side | Guided over the phone but observed an issue with the pH sensor. Need to change the sensor with a new one. | Guided over the phone but observed an issue with the pH sensor. Need to change the sensor with a new one. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD230744147 | 2023-08-17 | 2023-08-22 | Cust 481 | AHMEDABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | mA Output card for Multiparameter(AWA-4-20mA) | changed mA card for output and done calibration of multiparameter analyser. | changed mA card for output and done calibration of multiparameter analyser. | changed mA card for output and done calibration of multiparameter analyser. | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFAD230771348 | 2023-08-22 | 2023-08-23 | Cust 470 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | RSPCB camera streaming issue. | We have create a RTSP link and share to customer to proceed further to RSPCB camera configuration. so now we are closing this complaint | We have create a RTSP link and share to customer to proceed further to RSPCB camera configuration. so now we are closing this complaint | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD230774649 | 2023-08-23 | 2023-08-25 | Cust 467 | AHMEDABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Aqua4Trans | Aqua4Trans(1) | Installation of supplied FOC transmitter | Site visit made by Riteshbhai for installation of FOC supplied transmitter, refer attached MOM. | Site visit made by Riteshbhai for installation of FOC supplied transmitter, refer attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD230854250 | 2023-09-07 | 2023-09-08 | Cust 482 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | CPCB and RSPCB data uploaded issue. | Software crashed issue, hence we configured the software + Local + CPCB + RSPCB. now data is live on all the servers. | Software crashed issue, hence we configured the software + Local + CPCB + RSPCB. now data is live on all the servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD230857151 | 2023-09-08 | 2023-09-11 | Cust 473 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB camera streaming issue. | We took the access to customers system and did the camera configuration, now the camera is live on CPCB server. | We took the access to customers system and did the camera configuration, now the camera is live on CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD230869452 | 2023-09-11 | 2023-09-14 | Cust 473 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on RSPCB | Successfully did the software reinstallation local + CPCB + RSPCb configuration and now data is live on all the servers. | Successfully did the software reinstallation local + CPCB + RSPCb configuration and now data is live on all the servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD230875753 | 2023-09-12 | 2023-09-13 | Cust 474 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Need to create RTSP link for RSPCB configuration. | We create a RTSP link for RSPCB camera configuration and sent to customer, as of now we are closing this complaint. | We create a RTSP link for RSPCB camera configuration and sent to customer, as of now we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD230885854 | 2023-09-14 | 2023-09-15 | Cust 462 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation needs to be done. | We took the access to the system and successfully did the software reinstallation + local + CPCB + RSPCB configuration, and now data is live on all the servers. | We took the access to the system and successfully did the software reinstallation + local + CPCB + RSPCB configuration, and now data is live on all the servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD230886955 | 2023-09-14 | 2023-09-15 | Cust 483 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Commissioning and Troubleshooting of Aquamon SMARTPro 8966 with Sensor FM1106 | Site visit made for commissioning and troubleshooting, Observed issue with sensor. Suggested to send the same to Pune. The transmitter working ok. Refer attached MOM. | Site visit made for commissioning. Refer attached MOM | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD230886955 | 2023-09-14 | 2023-09-15 | Cust 483 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Commissioning and Troubleshooting of Aquamon SMARTPro 8966 with Sensor FM1106 | Site visit made for commissioning and troubleshooting, Observed issue with sensor. Suggested to send the same to Pune. The transmitter working ok. Refer attached MOM. | Site visit made for commissioning. Refer attached MOM | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD230886955 | 2023-09-14 | 2023-09-15 | Cust 483 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Commissioning and Troubleshooting of Aquamon SMARTPro 8966 with Sensor FM1106 | Site visit made for commissioning and troubleshooting, Observed issue with sensor. Suggested to send the same to Pune. The transmitter working ok. Refer attached MOM. | Site visit made for commissioning. Refer attached MOM | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD230886955 | 2023-09-14 | 2023-09-15 | Cust 483 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | Commissioning and Troubleshooting of Aquacon SMARTPro 8967 with Sensor CS42 | Site visit made for commissioning and troubleshooting, Observed issue with sensor. Suggested to send the same to Pune. The transmitter working ok. Refer attached MOM. | Site visit made for commissioning. Refer attached MOM | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD230886955 | 2023-09-14 | 2023-09-15 | Cust 483 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | Commissioning and Troubleshooting of Aquacon SMARTPro 8967 with Sensor CS42 | Site visit made for commissioning and troubleshooting, Calibrated the sensor with standards and found working ok. Refer attached MOM. | Site visit made for commissioning. Refer attached MOM | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD230886955 | 2023-09-14 | 2023-09-15 | Cust 483 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Commissioning and Troubleshooting of Aquamon SMARTPro 8966 with Sensor FM1106 | Site visit made for commissioning and troubleshooting, Calibrated with standards. The Analyser working ok. Refer attached MOM. | Site visit made for commissioning. Refer attached MOM | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD230886955 | 2023-09-14 | 2023-09-15 | Cust 483 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Commissioning and Troubleshooting of Aquamon SMARTPro 8966 with Sensor FM1106 | Site visit made for commissioning and troubleshooting, Calibrated with standards. The Analyser working ok. Refer attached MOM. | Site visit made for commissioning. Refer attached MOM | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD230886955 | 2023-09-14 | 2023-09-15 | Cust 483 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Commissioning and Troubleshooting of Aquamon SMARTPro 8966 with Sensor FM1106 | Site visit made for commissioning and troubleshooting, Calibrated with standards. The Analyser working ok. Refer attached MOM. | Site visit made for commissioning. Refer attached MOM | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD230886955 | 2023-09-14 | 2023-09-15 | Cust 483 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Commissioning and Troubleshooting of Aquamon SMARTPro 8966 with Sensor FM1106 | Site visit made for commissioning and troubleshooting, Calibrated with standards. The Analyser working ok. Refer attached MOM. | Site visit made for commissioning. Refer attached MOM | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD230886955 | 2023-09-14 | 2023-09-15 | Cust 483 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Commissioning and Troubleshooting of Aquamon SMARTPro 8966 with Sensor FM1106 | Site visit made for commissioning and troubleshooting, Calibrated with standards. The Analyser working ok. Refer attached MOM. | Site visit made for commissioning. Refer attached MOM | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD230889056 | 2023-09-15 | 2023-09-15 | Cust 484 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data mismatch issue | Took access of the system and found data not coming on modscan.Hence as of now we are cosing the complaint will repopen it after the hardware issue resolved from customer end. | Took access of the system and found data not coming on modscan.Hence as of now we are cosing the complaint will repopen it after the hardware issue resolved from customer end. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD230889157 | 2023-09-15 | 2023-09-15 | Cust 484 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data mismatch issue on cpcb server | Took access of the system and found data not coming on modscan.Hence as of now we are cosing the complaint will reopen it after the hardware issue resolved from customer end. | Took access of the system and found data not coming on modscan.Hence as of now we are cosing the complaint will repopen it after the hardware issue resolved from customer end. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD230889458 | 2023-09-15 | 2023-10-16 | Cust 483 | AHMEDABAD | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | Ph Electrode | FM1106(1) | No response of pH sensor. Total two sensor supplied as spare. Out to two one pH sensor having issue of no response when dipped in Various standard solutions. | A customer sent Three sensors only. for the same replacement was dispatched as MARK EXPRESS PRIVATE LIMITED, LR/Docket No. : 210070875. so closing the complaint for now. | A customer sent Three sensors only. for the same replacement was dispatched as MARK EXPRESS PRIVATE LIMITED, LR/Docket No. : 210070875. so closing the complaint for now. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD230889458 | 2023-09-15 | 2023-10-16 | Cust 483 | AHMEDABAD | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | Ph Electrode | FM1106(1) | Out of 8 supplied pH sensors, 04 Nos. pH sensor not responding with different buffer solutions. | A customer sent Three sensors only. for the same replacement was dispatched as MARK EXPRESS PRIVATE LIMITED, LR/Docket No. : 210070875. so closing the complaint for now. | A customer sent Three sensors only. for the same replacement was dispatched as MARK EXPRESS PRIVATE LIMITED, LR/Docket No. : 210070875. so closing the complaint for now. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD230889458 | 2023-09-15 | 2023-10-16 | Cust 483 | AHMEDABAD | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | Ph Electrode | FM1106(1) | Out of 8 supplied pH sensors, 04 Nos. pH sensor not responding with different buffer solutions. | A customer sent Three sensors only. for the same replacement was dispatched as MARK EXPRESS PRIVATE LIMITED, LR/Docket No. : 210070875. so closing the complaint for now. | A customer sent Three sensors only. for the same replacement was dispatched as MARK EXPRESS PRIVATE LIMITED, LR/Docket No. : 210070875. so closing the complaint for now. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD230889458 | 2023-09-15 | 2023-10-16 | Cust 483 | AHMEDABAD | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | Ph Electrode | FM1106(1) | Out of 8 supplied pH sensors, 04 Nos. pH sensor not responding with different buffer solutions. | A customer sent Three sensors only. for the same replacement was dispatched as MARK EXPRESS PRIVATE LIMITED, LR/Docket No. : 210070875. so closing the complaint for now. | A customer sent Three sensors only. for the same replacement was dispatched as MARK EXPRESS PRIVATE LIMITED, LR/Docket No. : 210070875. so closing the complaint for now. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD230889458 | 2023-09-15 | 2023-10-16 | Cust 483 | AHMEDABAD | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | Ph Electrode | FM1106(1) | Out of 8 supplied pH sensors, 04 Nos. pH sensor not responding with different buffer solutions. | A customer sent Three sensors only. for the same replacement was dispatched as MARK EXPRESS PRIVATE LIMITED, LR/Docket No. : 210070875. so closing the complaint for now. | A customer sent Three sensors only. for the same replacement was dispatched as MARK EXPRESS PRIVATE LIMITED, LR/Docket No. : 210070875. so closing the complaint for now. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD230891059 | 2023-09-15 | 2023-09-19 | Cust 477 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Troubleshooting | Urvesh Duided over phone and resolved the issue by Hardreset the Analyser. Refer to attached Mail. | Urvesh Duided over phone and resolved the issue by Hardreset the Analyser. Refer to attached Mail. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFAD230911560 | 2023-09-20 | 2023-09-21 | Cust 484 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue. | Hardware issue resolve from customer end and now data is live on CPCB. | Hardware issue resolve from customer end and now data is live on CPCB. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD230912361 | 2023-09-20 | 2023-09-22 | Cust 485 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation needs to be done. | now we are closing this complaint, we will reopen it after hardware issue is resolves from customers end. | now we are closing this complaint, we will reopen it after hardware issue is resolves from customers end. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD230917462 | 2023-09-21 | 2023-09-21 | Cust 486 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | RSPCB Camera streaming issue. | Took access of the system after changing the browser camera is live on RSPCB Server. | Took access of the system after changing the browser camera is live on RSPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD230925863 | 2023-09-24 | 2023-09-29 | Cust 484 | AHMEDABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | NOVAS CONVERTER(1) | ZLD FLOW METER AND MULTIPARAMETER DATA NOT COMING IN PC | VISITED SITE AND CABLE CHANGE AFTER COMING DATA ON PC AND RPCB. | VISITED SITE AND CABLE CHANGE AFTER COMING DATA ON PC AND RPCB. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFAD230929264 | 2023-09-25 | 2023-09-25 | Cust 473 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera streaming not working on CPCB and date time mismatch. | Camera was blocked and we installed java file. | Camera was blocked and we installed java file. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD230931365 | 2023-09-25 | 2023-09-28 | Cust 487 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | central portal data was not updating. | Issue has been resolved and data is live. | Issue has been resolved and data is live. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD230941666 | 2023-09-27 | 2023-09-28 | Cust 468 | AHMEDABAD | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 1st AMC Visit | Urvesh Visited site for AMC, Refer to attached MOM. | Urvesh visited site for AMC, Refer attached MOM | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFAD230947467 | 2023-09-27 | 2023-09-29 | Cust 488 | AHMEDABAD | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | Turbidity Meter | Optex(FMSC-P3) | CCDS temp check stat issue | Checked and it need to change . | Checked and it need to change . | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFAD230967568 | 2023-09-30 | 2023-10-02 | Cust 489 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Conductivity Transmitter Aquacon | Aquacon 2000(AQCN 2000 p) | Conductivity reading showing maximum | Urvesh guided the customer over the phone for troubleshooting. Observed issue with Cable. Rectified the same and found working ok. Refer to the attached mail. | Urvesh guided the customer over the phone for troubleshooting. Observed issue with Cable. Rectified the same and found working ok. Refer to the attached mail. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD231004869 | 2023-10-06 | 2023-10-09 | Cust 485 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Local + RSPCB data upload issue. | Cable connection is properly not connected to the local PC, sent a mail to customer regarding this, as of now we are closing this complaint. | Cable connection is properly not connected to the local PC, sent a mail to customer regarding this, as of now we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD231008970 | 2023-10-07 | 2023-10-11 | Cust 490 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Conductivity Transmitter Aquacon | Aquacon 2000(AQCN 2000 p) | Troubleshooting | Chiragbhai guided over phone, refer attached mail | Chiragbhai guided over phone, refer attached mail | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD231008970 | 2023-10-07 | 2023-10-11 | Cust 490 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Conductivity Transmitter Aquacon | Aquacon 2000(AQCN 2000 p) | Troubleshooting | Chiragbhai guided over phone, refer attached mail | Chiragbhai guided over phone, refer attached mail | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD231037471 | 2023-10-13 | 2023-10-16 | Cust 491 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Transmitter Polymetron | Ph Transmitter Polymetron 9135(9135) | commissioning complaint for pH & DO analyzer ( OA No:- 5524100451) and Multiparameter ( OA No:- 5523111564). | Guidelines for installation are provided over mail as attached. | Guidelines for installation are provided over mail as attached. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD231037471 | 2023-10-13 | 2023-10-16 | Cust 491 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Transmitter Polymetron | Ph Transmitter Polymetron 9135(9135) | commissioning complaint for pH & DO analyzer ( OA No:- 5524100451) and Multiparameter ( OA No:- 5523111564). | Guidelines for installation are provided over mail as attached. | Guidelines for installation are provided over mail as attached. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD231058272 | 2023-10-17 | 2023-10-19 | Cust 477 | AHMEDABAD | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 4th AMC visit | Urvesh visited site for troubleshooting. Recalibrated the positioners and found ok. Refer attached MOM. | Nimish Visited site for 4th AMC, Refer attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFAD231061073 | 2023-10-18 | 2023-10-22 | Cust 492 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Aqua4Trans | Turbidity Input card of Aqua4Trans(1) | Meter is not providing accurate reading. OA - 5523103784. | Urvesh visited the site for troubleshooting, observed O-ring missing in sensor installation, reading found ok. Refer attached MOM. | Urvesh visited site for Troubleshoooting, Refer attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD231062274 | 2023-10-18 | 2023-10-20 | Cust 493 | AHMEDABAD | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 1st AMC | Chiragbhai visited the site for 1st AMC, Refer attached MOM. | Chiragbhai visited site for AMC, Refer attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFAD231069975 | 2023-10-19 | 2023-10-23 | Cust 491 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Commissioning | Urvesh visited site for Commissioning, observed issue with Sensor, Transmitter working ok. Refer attached MOM. | Urvesh visited site for Commissioning, Refer attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD231069975 | 2023-10-19 | 2023-10-23 | Cust 491 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Commissioning | Urvesh visited site for Commissioning, observed issue with Sensor, Transmitter working ok. Refer attached MOM. | Urvesh visited site for Commissioning, Refer attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD231069975 | 2023-10-19 | 2023-10-23 | Cust 491 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Commissioning | Urvesh visited site for Commissioning, observed issue with Sensor, Transmitter working ok. Refer attached MOM. | Urvesh visited site for Commissioning, Refer attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD231069975 | 2023-10-19 | 2023-10-23 | Cust 491 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | Commissioning | Urvesh visited site for Commissioning, Refer attached MOM. | Urvesh visited site for Commissioning, Refer attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD231069975 | 2023-10-19 | 2023-10-23 | Cust 491 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | Commissioning | Urvesh visited site for Commissioning, Refer attached MOM. | Urvesh visited site for Commissioning, Refer attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD231070376 | 2023-10-19 | 2023-10-23 | Cust 494 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Silica Analyser | Silica 8891(8891) | Swan make Silica Analyser Commissioning | Chiragbhai and Riteshbhai visited the site to commission SWAN to make SILICA Analyser with Sequencer, Refer to attached MOM. | Chiragbhai and Riteshbhai visited site for commissioning, Refer attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD231070577 | 2023-10-19 | 2023-11-10 | Cust 491 | AHMEDABAD | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | Ph Electrode | FM1106(1) | Observed issue with all supplied pH sensors as sensors not responding with standard solutions. | Replacement sensors dispatched as per attached mail. | Replacement sensors dispatched as per attached mail. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD231070577 | 2023-10-19 | 2023-11-10 | Cust 491 | AHMEDABAD | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | Ph Electrode | FM1106(1) | Observed issue with all supplied pH sensors as sensors not responding with standard solutions. | Replacement sensors dispatched as per attached mail. | Replacement sensors dispatched as per attached mail. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD231070577 | 2023-10-19 | 2023-11-10 | Cust 491 | AHMEDABAD | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | Ph Electrode | FM1106(1) | Observed issue with all supplied pH sensors as sensors not responding with standard solutions. | Replacement sensors dispatched as per attached mail. | Replacement sensors dispatched as per attached mail. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD231088278 | 2023-10-25 | 2023-10-27 | Cust 464 | AHMEDABAD | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 3rd AMC visit | Nimish and Krutik visited the site for AMC, Refer attached MOM. | Nimish and Krutik visited site for AMC, Refer attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFAD231093579 | 2023-10-26 | 2023-10-27 | Cust 478 | AHMEDABAD | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 1st AMC Visit | Nimish and Krutik visited site for AMC, Refer attached MOM. | Nimish and Krutik visited site for AMC, Refer attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFAD231096880 | 2023-10-26 | 2023-10-31 | Cust 495 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Smartpro-con) | Commissioning of TDS sensor | Did the wiring and commission the system. | Did the wiring and commission the system. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFAD231101681 | 2023-10-27 | 2023-11-01 | Cust 496 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Software reinstallation needs to be done. | We did CPCb reconfiguration and now data is live on CPCB server. | We did CPCb reconfiguration and now data is live on CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD231146982 | 2023-11-01 | 2023-11-02 | Cust 470 | AHMEDABAD | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 2nd AMC Visit | Nimish visited site for 2nd AMC Visit, Refer attached MOM. | Nimish Visited site for 2nd AMC, Refer attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFAD231150483 | 2023-11-02 | 2023-11-03 | Cust 497 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB Central Portal issue. | The network issue has been resolved and data is live on the server. | The network issue has been resolved and data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD231166184 | 2023-11-06 | 2023-11-09 | Cust 478 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | RSPCB data uploaded issue | We have checked the reports and for COD R Flag is generating which means the negative values are uploading on the CPCB server and the RSPCB server will not accept the negative values hence it is showing NA on the RSPCB server so please check this issue from the hardware side | for COD R Flag is generating which means the negative values are uploading on the CPCB server and the RSPCB server will not accept the negative values hence it is showing NA on the RSPCB server | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD231176485 | 2023-11-08 | 2023-11-10 | Cust 491 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning of analyser supplied to M/s Aarti industries Bhachau. | Urvesh and Krutik visited site M/s Aarti for the Commissioning of the analyser, Refer attached MOM. | Urvesh and Krutik visited site for Commissioning, Refer attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD231232586 | 2023-11-22 | 2023-11-24 | Cust 498 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Analyser Servicing | Urvesh and Krutik visited the site for troubleshooting and made this analyser operational. Refer attached MOM. | Urvesh and Krutik visited site for troubleshooting, Refer attached MOM. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFAD231232586 | 2023-11-22 | 2023-11-24 | Cust 498 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Analyser Servicing | Urvesh and Krutik visited site for troubleshooting, as most of spare used form this analyser to make another analyser operational, Spares required to make this analyser healthy. Refer attached MOM. | Urvesh and Krutik visited site for troubleshooting, Refer attached MOM. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFAD231263887 | 2023-11-27 | 2023-11-28 | Cust 499 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | Commissioning of DO transmitter | Urvesh guided over the phone and commissioned the Analyser however DIP type system required to install in Tank, Refer to attached MOM. | Urvesh guided over the phone and commissioned the Analyser however DIP type system required to install in Tank, Refer to attached MOM. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD231263988 | 2023-11-27 | 2023-11-28 | Cust 500 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4Trans(1) | Reading issue while checking newly received Conductivity transmeter | Urvesh guide over phone and resolved the issue, Refer to attached MOM. | Urvesh guide over phone and resolved the issue, Refer to attached MOM. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD231285989 | 2023-11-30 | 2024-01-23 | Cust 501 | AHMEDABAD | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | MULTIPARAMETER ANALYSER | Mother board-PCX2 of Multiparameter Analyzer(1) | Sample pump motor not working | Received buyback order for CX-2000 against this analyser. Refer attached PO. Closing the complaint with this. | Received buyback order for CX-2000 against this analyser. Refer attached PO. Closing the complaint with this. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD231300590 | 2023-12-01 | 2023-12-04 | Cust 502 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data mismatch issue on the CPCB server | now the Sox and Nox readings are uploading correctly | now the Sox and Nox readings are uploading correctly | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD231301191 | 2023-12-01 | 2023-12-03 | Cust 503 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | VISI pro DO(VISI pro DO) | AQUAMAX DOT U16520211694 with VCwater DO P Arc FC10 | Chiragbhai guided over phone for commissioning, Refer attached Mail. | Chiragbhai guided over phone for commissioning, Refer attached Mail. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFAD231308892 | 2023-12-04 | 2023-12-05 | Cust 504 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Troubleshooting | Chiragbhai guided over the phone and observed an issue with the transmitter, Refer to the attached mail. | Chiragbhai guided over phone and observed issue with transmitter, Refer to attached mail. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFAD231311293 | 2023-12-04 | 2023-12-05 | Cust 473 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | RSPCB Page not working. | Took access to the system and now RSPSB Page is opening and data is live on RSPCB Server. | Took access to the system and now RSPSB Page is opening and data is live on RSPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD231311394 | 2023-12-04 | 2023-12-07 | Cust 505 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on RSPCB and date not updating even when data is live.Taken up with RSPCB. | We took access to the system and data is live on RSPCB Server.Date is not updating we have takenup till issue with RSPCB. | We took access to the system and data is live on RSPCB Server.Date is not updating we have takenup till issue with RSPCB. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD231312395 | 2023-12-04 | 2023-12-07 | Cust 506 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | RSPCB data upload issue | data is updating on RSPCB Server but the site status is showing Inactive. We have taken up with RSPCB. | data is updating on RSPCB Server but the site status is showing Inactive. We have taken up with RSPCB. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD231316296 | 2023-12-05 | 2023-12-08 | Cust 507 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not Playing on RSPCB server. | As of now closing this complaint customer did not provide static IP to create RTSP link to share RSPCB and at customer end there is camera issue. | As of now closing this complaint customer did not provide static IP to create RTSP link to share RSPCB and at customer end there is camera issue. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD231316897 | 2023-12-05 | 2023-12-08 | Cust 462 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on CPCB and RSPCB. | Took access to the system and now camera is plying on CPCB server and for RSPCB server. rtsp link created and shared with customer. | Took access to the system and now camera is plying on CPCB server and for RSPCB server. rtsp link created and shared with customer. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD231317398 | 2023-12-05 | 2023-12-07 | Cust 508 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | pH Transmitter Aquamon | Ph Transmitter Aquamon 2000(PH2000) | Reading issue | Urvesh guided over phone for troubleshooting, Suggested to change pH sensor, Refer attached Mail. | Urvesh guided over phone for troubleshooting, Suggested to change pH sensor, Refer attached Mail. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD231325999 | 2023-12-07 | 2023-12-08 | Cust 462 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data mismatch issue for COD and BOD. | Took access to the system and changed the address and now data is matching. | Took access to the system and changed the address and now data is matching. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2313447100 | 2023-12-11 | 2023-12-16 | Cust 497 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | Conductivity Transmitter not responding to the sensor. | Urvesh guided over the phone and observed an issue with the transmitters, Suggested repairing or replacing it. Refer to the attached Mail. | Urvesh guided over the phone and observed an issue with the transmitters, Suggested repairing or replacing it. Refer to the attached Mail. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2313447100 | 2023-12-11 | 2023-12-16 | Cust 497 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | Conductivity Transmitter not responding to the sensor. | Urvesh guided over the phone and observed an issue with the transmitters, Suggested repairing or replacing it. Refer to the attached Mail. | Urvesh guided over the phone and observed an issue with the transmitters, Suggested repairing or replacing it. Refer to the attached Mail. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2313602101 | 2023-12-14 | 2023-12-15 | Cust 507 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing in RSPCB. | We have created the rtsp link which is playing out of customer network which is shared by rspcb by the customer. | We have created the rtsp link which is playing out of customer network which is shared by rspcb by the customer. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2313780102 | 2023-12-18 | 2023-12-21 | Cust 485 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | RSPCB page not opening in Customers PC. | We took access to the system and data now rspcb page is working. | We took access to the system and data now rspcb page is working. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2313937103 | 2023-12-20 | 2023-12-22 | Cust 509 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | Calibration and mA issue | Urvesh guided customer over phone and resolved the issue, Refer attached mail. | Urvesh guided customer over phone and resolved the issue, Refer attached mail. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2313938104 | 2023-12-20 | 2023-12-23 | Cust 486 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera is not playing in RSPCB. | RTSP link created and shared with the customer which is playing out of their Network. | RTSP link created and shared with the customer which is playing out of their Network. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2314003105 | 2023-12-21 | 2023-12-25 | Cust 465 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Blackbox commissioning | Ajay and Krutik visited the site for the commissioning of Blackbox, Refer attached MOM. | Ajay and Krutik visited site for commissioning of Blackbox, Refer attached MOM. | YES | Paid | Core | N |  | Analytical - CHAKAN
AFAD2314027106 | 2023-12-22 | 2023-12-23 | Cust 481 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | eversense software installation + local + CPCB + GPCB configuration needs to be done. | The software configuration and CPCB + GPCB connectivity has been completed and data is live on the server. | The software configuration and CPCB + GPCB connectivity has been completed and data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD2314035107 | 2023-12-22 | 2023-12-23 | Cust 510 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | RSPCB data availability is very less. | Took access to the system and found negative readings are updating due to which data availability is less. | Took access to the system and found negative readings are updating due to which data availability is less. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2314226108 | 2023-12-26 | 2023-12-27 | Cust 458 | AHMEDABAD | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | Conductivity Transmitter Aquacon | Aquacon 2000(AQCN 2000 p) | Installation support and sensor wiring | Urvesh guided the customer over the phone to wire all the sensors as per standard. All the instruments are still at the customer plant and end user site is at Bhavnagar. Refer to the attached mail. | Urvesh guided the customer over the phone to wire all the sensors as per standard. All the instruments are still at the customer plant and end user site is at Bhavnagar. Refer to the attached mail. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2314226108 | 2023-12-26 | 2023-12-27 | Cust 458 | AHMEDABAD | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | Conductivity Transmitter Aquacon | Aquacon 2000(AQCN 2000 p) | Installation support and sensor wiring | Urvesh guided the customer over the phone to wire all the sensors as per standard. All the instruments are still at the customer plant and end user site is at Bhavnagar. Refer to the attached mail. | Urvesh guided the customer over the phone to wire all the sensors as per standard. All the instruments are still at the customer plant and end user site is at Bhavnagar. Refer to the attached mail. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2314226108 | 2023-12-26 | 2023-12-27 | Cust 458 | AHMEDABAD | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | pH Transmitter Aquamon | Ph Transmitter Aquamon 2000(PH-Trans Aqu-2000) | Installation support and sensor wiring | Urvesh guided the customer over the phone to wire all the sensors as per standard. All the instruments are still at the customer plant and end user site is at Bhavnagar. Refer to the attached mail. | Urvesh guided the customer over the phone to wire all the sensors as per standard. All the instruments are still at the customer plant and end user site is at Bhavnagar. Refer to the attached mail. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2314226108 | 2023-12-26 | 2023-12-27 | Cust 458 | AHMEDABAD | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | Aqua4Trans | Aqua4Trans(1) | Installation guidance for Turbidity Meter | Urvesh guided the customer over the phone to wire all the sensors as per standard. All the instruments are still at the customer plant and end user site is at Bhavnagar. Refer to the attached mail. | Urvesh guided the customer over the phone to wire all the sensors as per standard. All the instruments are still at the customer plant and end user site is at Bhavnagar. Refer to the attached mail. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2314226108 | 2023-12-26 | 2023-12-27 | Cust 458 | AHMEDABAD | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | Aqua4Trans | Aqua4Trans(1) | Installation guidance for Turbidity Meter | Urvesh guided the customer over the phone to wire all the sensors as per standard. All the instruments are still at the customer plant and end user site is at Bhavnagar. Refer to the attached mail. | Urvesh guided the customer over the phone to wire all the sensors as per standard. All the instruments are still at the customer plant and end user site is at Bhavnagar. Refer to the attached mail. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2314273109 | 2023-12-27 | 2023-12-30 | Cust 511 | AHMEDABAD | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | Reported issue of Conductivity not showing true value. Checked and found cable loose from sensor side and Done calibration with standard solution and now working OK. | Reported issue of Conductivity not showing true value. Checked and found cable loose from sensor side and Done calibration with standard solution and now working OK. | Reported issue of Conductivity not showing true value. Checked and found cable loose from sensor side and Done calibration with standard solution and now working OK. | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFAD2314295110 | 2023-12-27 | 2023-12-30 | Cust 512 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on RSPCB. | We took access to the system and created RTSP link and shared with customer which is playing out of their network. | We took access to the system and created RTSP link and shared with customer which is playing out of their network. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2314758111 | 2024-01-03 | 2024-01-04 | Cust 485 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | RSPCB site is not opening | RSPCB server is opening now on customers PC. | RSPCB server is opening now on customers PC. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD2314759112 | 2024-01-03 | 2024-01-04 | Cust 513 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Reconfiguration of Local + CPCB + RSPCB | We have completed Local + CPCB + RSPCB connectivity and data is live on the same. | We have completed Local + CPCB + RSPCB connectivity and data is live on the same. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD2314859113 | 2024-01-04 | 2024-01-08 | Cust 514 | AHMEDABAD | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | Erection and Installation | Alpeshbhai visited the site - WTP Gift city for the Installation/erection of Supplied Analytical instruments, Refer attached mail. | Alpeshbhai visited site for Installation/erection of Supplied Analytical instruments, Refer attached mail. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2314859113 | 2024-01-04 | 2024-01-08 | Cust 514 | AHMEDABAD | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | Erection and Installation | Alpeshbhai visited the site - WTP Gift city for the Installation/erection of Supplied Analytical instruments, Refer attached mail. | Alpeshbhai visited site for Installation/erection of Supplied Analytical instruments, Refer attached mail. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2314859113 | 2024-01-04 | 2024-01-08 | Cust 514 | AHMEDABAD | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | Erection and Installation | Alpeshbhai visited the site - WTP Gift city for the Installation/erection of Supplied Analytical instruments, Refer attached mail. | Alpeshbhai visited site for Installation/erection of Supplied Analytical instruments, Refer attached mail. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2315034114 | 2024-01-08 | 2024-01-09 | Cust 514 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Commissioning of ORP Transmitter SmartPro-8966 | Alpeshbhai visited the site for commissioning, Refer to the attached mail. | Alpeshbhai visited the site for Commissioning, Refer attached Mail. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2315034114 | 2024-01-08 | 2024-01-09 | Cust 514 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Commissioning of pH transmitter SmartPro-8966 | Alpeshbhai visited the site for Commissioning, Refer to the attached Mail. | Alpeshbhai visited the site for Commissioning, Refer attached Mail. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2315034114 | 2024-01-08 | 2024-01-09 | Cust 514 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | Commissioning of pH transmitter SmartPro-8967 | Alpeshbhai visited the site for Commissioning, Refer attached Mail. | Alpeshbhai visited the site for Commissioning, Refer attached Mail. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2315155115 | 2024-01-10 | 2024-01-12 | Cust 515 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | unable to check the real-time Dashboard on RSPCB | There is issue from RSPCB server end and we have taken up with RSPCB. | There is issue from RSPCB server end and we have taken up with RSPCB. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD2315200116 | 2024-01-11 | 2024-01-12 | Cust 478 | AHMEDABAD | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 2nd AMC Visit | Chiragbhai visited the site for the AMC visit, Refer attached MOM. | Chiragbhai visited site for AMC, Refer attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFAD2315201117 | 2024-01-11 | 2024-01-12 | Cust 464 | AHMEDABAD | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC Visit | Riteshbhai visited the site for AMC. Refer attached MOM. | Riteshbhai visited site for AMC. Refer attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFAD2315280118 | 2024-01-12 | 2024-01-12 | Cust 516 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not laying on RSPCB Server. | Created RTSP link and shared with customer which is playing out of their network. | Created RTSP link and shared with customer which is playing out of their network. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2315748119 | 2024-01-22 | 2024-01-25 | Cust 482 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Reinstallation and configuration of eversense and rspcb | We took access to the system and successfully completed the reinstallation and configuration of CPCB and RSPCB. | We took access to the system and successfully completed the reinstallation and configuration of CPCB and RSPCB. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2315793120 | 2024-01-23 | 2024-01-25 | Cust 466 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software installation and CPCB + RSPCB connectivity need to be done. | here the hardware side issue, hence As of now we are closing this complaint, and will reopen it after the hardware side issue is resolved from customers end | here the hardware side issue, hence As of now we are closing this complaint, and will reopen it after the hardware side issue is resolved from customers end | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD2315794121 | 2024-01-23 | 2024-01-25 | Cust 515 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | stack data is not coming on local server | As of now we are closing this complaint, we will reopen this once the access is given by the customer. | As of now we are closing this complaint, we will reopen this once the access is given by the customer. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD2315961122 | 2024-01-27 | 2024-01-29 | Cust 484 | AHMEDABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Callibration Done | Callibration Done | Callibration Done | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFAD2316014123 | 2024-01-29 | 2024-02-01 | Cust 496 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on RSPCB | RTSP Link shared with the customer for camera connectivity on RSPCB | RTSP Link shared with the customer for camera connectivity on RSPCB | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2316055124 | 2024-01-29 | 2024-02-01 | Cust 515 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not updating on local and RSPCB server. | We took access to the system and re-installed the software with the updated version. Also we have completed the Local and RSPCB connectivity and the data is now live on the same servers. | We took access to the system and re-installed the software with the updated version. Also we have completed the Local and RSPCB connectivity and the data is now live on the same servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD2316482125 | 2024-02-02 | 2024-02-02 | Cust 462 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on RSPCB Server. | We created the rtsp link which is playing out of customers network and shared with the customer. | We created the rtsp link which is playing out of customers network and shared with the customer. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2316487126 | 2024-02-02 | 2024-02-05 | Cust 486 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | RTSP Link shared previously is now not playing and camera is not live on RSPCB Server. | As of now closing this issue as static IP Provided by customer is not playing out of customers network. | As of now closing this issue as static IP Provided by customer is not playing out of customers network. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2316555127 | 2024-02-05 | 2024-02-06 | Cust 514 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | AQUAMAX DOT(1) | Commissioning | Chiragbhai visited the site for commissioning, Refer to the attached mail. | Chiragbhai visited site for commissioning, Refer attached Mail | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2316555127 | 2024-02-05 | 2024-02-06 | Cust 514 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | AQUAMAX DOT(1) | Commissioning | Chiragbhai visited the site for commissioning, Refer to the attached Mail | Chiragbhai visited site for commissioning, Refer attached Mail | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2316570128 | 2024-02-05 | 2024-02-06 | Cust 461 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | Reading difference with reference. | Urvesh guided over the phone for calibration and found ok. Refer attached mail. | Urvesh guided over the phone for calibration and found ok. Refer attached mail. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2316621129 | 2024-02-06 | 2024-02-09 | Cust 484 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on RSPCB Server. | We took access to the system and now data is live on RSPCB Server. | We took access to the system and now data is live on RSPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2316696130 | 2024-02-07 | 2024-02-09 | Cust 510 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Not able to login to enviroSENSE showing invalid credentials. | We took access to the system and now RSPCB Page is accessible. | We took access to the system and now RSPCB Page is accessible. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2316717131 | 2024-02-07 | 2024-02-08 | Cust 517 | AHMEDABAD | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | Ph Electrode | Inchtrode N75P (3/4 NPT with pH bulb protection, Ryton)(HT238 342) | Install 2 new pH Sensor. it is not working properly. pH sensor need to send pune for reparing. | checked the transmmitter found ok. Issue found in pH sensor | checked the transmmitter found ok. Issue found in pH sensor | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFAD2316753132 | 2024-02-08 | 2024-02-09 | Cust 505 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Need custom report to download flow data from LOCAL Server. | We took access to the system and configure custom report. | We took access to the system and configure custom report. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2316754133 | 2024-02-08 | 2024-02-09 | Cust 496 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | RSPCB Page is not accessible and Reports need to download from LOCAL SEVER | We took access to the system and configure custom report and now RSPCB Page is accessible | We took access to the system and configure custom report and now RSPCB Page is accessible | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2316756134 | 2024-02-08 | 2024-02-09 | Cust 510 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Negative reading updating on RSPCB. | Ankit Prajapati visited the site and changed some hardware side configuration and now negative reading has stopped updating oN RSPCB Server. | Ankit Prajapati visited the site and changed some hardware side configuration and now negative reading has stopped updating oN RSPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2316809135 | 2024-02-09 | 2024-02-09 | Cust 493 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | DATCON | DATCON(DATCON) | DIsplay Board configuration | We completed the Display Board configuration and now CSV files are generating in the local PC >>> D Drive >> EHSFTP folder | We completed the Display Board configuration and now CSV files are generating in the local PC &amp;gt;&amp;gt;&amp;gt; D Drive &amp;gt;&amp;gt; EHSFTP folder | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD2316810136 | 2024-02-09 | 2024-02-12 | Cust 486 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on RSPCB Server | There was internet guided customer and issue is resolved now. | There was internet guided customer and issue is resolved now. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2316888137 | 2024-02-12 | 2024-02-15 | Cust 474 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on RSPCB Server, Addition of new camera to CPCB and RSPCB Server, Contact details changes to be done on CPCB server. | Created RTSP link for ZLD camera and shared with customer. For addition of new Camera and changes in Contact details shared procedure with the customer. | Created RTSP link for ZLD camera and shared with customer. For addition of new Camera and changes in Contact details shared procedure with the customer. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2316915138 | 2024-02-12 | 2024-02-15 | Cust 510 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | RSWM -Gulabpur Data availability is very less | Ankit Prajapati visited the site and after some hardware side configuration Data availability has increased. | Ankit Prajapati visited the site and after some hardware side configuration Data availability has increased. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2316916139 | 2024-02-12 | 2024-02-15 | Cust 507 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on RSPCB Server. | Shared the rtsp link with customer which is playing out of customers network and shared with customer. | Shared the rtsp link with customer which is playing out of customers network and shared with customer. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2316939140 | 2024-02-13 | 2024-02-16 | Cust 497 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Temp. value needs to be checked as still showing ?0? on CPCB server. | we took access of the BB and we are getting 0 readings for Temperature till the Black Box, hence as of now we are closing this compliant and will reopen it after hardware side issue is resolved | we took access of the BB and we are getting 0 readings for Temperature till the Black Box, hence as of now we are closing this compliant and will reopen it after hardware side issue is resolved | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD2316942141 | 2024-02-13 | 2024-02-16 | Cust 510 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | RSWM -Mandpam data was not updating on CPCB and RSPCB Server. | We took access to the system and data is live on CPCB and MPCB. | We took access to the system and data is live on CPCB and MPCB. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2316947142 | 2024-02-13 | 2024-02-16 | Cust 518 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on RSPCB Server. | We have created the RTSP link and shared with the customer. | We have created the RTSP link and shared with the customer. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2316948143 | 2024-02-13 | 2024-02-16 | Cust 519 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB and RSPCB Server and Local host Page is not accessible. | We took access to the system and now data is live CPCB and RSPCB and local server is also now accessible. | We took access to the system and now data is live CPCB and RSPCB and local server is also now accessible. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2316955144 | 2024-02-13 | 2024-02-13 | Cust 486 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on RSPCB Server. | We have created the RTSP link which is playing out of customers network and shared with customer. | We have created the RTSP link which is playing out of customers network and shared with customer. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2317012145 | 2024-02-14 | 2024-03-29 | Cust 502 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera streaming issue on CPCB server. | There is the issue is from customers network side ZSCALLER is installed, after removing the same now camera is live on CPCB server. | There is the issue is from customers network side ZSCALLER is installed, after removing the same now camera is live on CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD2317027146 | 2024-02-14 | 2024-02-16 | Cust 520 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on RSPCB Server | We have created the rtsp link which is playing out of customers network and shared with customer. | We have created the rtsp link which is playing out of customers network and shared with customer. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2317228147 | 2024-02-19 | 2024-02-23 | Cust 505 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not updating on LOCAL+RSPCB server. | We took access to the system and data is live on CPCB and MPCB Server. | We took access to the system and data is live on CPCB and MPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2317255148 | 2024-02-19 | 2024-02-21 | Cust 490 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Conductivity Transmitter Aquacon | Conductivity Transmitter Aquamon 2000(CON-TR-2000) | AquaCon 2000Pro Transmitter not powering up | Urvesh checked the transmitter at the office and suspected an issue with the Power section, Suggested sending the same to Pune for further checking, Refer to the attached mail. | Urvesh checked the transmitter at the office and suspected an issue with the Power section, Suggested sending the same to Pune for further checking, Refer to the attached mail. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2317412149 | 2024-02-21 | 2024-02-22 | Cust 521 | AHMEDABAD | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | pH sensor is faulty | pH sensor is faulty | pH sensor is faulty | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFAD2317413150 | 2024-02-21 | 2024-02-22 | Cust 521 | AHMEDABAD | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | TDS sensor faulty | TDS sensor faulty | TDS sensor faulty | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFAD2317603151 | 2024-02-25 | 2024-02-28 | Cust 473 | AHMEDABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(Ph-Trans-Smartpro) | Calibration of PH and TSS sensor. Found PH meter faulty and sent to Division. DOne calibration of TSS sensor | Calibration of PH and TSS sensor. Found PH meter faulty and sent to Division. DOne calibration of TSS sensor | Calibration of PH and TSS sensor. Found PH meter faulty and sent to Division. DOne calibration of TSS sensor | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFAD2317603151 | 2024-02-25 | 2024-02-28 | Cust 473 | AHMEDABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(Ph-Trans-Smartpro) | Calibration of PH and TSS sensor. Found PH meter faulty and sent to Division. DOne calibration of TSS sensor | Calibration of PH and TSS sensor. Found PH meter faulty and sent to Division. DOne calibration of TSS sensor | Calibration of PH and TSS sensor. Found PH meter faulty and sent to Division. DOne calibration of TSS sensor | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFAD2317663152 | 2024-02-26 | 2024-03-07 | Cust 522 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not updating on RSPCB server | PC time was taking future time stamp after certain time. We took up this issue with RnD team and now data is live and continuously updating on RSPCB Server. | PC time was taking future time stamp after certain time. We took up this issue with RnD team and now data is live and continuously updating on RSPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2317710153 | 2024-02-27 | 2024-03-01 | Cust 512 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera streaming issue. | The time sync issue of the PC has been resolved. We have provided RTSP link to customer for RSPCB camera configuration. | The time sync issue of the PC has been resolved. We have provided RTSP link to customer for RSPCB camera configuration. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2317853154 | 2024-02-28 | 2024-02-29 | Cust 496 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Customer have changed the camera and want new rtsp link for camera configuration on RSPCB Server. | We have created RTSP link and shared with the customer. | We have created RTSP link and shared with the customer. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2318006155 | 2024-02-29 | 2024-03-02 | Cust 523 | AHMEDABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Tranmitter Smartpro | Smartpro Electronics Assembly(Smartpro-Elc-Assembly) | Calibration | Chiragbhai and Ajay visited the site for Calibration, Refer attached MOM | Chiragbhai and Ajay visited site for Calibration, Refer attached MOM. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFAD2318008156 | 2024-02-29 | 2024-03-01 | Cust 524 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | pH Transmitter Aquamon | Ph Transmitter Aquamon 2000(PH-Trans Aqu-2000) | Transmitter not powering up during commissioning. | Guided over phone and suspected issue with electronics, Recommended to send the same to Pune for further checking. Refer to the attached mail. | Guided over phone and suspected issue with electronics, Recommended to send the same to Pune for further checking. Refer to the attached mail. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2318165157 | 2024-03-04 | 2024-03-07 | Cust 464 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CAMERA ISSUE. | Now the Camera streaming is live on the CPCB server. | Now the Camera streaming is live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2318188158 | 2024-03-04 | 2024-03-07 | Cust 485 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | RSPCB data upload issue | We took access to the system and now the data is continuously reflecting to the RSPCB and Local server. | We took access to the system and now the data is continuously reflecting to the RSPCB and Local server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2318213159 | 2024-03-05 | 2024-03-08 | Cust 496 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Getting spikes in readings. | There was hardware side issue after doing changes for negative readings and zero calibration from hardware end data is live without spikes. | There was hardware side issue after doing changes for negative readings and zero calibration from hardware end data is live without spikes. | YES | No Trip | Process | Y |  | Analytical - CHAKAN
AFAD2318263160 | 2024-03-06 | 2024-03-13 | Cust 464 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on CPCB Server. | We took access to the system and Camera is live on CPCB Server. | We took access to the system and Camera is live on CPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2318413161 | 2024-03-11 | 2024-03-14 | Cust 468 | AHMEDABAD | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 2nd AMC visit | Urvesh and Krutik visited site for AMC of Analyser, Refer attached MOM. | Urvesh and Krutik visited site for AMC of Analyser, Refer attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFAD2318447162 | 2024-03-11 | 2024-03-15 | Cust 501 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning of CX2000 | Urvesh and Krutik visited the site for the commissioning of CX2000 Analyser, Refer to the attached MOM. | Urvesh and Krutik visited the site for commissioning of CX2000 Analyser, Refer to the attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2318527163 | 2024-03-13 | 2024-03-14 | Cust 510 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on RSPCB Server. | We took access to the system and camera is playing on RSPCB Server and have shared the rtsp link with the customer. | We took access to the system and camera is playing on RSPCB Server and have shared the rtsp link with the customer. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2318580164 | 2024-03-14 | 2024-03-15 | Cust 525 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Not responding to process variation at all. | Riteshbhai checked the analyser at the office and observed an issue with the sensor, Refer to the attached MOM. | Riteshbhai checked the Analyser at the office and observed an issue with sensor, Refer attached MOM. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2318613165 | 2024-03-15 | 2024-03-15 | Cust 459 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Black Box commissioning for Jodhpur Pradushan Niwaran Trust | we have completed the Black Box configuration and now data is coming till the BB and the same is verified by you also. | we have completed the Black Box configuration and now data is coming till the BB and the same is verified by you also. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD2318782166 | 2024-03-18 | 2024-03-21 | Cust 464 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on cpcb server and rspcb server | We took access to the system and camera is live on CPCB Server. | We took access to the system and camera is live on CPCB Server. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2318795167 | 2024-03-18 | 2024-03-22 | Cust 493 | AHMEDABAD | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 2nd AMC Visit | Krutik and Ajay visited the site for the AMC of Multiparameter Analyser, Refer attached MOM. | Krutik and Ajay visited site for AMC, Refer attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFAD2318795167 | 2024-03-18 | 2024-03-22 | Cust 493 | AHMEDABAD | Analytical - CHAKAN | CLOSED | AMC | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Smartpro-con) | 2nd AMC Visit | Krutik and Ajay visited site for AMC, Refer attached MOM. | Krutik and Ajay visited site for AMC, Refer attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFAD2318893168 | 2024-03-19 | 2024-03-22 | Cust 519 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB+LOCAL+RSPCB | We took access to the system and data is live on CPCB+LOCAL+MPCB. | We took access to the system and data is live on CPCB+LOCAL+MPCB. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2318975169 | 2024-03-20 | 2024-03-22 | Cust 473 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on RSPCB Server. | We took access to the system and data is live on RSPCB. | We took access to the system and data is live on RSPCB. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2319047170 | 2024-03-21 | 2024-03-22 | Cust 505 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on RSPCB Server. | We took access to the system and data is live on RSPCB. | We took access to the system and data is live on RSPCB. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2319072171 | 2024-03-21 | 2024-03-22 | Cust 464 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on RSPCB Server. | We took access to the system and data is live on RSPCB. | We took access to the system and data is live on RSPCB. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2319197172 | 2024-03-25 | 2024-03-29 | Cust 526 | AHMEDABAD | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | AQUADO Analyzer | AQUAMAX DOT(1) | most of the analysers at the site do not show the correct reading | Calibrate the analyser & working found satisfactory. | Calibrate the analyser &amp;amp; working found satisfactory. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFAD2319371173 | 2024-03-27 | 2024-03-29 | Cust 478 | AHMEDABAD | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 3rd AMC Visit | Urvesh and Chiragbhai visited the site and carried out AMC activities, Refer to the attached MOM. | Urvesh and Chiragbhai visited site for AMC, Refer attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFAD2319474174 | 2024-03-28 | 2024-03-30 | Cust 512 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Data is not updating to CPCB and RSPCB server. | Issue has been resolved and data is now live on the CPCB and RSPCB servers | Issue has been resolved and data is now live on the CPCB and RSPCB servers | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2319490175 | 2024-03-28 | 2024-03-28 | Cust 527 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on RSPCB Server. | Checked rtsp link playing out of network guided customer to take up this issue with RSPCB. | Checked rtsp link playing out of network guided customer to take up this issue with RSPCB. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2319612176 | 2024-03-29 | 2024-03-30 | Cust 528 | AHMEDABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Turbidity Meter | Thermo series(Thermo series) | Calibrate the 5Nos of Turbidity Analyser. | Calibrate the 5Nos of Turbidity Analyser. | Calibrate the 5Nos of Turbidity Analyser. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFAD2319613177 | 2024-03-29 | 2024-03-30 | Cust 528 | AHMEDABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | pH Transmitter Aquamon | Ph Transmitter Aquamon 2000(PH-Trans Aqu-2000) | Calibrate the ph sensor & transmitter | Calibrate the ph sensor & transmitter | Calibrate the ph sensor &amp;amp; transmitter | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFAD2319614178 | 2024-03-29 | 2024-03-30 | Cust 528 | AHMEDABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Transmitter Aquacon | Conductivity Transmitter Aquamon 2000(CON-TR-2000) | calibrate the sensor & Transmitter | calibrate the sensor & Transmitter | calibrate the sensor &amp;amp; Transmitter | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFAD2319615179 | 2024-03-29 | 2024-03-30 | Cust 528 | AHMEDABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | pH Transmitter Aquamon | Ph Transmitter Aquamon 4000(PH-Trans Aqu-4000) | Calibrate PH-ORP transmitter with simulator | Calibrate PH-ORP transmitter with simulator | Calibrate PH-ORP transmitter with simulator | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFAD2319662180 | 2024-03-30 | 2024-03-30 | Cust 512 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera streaming issue | Sent RTSP link to the customer. | Sent RTSP link to the customer. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2319776181 | 2024-03-31 | 2024-03-31 | Cust 528 | AHMEDABAD | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | pH Transmitter Aquamon | Ph Transmitter Aquamon 4000(PH-Trans-4000) | Checked 1 PH/ORP Meter by feeding mV from PH Simlulator followed by Nimish's Visit. | Checked 1 PH/ORP Meter by feeding mV from PH Simlulator followed by Nimish's Visit. | Checked 1 PH/ORP Meter by feeding mV from PH Simlulator followed by Nimish's Visit. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFAD2400062182 | 2024-04-02 | 2024-04-05 | Cust 502 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera Streaming issue on the CPCB server. | We have taken access and it is found that streaming windows are not running in your system and after reopening the same the camera streaming started uploading to the server. | We have taken access and it is found that streaming windows are not running in your system and after reopening the same the camera streaming started uploading to the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD2400139183 | 2024-04-03 | 2024-04-04 | Cust 464 | AHMEDABAD | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 3rd AMC | Krutik visited site for AMC of Analyser, Refer attached MOM for the activities carried out during visit. | Krutik visited site for 3rd AMC of Analyser, Refer attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFAD2400175184 | 2024-04-04 | 2024-04-05 | Cust 464 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Reconfiguration of software in new PC. | We took access to the system and successfully reconfigured the software in new PC. | We took access to the system and successfully reconfigured the software in new PC. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2400187185 | 2024-04-04 | 2024-04-05 | Cust 484 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on RSPCB Server. | We took access to the system and now data is updating on MPCB Server. | We took access to the system and now data is updating on MPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2400337186 | 2024-04-08 | 2024-04-09 | Cust 468 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Troubleshooting of reading in CX2000-3922 | Chiragbhai visited the site for troubleshooting, and after recalibration of the analyser, found it working ok. Refer attached MOM. | Chiragbhai visited the site for troubleshooting, Refer attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2400382187 | 2024-04-09 | 2024-04-11 | Cust 529 | AHMEDABAD | Analytical - CHAKAN | CLOSED | SALES PROMOTION | Aqua4Trans | Aqua4 trans-A4TW033000(1) | Commissioning of Demo Unit | Urvesh visited the site for the commissioning of the Demo Conductivity analyser, Refer attached mail. | Urvesh visited site for commissioning of Demo Analyser, Refer attached Mail. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2400409188 | 2024-04-09 | 2024-04-10 | Cust 507 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on RSPCB Server. | We took access to the system and now data is updating on RSPCB Server. | We took access to the system and now data is updating on RSPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2400478189 | 2024-04-10 | 2024-04-13 | Cust 471 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | RSPCB data uploaded issue. | now the data is live on the RSPCB server. | now the data is live on the RSPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD2400484190 | 2024-04-10 | 2024-04-11 | Cust 530 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Turbidity Meter | Optex(FMSC-P3) | Error 6 on display | Urvesh visited the site for Troubleshooting, Suspected an issue with Electronics, Refer to attached Mail. | Urvesh visited site for Troubleshooting, Refer attached Mail. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2400486191 | 2024-04-10 | 2024-04-11 | Cust 530 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Sludge Level Meter - Controller | Sludge Depth Meter 602 with Multiple Outputs(Sludge Controller 602) | Sludge Density meter of Markland speciality | Urvesh visited the site for commissioning, Refer to the attached Mail. | Urvesh visited site for commissioning, Refer attached Mail. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2400486191 | 2024-04-10 | 2024-04-11 | Cust 530 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Sludge Level Meter - Controller | Sludge Depth Meter 602 with Multiple Outputs(Sludge Controller 602) | Sludge Density meter of Markland speciality | Urvesh visited the site for commissioning, Refer to the attached Mail. | Urvesh visited site for commissioning, Refer attached Mail. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2400528192 | 2024-04-11 | 2024-04-12 | Cust 530 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Reading issue | Urvesh guided over the phone and resolved the issue after calibration with a standard solution. Refer attached mail. | Urvesh guided over the phone and resolved the issue after calibration with a standard solution. Refer attached mail. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2400560193 | 2024-04-12 | 2024-04-13 | Cust 531 | AHMEDABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | pH Transmitter Aquamon | Ph Transmitter Aquamon 2000(PH-Trans Aqu-2000) | pH Calibration | Alpeshbhai guided over the phone for pH calibration and resolved the issue, Refer attached Mail. | Alpeshbhai guided over the phone for pH calibration and resolved the issue, Refer attached Mail. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2400563194 | 2024-04-12 | 2024-04-12 | Cust 468 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Internet configuration needs to be done. | We have configured the internet IP details as provided by you so now the internet is pinging till the Black Box. | We have configured the internet IP details as provided by you so now the internet is pinging till the Black Box. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD2400992195 | 2024-04-23 | 2024-04-26 | Cust 496 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on RSPCB Server. | As of now we are closing this complaint. We will reopen this complaint once we get access. | As of now we are closing this complaint. We will reopen this complaint once we get access. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2401031196 | 2024-04-23 | 2024-04-24 | Cust 524 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | pH Transmitter Aquamon | Ph Transmitter Aquamon 2000(PH2000) | mA output range setting | Urvesh guided over phone for necessary configuration and help in commissioning, Refer attached mail. | Urvesh guided over phone for necessary configuration and help in commissioning, Refer attached mail. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2401031196 | 2024-04-23 | 2024-04-24 | Cust 524 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4 trans-A4TW033000(1) | mA output and TDS configuration | Urvesh guided over phone for necessary configuration and help in commissioning, Refer attached mail. | Urvesh guided over phone for necessary configuration and help in commissioning, Refer attached mail. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2401064197 | 2024-04-24 | 2024-04-26 | Cust 532 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data availability is less on RSPCB | We took access to the system and found due to internet connectivity data is breaking on RSPCB Server guided customer to check from their end. | We took access to the system and found due to internet connectivity data is breaking on RSPCB Server guided customer to check from their end. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2401077198 | 2024-04-24 | 2024-04-26 | Cust 510 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data availability is less on RSPCB Server. | We checked CPCB and RSPCB data. Data was breaking on both the server. Now data is live and data availability is 99.86% now. | We checked CPCB and RSPCB data. Data was breaking on both the server. Now data is live and data availability is 99.86% now. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2401106199 | 2024-04-24 | 2024-04-25 | Cust 483 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Reading issue | Urvesh guided over the phone and observed an issue with the Transmitter. Suggested to send the same to Pune for further check/repair. Refer attached mail. | Urvesh guided over the phone and observed an issue with the Transmitter. Suggested to send the same to Pune for further check/repair. Refer attached mail. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2401150200 | 2024-04-25 | 2024-04-26 | Cust 507 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB and RSPCB data uploaded issue. | Today we have reinstalled the software and completed the connectivity so now the data is live on the RSPCB server | Today we have reinstalled the software and completed the connectivity so now the data is live on the RSPCB server | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD2401178201 | 2024-04-25 | 2024-04-27 | Cust 529 | AHMEDABAD | Analytical - CHAKAN | CLOSED | SALES PROMOTION | Aqua4Trans | Aqua4 trans-A4TW033000(1) | Discussion about Demo unit conductivity reading issue | Chiragbhai alongwith Ankitbhai and Mikita visited site for resolution of customer query in Demo Analyser, Refer attached Mail. | Chiragbhai alongwith Ankitbhai and Mikita visited site for resolution of customer query in Demo Analyser, Refer attached Mail. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2401189202 | 2024-04-25 | 2024-04-30 | Cust 458 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | AQUAMAX DOT(1) | Commissioning | Urvesh guided over the phone for sensor installation and wiring however power-up is pending from the customer side, Refer attached MOM. | Urvesh guided over the phone for sensor installation and wiring however power-up is pending from the customer side, Refer attached MOM. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2401361203 | 2024-04-29 | 2024-05-01 | Cust 486 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera is not playing on RSPCB | Issue is resolved | Issue is resolved | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2401425204 | 2024-04-30 | 2024-05-06 | Cust 520 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Eversense software installation and configuration. | We took access to your system and now data is live on CPCB + Local + RSPCB. Please find below attached screenshot for your reference. | We took access to your system and now data is live on CPCB + Local + RSPCB. Please find below attached screenshot for your reference. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2401459205 | 2024-04-30 | 2024-05-01 | Cust 486 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | RSPCB Camera IP changed by customer need to configure new rtsp link for camera. | We have created new rtsp link which is playing out of customers network and shared with customer. | We have created new rtsp link which is playing out of customers network and shared with customer. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2401464206 | 2024-04-30 | 2024-05-03 | Cust 533 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to RSPCB. | We have updated (EnviroConnectDAS),and now data is live on both the server LOCAL + RSPCB. | We have updated (EnviroConnectDAS),and now data is live on both the server LOCAL + RSPCB. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2401588207 | 2024-05-01 | 2024-05-04 | Cust 504 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4Trans(1) | pH transmitter with Relay output | Urvesh guided over phone and completed the commissioning of Analyser, Refer to the attached Mail. | Urvesh guided over phone and completed the commissioning of Analyser, Refer to the attached Mail. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2401684208 | 2024-05-04 | 2024-05-07 | Cust 484 | AHMEDABAD | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration Done | Calibration Done | Calibration Done | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFAD2401811209 | 2024-05-07 | 2024-05-10 | Cust 459 | AHMEDABAD | DCS - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | MICROCON+ER SERIES | CONTROLLER MODULE (ER-201)(ER-201) | Data is not uploading to RSPCB server | As of now we are closing this complaint, we will reopen this once the internet issue is resolved from the customer end. | As of now we are closing this complaint, we will reopen this once the internet issue is resolved from the customer end. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2401830210 | 2024-05-07 | 2024-05-13 | Cust 534 | AHMEDABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Smartpro Electronics Assembly(Smartpro-Elc-cbl) | The instruments supplied against 1243/17 pH & DO meter are not showing proper reading. So please send your person to rectify the same. Site Address: HINDUSTAN COCA-COLA BEVERAGES PVT LTD. HCCB Sanand (Guj) Plant, E-569, Sanand Industrial Estate #II, Ahmedabad - 382110 24 Gujarat (India) Site person contact details: Mr. Prokash :86174 48402 Mr. Krishna : 93047 17787 | Visited site and checked all meters. attached MOM For reference | Visited site and checked all meters. attached MOM For reference | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFAD2401871211 | 2024-05-08 | 2024-05-09 | Cust 478 | AHMEDABAD | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 4th AMC visit and troubleshooting | Urvesh and Krutik visited the site for troubleshooting and AMC, Refer attached MOM. | Urvesh and Krutik visited site for troubleshooting and AMC, Refer attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFAD2401888212 | 2024-05-08 | 2024-11-11 | Cust 494 | AHMEDABAD | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | TOC ANALYSER | TOC 4000 Dual Stream ANALYSER ANATEL(4000 TOC) | Commissioning | Visited site with new detector card. Observed one more issue with the connector of cable connecting detector board with Micro Board. With temporary adjustments, commissioned the analyser. However need new cable. The new required cable was received at the Ahmedabad Branch on 1st week of November. Hence closing this complaint. | The new required cable was received at the Ahmedabad Branch. Hence closing this complaint. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2401947213 | 2024-05-10 | 2024-05-10 | Cust 493 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | DATCON | DATCON(DATCON) | CPCB and GPCB data uploded issue. | Problem has been sorted out with some minor change in loose cabling, now data transfer started and displayed from CPCB server. | Problem has been sorted out with some minor change in loose cabling, now data transfer started and displayed from CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD2402076214 | 2024-05-14 | 2024-05-17 | Cust 462 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is mismach of COD BOD on CPCB server. | We took access to the system and checked data on modscan and same data is updating on CPCB Server. | We took access to the system and checked data on modscan and same data is updating on CPCB Server. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2402387215 | 2024-05-21 | 2024-05-23 | Cust 477 | AHMEDABAD | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 1st AMC | Ajay visited the site for the 1st AMC of Multiparameter Analyser, Refer to the attached MOM. | Ajay visited site for AMC, Refer attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFAD2402412216 | 2024-05-21 | 2024-05-23 | Cust 480 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Smartpro-con) | Reading Zero issue | Urvesh guided over phone and resolved the issue, Refer attached Mail. | Urvesh guided over phone and resolved the issue, Refer attached Mail. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2402414217 | 2024-05-21 | 2024-05-24 | Cust 474 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | RSPCB camera confguration. | We created rtsp link and shared with the customer. | We created rtsp link and shared with the customer. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2402442218 | 2024-05-22 | 2024-05-23 | Cust 470 | AHMEDABAD | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 3rd AMC Visit | Urvesh visited the site for the 3rd AMC of Multiparameter Analyser at Bithuja CETP, Refer to the attached MOM. | Urvesh visited site for 3rd AMC, Refer attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFAD2402456219 | 2024-05-22 | 2024-05-23 | Cust 535 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4Trans(1) | Chlorine Analyser installation | Chiragbhai visited the site for installation checking, Observed lot of deposition in flow through chamber. Recommended to provide strainer in sample line. Refer to the attached mail. | Chiragbhai visited site for installation checking, Refer attached mail. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2402456219 | 2024-05-22 | 2024-05-23 | Cust 535 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Turbidity Input card of Aqua4Trans(1) | Turbidity Analyser Commissioning | Chiragbhai visited the site for commissioning and checked the reading with 0 and 35 NTU solutions. Refer attached mail. | Chiragbhai visited site for installation checking, Refer attached mail. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2402456219 | 2024-05-22 | 2024-05-23 | Cust 535 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Turbidity Input card of Aqua4Trans(1) | Turbidity Analyser Commissioning | Chiragbhai visited the site for commissioning and checked the reading with 0 and 35 NTU solution. Refer attached mail. | Chiragbhai visited site for installation checking, Refer attached mail. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2402669220 | 2024-05-27 | 2024-05-30 | Cust 528 | AHMEDABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Done calibration of various PH meters, Conductivity meters. | Done calibration of various PH meters, Conductivity meters. | Done calibration of various PH meters, Conductivity meters. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFAD2402812221 | 2024-05-28 | 2024-05-30 | Cust 474 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not streaming on RSPCB Server. | Created RTSP link and shared with customer. | Created RTSP link and shared with customer. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2402886222 | 2024-05-29 | 2024-06-14 | Cust 471 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | RSPCB data uploading issue. | We have made some changes from the software end. Also there was an issue with FTP. We have added an FTP port in the firewall. | We have made some changes from the software end. Also there was an issue with FTP. We have added an FTP port in the firewall. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2402897223 | 2024-05-29 | 2024-05-31 | Cust 474 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on SPCB Server. | Created rtsp link and shared with customer. | Created rtsp link and shared with customer. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2402900224 | 2024-05-29 | 2024-05-31 | Cust 470 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CBCB + RSPCB data upload issue. | As of now we are closing this complaint and will reopen it once the hardware side issue is resolved. | As of now we are closing this complaint and will reopen it once the hardware side issue is resolved. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD2403027225 | 2024-05-31 | 2024-06-01 | Cust 512 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | .PTZ Camera live Streaming Inactive. It is not connected to the RSPCB server | The camera is not playing in customers network. | The camera is not playing in customers network. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2403044226 | 2024-05-31 | 2024-06-03 | Cust 536 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | Commissioning | Chiragbhai and Ajay visited Renew Photovoltaic site for commissioning, Refer to the attached Mail for details. | Chiragbhai and Ajay visited Renew Photovoltaic site for commissioning, Refer to the attached Mail for details. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2403044226 | 2024-05-31 | 2024-06-03 | Cust 536 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | Commissioning | Chiragbhai and Ajay visited Renew Photovoltaic site for commissioning, Refer to the attached Mail for details. | Chiragbhai and Ajay visited Renew Photovoltaic site for commissioning, Refer to the attached Mail for details. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2403044226 | 2024-05-31 | 2024-06-03 | Cust 536 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | Commissioning | Chiragbhai and Ajay visited Renew Photovoltaic site for commissioning, Refer to the attached Mail for details. | Chiragbhai and Ajay visited Renew Photovoltaic site for commissioning, Refer to the attached Mail for details. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2403044226 | 2024-05-31 | 2024-06-03 | Cust 536 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(Ph-Trans-Smartpro) | Commissioning | Chiragbhai and Ajay visited Renew Photovoltaic site for commissioning, Refer to the attached Mail for details. | Chiragbhai and Ajay visited Renew Photovoltaic site for commissioning, Refer to the attached Mail for details. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2403044226 | 2024-05-31 | 2024-06-03 | Cust 536 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(Ph-Trans-Smartpro) | Commissioning | Chiragbhai and Ajay visited Renew Photovoltaic site for commissioning, Refer to the attached Mail for details. | Chiragbhai and Ajay visited Renew Photovoltaic site for commissioning, Refer to the attached Mail for details. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2403044226 | 2024-05-31 | 2024-06-03 | Cust 536 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(Ph-Trans-Smartpro) | Commissioning | Chiragbhai and Ajay visited Renew Photovoltaic site for commissioning, Refer to the attached Mail for details. | Chiragbhai and Ajay visited Renew Photovoltaic site for commissioning, Refer to the attached Mail for details. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2403044226 | 2024-05-31 | 2024-06-03 | Cust 536 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(Ph-Trans-Smartpro) | Commissioning | Chiragbhai and Ajay visited Renew Photovoltaic site for commissioning, Refer to the attached Mail for details. | Chiragbhai and Ajay visited Renew Photovoltaic site for commissioning, Refer to the attached Mail for details. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2403044226 | 2024-05-31 | 2024-06-03 | Cust 536 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Commissioning of ORP transmitter | Chiragbhai and Ajay visited Renew Photovoltaic site for commissioning, Refer to the attached Mail for details. | Chiragbhai and Ajay visited Renew Photovoltaic site for commissioning, Refer to the attached Mail for details. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2403044226 | 2024-05-31 | 2024-06-03 | Cust 536 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Turbidity Input card of Aqua4Trans(1) | Commissioning of Turbidity Analyer | Chiragbhai and Ajay visited Renew Photovoltaic site for commissioning, Refer to the attached Mail for details. | Chiragbhai and Ajay visited Renew Photovoltaic site for commissioning, Refer to the attached Mail for details. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2403044226 | 2024-05-31 | 2024-06-03 | Cust 536 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Turbidity Input card of Aqua4Trans(1) | Commissioning of Turbidity Analyer | Chiragbhai and Ajay visited Renew Photovoltaic site for commissioning, Refer to the attached Mail for details. | Chiragbhai and Ajay visited Renew Photovoltaic site for commissioning, Refer to the attached Mail for details. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2403103227 | 2024-05-31 | 2024-05-31 | Cust 528 | AHMEDABAD | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Smartpro-con) | Reset the smartpro & calibrate the sensor working found satisfactory | Reset the smartpro & calibrate the sensor working found satisfactory | Reset the smartpro &amp;amp; calibrate the sensor working found satisfactory | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFAD2403255228 | 2024-06-04 | 2024-06-07 | Cust 505 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Hardware changed from customer end need to update camera link on RSPCB Server. | Created RTSP link and shared with customer. | Created RTSP link and shared with customer. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2403258229 | 2024-06-04 | 2024-06-07 | Cust 537 | AHMEDABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Turbidity Meter | Optex turbidity sensor(1) | Calibration | Chiragbhai and Ajay visited the site for the calibration of analysers, Refer to the attached MOM. | Chiragbhai and Ajay visited site for calibration of analysers, Refer attached MOM. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFAD2403258229 | 2024-06-04 | 2024-06-07 | Cust 537 | AHMEDABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Chlorine Analyzer-MAC 100 | MAC 100(1) | Calibration | Chiragbhai and Ajay visited the site for the calibration of analysers, Refer to the attached MOM. | Chiragbhai and Ajay visited site for calibration of analysers, Refer attached MOM. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFAD2403258229 | 2024-06-04 | 2024-06-07 | Cust 537 | AHMEDABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Calibration | Chiragbhai and Ajay visited the site for the calibration of analysers, Refer to the attached MOM. | Chiragbhai and Ajay visited site for calibration of analysers, Refer attached MOM. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFAD2403258229 | 2024-06-04 | 2024-06-07 | Cust 537 | AHMEDABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Calibration | Chiragbhai and Ajay visited the site for the calibration of analysers, Refer to the attached MOM. | Chiragbhai and Ajay visited site for calibration of analysers, Refer attached MOM. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFAD2403258229 | 2024-06-04 | 2024-06-07 | Cust 537 | AHMEDABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Calibration | Chiragbhai and Ajay visited the site for the calibration of analysers, Refer to the attached MOM. | Chiragbhai and Ajay visited site for calibration of analysers, Refer attached MOM. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFAD2403258229 | 2024-06-04 | 2024-06-07 | Cust 537 | AHMEDABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Aqua4Trans | Turbidity Input card of Aqua4Trans(1) | Calibration | Chiragbhai and Ajay visited the site for the calibration of analysers, Refer to the attached MOM. | Chiragbhai and Ajay visited site for calibration of analysers, Refer attached MOM. | YES | Paid | Core | N |  | Analytical - CHAKAN
AFAD2403258229 | 2024-06-04 | 2024-06-07 | Cust 537 | AHMEDABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Aqua4Trans | Turbidity Input card of Aqua4Trans(1) | Calibration | Chiragbhai and Ajay visited the site for the calibration of analysers, Refer to the attached MOM. | Chiragbhai and Ajay visited site for calibration of analysers, Refer attached MOM. | YES | Paid | Core | N |  | Analytical - CHAKAN
AFAD2403258229 | 2024-06-04 | 2024-06-07 | Cust 537 | AHMEDABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Aqua4Trans | Turbidity Input card of Aqua4Trans(1) | Calibration | Chiragbhai and Ajay visited the site for the calibration of analysers, Refer to the attached MOM. | Chiragbhai and Ajay visited site for calibration of analysers, Refer attached MOM. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFAD2403258229 | 2024-06-04 | 2024-06-07 | Cust 537 | AHMEDABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Aqua4Trans | Aqua4Trans(1) | Calibration of Chlorine Analyser | Chiragbhai and Ajay visited the site for the calibration of analysers, Refer to the attached MOM. | Chiragbhai and Ajay visited site for calibration of analysers, Refer attached MOM. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFAD2403258229 | 2024-06-04 | 2024-06-07 | Cust 537 | AHMEDABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Aqua4Trans | Aqua4Trans(1) | Calibration of Chlorine Analyser | Chiragbhai and Ajay visited the site for the calibration of analysers, Refer to the attached MOM. | Chiragbhai and Ajay visited site for calibration of analysers, Refer attached MOM. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFAD2403298230 | 2024-06-05 | 2024-06-10 | Cust 520 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Eversense software is not working, and RSPCB server camera issue. | Now the camera streaming is live on the CPCB server. | We have checked that eversense software is working properly, and the data is also live on CPCB + Local + RSPCB server and ,Now the camera streaming is live on the CPCB server. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2403510231 | 2024-06-10 | 2024-06-13 | Cust 538 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4Trans(1) | Commissioning of Aqua4trans with 02 Nos. Conductivity Sensor. | Urvesh guided over the phone for commissioning and found the analysers working ok. Refer attached mail. | Urvesh guided over the phone for commissioning and found the analysers working ok. Refer attached mail. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2403510231 | 2024-06-10 | 2024-06-13 | Cust 538 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4Trans(1) | Commissioning of Aqua4trans with pH and ORP Sensor. | Urvesh guided over the phone for commissioning and found the analysers working ok. Refer attached mail. | Urvesh guided over the phone for commissioning and found the analysers working ok. Refer attached mail. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2403531232 | 2024-06-10 | 2024-06-14 | Cust 497 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on CPCB | We have checked and observed that the internet is not pinging till the Blackbox hence as of now we are closing this complaint and will reopen this case once the internet issue has been resolved from the customer end. | as of now we are closing this complaint and will reopen this case once the internet issue has been resolved from the customer end. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2403755233 | 2024-06-16 | 2024-06-20 | Cust 539 | AHMEDABAD | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | MULTIPARAMETER ANALYSER | Xenon Lamp Power supply for Multiparameter(AWA-XLAMPPS) | Checked 12V in Xenon lamp power supply it showing zero. After change the card Analyser working found satisfactory | Checked 12V in Xenon lamp power supply it showing zero. After change the card Analyser working found satisfactory | Checked 12V in Xenon lamp power supply it showing zero. After change the card Analyser working found satisfactory | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFAD2403757234 | 2024-06-16 | 2024-06-20 | Cust 528 | AHMEDABAD | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | Chlorine Analyzer-MAC 100 | MAC 100(1) | Clean the sensor & calibrate the sensor. Chlorine Analyser working found satisfactory | Clean the sensor & calibrate the sensor. Chlorine Analyser working found satisfactory | Clean the sensor &amp;amp; calibrate the sensor. Chlorine Analyser working found satisfactory | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFAD2403828235 | 2024-06-17 | 2024-06-20 | Cust 522 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera link changed need to configure camera on RSPCB Server. | We have created and shared the rtsp link with the customer. | We have created and shared the rtsp link with the customer. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2403897236 | 2024-06-18 | 2024-06-24 | Cust 473 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(Ph-Trans-Smartpro) | Repaired PH Transmitter has been installed and commissioned with newly supplied PH Electrode. | Repaired PH Transmitter has been installed and commissioned with newly supplied PH Electrode. | Repaired PH Transmitter has been installed and commissioned with newly supplied PH Electrode. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFAD2403930237 | 2024-06-18 | 2024-06-20 | Cust 512 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Camera Streaming issue | We have created an RTSP link and send to the customer for RSPCB configuration | We have created an RTSP link and send to the customer for RSPCB configuration | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2404052238 | 2024-06-21 | 2024-06-21 | Cust 462 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on RSPCB server. | Now data is live on the RSPCB server. | Now data is live on the RSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2404067239 | 2024-06-21 | 2024-06-21 | Cust 505 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera showing inactive on RSPCB Server. | We have created the rtsp link and shared with the RSPCB for configuration. | We have created the rtsp link and shared with the RSPCB for configuration. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2404098240 | 2024-06-22 | 2024-06-23 | Cust 528 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | Dissolved Oxygen Analyzer Panel mounted(AQUADO4000P) | Install the sensor. Sensor working found properly. | Install the sensor. Sensor working found properly. | Install the sensor. Sensor working found properly. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFAD2404099241 | 2024-06-22 | 2024-06-23 | Cust 528 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Calibrate the sensor & found ok. also Checked in lab & Found ok. | Calibrate the sensor & found ok. also Checked in lab & Found ok. | Calibrate the sensor &amp;amp; found ok. also Checked in lab &amp;amp; Found ok. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFAD2404108242 | 2024-06-22 | 2024-06-27 | Cust 528 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Electrode | Polilyte PRO 120(HT238 411) | Checked the electrode dip in 4,7,10 bt Sensor not response. Also checked transmmitter & found ok | New pH Sensor Is dispatch from Pune to customer. | New pH Sensor Is dispatch from Pune to customer. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFAD2404185243 | 2024-06-24 | 2024-06-27 | Cust 510 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | optex series(optex series) | Optex Transmitter and Sensor Commissioning and done configuration for Blower settings. | Optex Transmitter and Sensor Commissioning and done configuration for Blower settings. | Optex Transmitter and Sensor Commissioning and done configuration for Blower settings. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFAD2404257244 | 2024-06-25 | 2024-06-29 | Cust 464 | AHMEDABAD | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC visit | Krutik and Ajay visited site on 25.06.24 and completed AMC activity, later issue was observed in analyst and the same reset by Customer. Urvesh visited site on 28.06.24 for calibration of analyser. Refer attached MOMs. | Urvesh, Krutik and Ajay visited site for AMC , Refer attached MOMs. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFAD2404274245 | 2024-06-25 | 2024-07-01 | Cust 524 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Aqua4Trans | Display card of Aqua4Trans(1) | Display flickering issue reported by customer | Urvesh guided the customer over the phone and resolved the issue by decreasing the contrast of the display, Refer attached Mail. | Urvesh guided the customer over the phone and resolved the issue by decreasing the contrast of the display, Refer attached Mail. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2404274245 | 2024-06-25 | 2024-07-01 | Cust 524 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Aqua4Trans | Display card of Aqua4Trans(1) | Display flickering issue reported by customer | Urvesh guided the customer over the phone and resolved the issue by decreasing the contrast of the display, Refer attached Mail. | Urvesh guided the customer over the phone and resolved the issue by decreasing the contrast of the display, Refer attached Mail. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2404300246 | 2024-06-26 | 2024-06-29 | Cust 464 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB + RSPCB ISSUE. | now the Data is live on the CPCB server. | now the Data is live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2404309247 | 2024-06-26 | 2024-06-29 | Cust 505 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera streaming is not live on RSPCB server. | ETP CAMERA and ZLD CAMERA is playing out of customer's network. | ETP CAMERA and ZLD CAMERA is playing out of customer's network. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2404385248 | 2024-06-28 | 2024-06-29 | Cust 464 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | Aquamax system(1) | Checking of DO sensors | Urvesh visited the site for checking of DO sensors, Refer attached MOM. | Urvesh visited site for Checking of DO sensors, Refer attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2404385248 | 2024-06-28 | 2024-06-29 | Cust 464 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | Aquamax system(1) | Checking of DO sensors | Urvesh visited site for Checking of DO sensors, Refer attached MOM. | Urvesh visited site for Checking of DO sensors, Refer attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2404497249 | 2024-06-29 | 2024-06-29 | Cust 502 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | REMOVE DECIMAL POINT from the server. | we have made modifications so now the correct readings are uploaded. | we have made modifications so now the correct readings are uploaded. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD2404811250 | 2024-07-03 | 2024-07-04 | Cust 540 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | Reading issue. Observed very low reading | Urvesh guided over the phone and observed an issue with the Cable connector, Resolved the issue and found working ok. Refer attached Mail. | Urvesh guided over the phone and observed an issue with the Cable connector, Resolved the issue and found working ok. Refer attached Mail. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2405088251 | 2024-07-09 | 2024-07-10 | Cust 501 | AHMEDABAD | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 3rd AMC Visit | Ajay and Jigar visited the site for Preventive Maintenance, Refer to the attached MOM. | Ajay and Jigar visited site for Preventive Maintenance, Refer attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFAD2405406252 | 2024-07-15 | 2024-07-19 | Cust 462 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on CPCB + Local + RSPCB. | We took access to the system and successfully completed the software re-installation. And data is live on CPCB server for ETP Outlet. But currently data is not communicating with the system for CETP Inlet Flow and CETP Outlet Flow.so the issue is from hardware side. so we are closing this complaint. | data is live on CPCB server for ETP Outlet. CETP Inlet Flow and CETP Outlet Flow.There is a issue is from hardware side. so we are closing this complaint. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2405412253 | 2024-07-15 | 2024-07-19 | Cust 495 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to CPCB and RSPCB server. | Data is now live on the CPCB & RSPCB server. | Data is now live on the CPCB &amp;amp; RSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2405614254 | 2024-07-19 | 2024-07-20 | Cust 480 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | The transmitter not powering up | Urvesh guided over the phone and suspected an issue with the electronics, Suggested sending the electronics to Pune for further checking/repair. Refer to the attached mail. | Urvesh guided over the phone and suspected an issue with the electronics, Suggested sending the electronics to Pune for further checking/repair. Refer to the attached mail. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2405813255 | 2024-07-23 | 2024-07-25 | Cust 541 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | CX2000 Analyser Commissioning | Chiragbhai with Ajay visited the site for Commissioning, Refer attached MOM. | Chiragbhai with Ajay visited site for Commissioning, Refer attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2405921256 | 2024-07-25 | 2024-07-26 | Cust 462 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera streaming issue | As of now we are closing this complaint. We will reopen this once we get camera IP details | As of now we are closing this complaint. We will reopen this once we get camera IP details | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2405965257 | 2024-07-25 | 2024-07-27 | Cust 541 | AHMEDABAD | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | During the Commissioning of CX2000, observed Modbus output connections reversed at the card end | Chiragbhai has changed the wiring of the connector during his commissioning visit, Refer to the attached visit MOM. | Chiragbhai has changed the wiring of the connector during his commissioning visit, Refer to the attached visit MOM. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2406028258 | 2024-07-26 | 2024-07-29 | Cust 502 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Blackbox configuration and CPCB + GPCB connectivity needs to be done. | we have completed the CPCB and GPCB connectivity through FMBB for 43 MW and 80 MW and data is live on the server for both the units. | we have completed the CPCB and GPCB connectivity through FMBB for 43 MW and 80 MW and data is live on the server for both the units. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD2406212259 | 2024-07-29 | 2024-07-30 | Cust 484 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | RSPCB configuration needs to be done. | we took access to your system and completed RSPCB configuration and now data is live on the RSPCB server. | we took access to your system and completed RSPCB configuration and now data is live on the RSPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD2406247260 | 2024-07-29 | 2024-08-01 | Cust 474 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Data not updating on CPCB and Local server. | We took access to the system and now data is live on CPCB+LOCAL Server. | We took access to the system and now data is live on CPCB+LOCAL Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2406324261 | 2024-07-30 | 2024-08-02 | Cust 513 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | RSPCB data uploading issue. | Issue has been resolved and the data is now live on the RSPCB server. | Issue has been resolved and the data is now live on the RSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2406514262 | 2024-08-01 | 2024-08-02 | Cust 542 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | pH input card-FM of Multiparameter(1) | Multiparameter analyzer commissioning . | Site details not received through complaint or mail So closing the complaint for now. | Site details not received through complaint or mail So closing the complaint for now. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2406523263 | 2024-08-01 | 2024-08-02 | Cust 522 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Data avaliability is very low data not updating on the CPCB server. | There was Hardware side issue from Customer end when resolved data stared updating on CPCB+LOCAL+RSCB server. | There was Hardware side issue from Customer end when resolved data stared updating on CPCB+LOCAL+RSCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2406797264 | 2024-08-07 | 2024-08-12 | Cust 542 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning of CX2000 | Chiragbhai visited the site for commissioning, Refer attached MOM. | Chiragbhai visited site for Commissioning, Refer atatched MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2406840265 | 2024-08-08 | 2024-08-09 | Cust 513 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera connectivity issue to RSPCB server | We have created an RTSP link which is playing out of customers network and sent a mail. | We have created an RTSP link which is playing out of customers network and sent a mail. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2406871266 | 2024-08-08 | 2024-08-11 | Cust 537 | AHMEDABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Chlorine Analyzer-MAC 100 | MAC 100(1) | Calibration | Urvesh visited the site and calibrated the analyser, Refer to the attached MOM. | Urvesh visited the site and calibrated the analyser, Refer to the attached MOM. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFAD2406871266 | 2024-08-08 | 2024-08-11 | Cust 537 | AHMEDABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Turbidity Meter | Optex(FMSC-P3) | Calibration | Urvesh visited the site and calibrated the analyser, Refer to the attached MOM. | Urvesh visited the site and calibrated the analyser, Refer to the attached MOM. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFAD2406871266 | 2024-08-08 | 2024-08-11 | Cust 537 | AHMEDABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Aqua4Trans | Turbidity Input card of Aqua4Trans(1) | Calibration | Urvesh visited the site and calibrated the analyser, Refer to the attached MOM. | Urvesh visited the site and calibrated the analyser, Refer to the attached MOM. | YES | Paid | Core | N |  | Analytical - CHAKAN
AFAD2406871266 | 2024-08-08 | 2024-08-11 | Cust 537 | AHMEDABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Aqua4Trans | Turbidity Input card of Aqua4Trans(1) | Calibration | Urvesh visited the site and calibrated the analyser, Refer to the attached MOM. | Urvesh visited the site and calibrated the analyser, Refer to the attached MOM. | YES | Paid | Core | N |  | Analytical - CHAKAN
AFAD2406871266 | 2024-08-08 | 2024-08-11 | Cust 537 | AHMEDABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Aqua4Trans | Turbidity Input card of Aqua4Trans(1) | Calibration | Urvesh visited the site and calibrated the analyser, Refer to the attached MOM. | Urvesh visited the site and calibrated the analyser, Refer to the attached MOM. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFAD2406871266 | 2024-08-08 | 2024-08-11 | Cust 537 | AHMEDABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Aqua4Trans | Aqua4Trans(1) | Calibration of Chlorine sensor installed with Aqua4trans | Urvesh visited the site and calibrated the analyser, Refer to the attached MOM. | Urvesh visited the site and calibrated the analyser, Refer to the attached MOM. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFAD2406871266 | 2024-08-08 | 2024-08-11 | Cust 537 | AHMEDABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Aqua4Trans | Aqua4Trans(1) | Calibration of Chlorine sensor installed with Aqua4trans | Urvesh visited the site and calibrated the analyser, Refer to the attached MOM. | Urvesh visited the site and calibrated the analyser, Refer to the attached MOM. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFAD2406980267 | 2024-08-12 | 2024-08-14 | Cust 543 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Eversense installation and configuration. | As of now closing this issue. We will reopen once camera issue is resolved from customer end. | As of now closing this issue. We will reopen once camera issue is resolved from customer end. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2407000268 | 2024-08-12 | 2024-08-13 | Cust 474 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Reinstallation and configuration of software old PC Hang issue. | We took access to the system and successfully completed the software configuration and now data is live on LOCAL+CPCB+RSPCB Server. | We took access to the system and successfully completed the software configuration and now data is live on LOCAL+CPCB+RSPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2407098269 | 2024-08-14 | 2024-08-20 | Cust 544 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | Service support is required for the newly supplied analyser . | As per discussion over the phone with the Site person Mr. Raja, The installation of all analytical instruments is ongoing and will inform us after completion of the same for further commissioning. So closing the complaint for now. | As per discussion over the phone with the Site person Mr. Raja, The installation of all analytical instruments is ongoing and will inform us after completion. So closing the complaint for now. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2407107270 | 2024-08-14 | 2024-08-16 | Cust 512 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on RSPCB Server. | As of now we are closing this complaint and we will reopen this once the camera issue is resolved from customer end. | As of now we are closing this complaint and we will reopen this once the camera issue is resolved from customer end. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2407111271 | 2024-08-14 | 2024-08-16 | Cust 542 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning | Chiragbhai and Ajay visited the site for commissioning, refer attached MOM. | Chiragbhai and Ajay visited site for commissioning, refer attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2407118272 | 2024-08-14 | 2024-08-16 | Cust 542 | AHMEDABAD | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | MULTIPARAMETER ANALYSER | Xenon Lamp Power supply for Multiparameter(AWA-XLAMPPS) | Found Xenon Lamp Power supply faulty during commissioning. Refer attached images | A New Power supply was sent to the Branch which was replaced during commissioning as per the attached MOM. | Chiragbhai visited site and replaced the faulty Power supply, Refer attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2407288273 | 2024-08-19 | 2024-08-22 | Cust 486 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Software installation and Configuration. | We took access to your system and successfully completed the installation and configuration of the software. | We took access to your system and successfully completed the installation and configuration of the software. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2407425274 | 2024-08-21 | 2024-08-23 | Cust 512 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Camera not playing on RSPCB Server. | STATIC IP not playing out of network & static IP in customers network for PTZ Cameras ID and PASSWORD is showing incorrect. | STATIC IP not playing out of network &amp; static IP in customers network for PTZ Cameras ID and PASSWORD is showing incorrect. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2407436275 | 2024-08-21 | 2024-08-24 | Cust 470 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on CPCB + Local + RSPCB server | from the hardware side | The issue is from the hardware side | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2407495276 | 2024-08-22 | 2024-08-25 | Cust 544 | AHMEDABAD | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | Aqua4Trans | Turbidity Input card of Aqua4Trans(1) | Instalation/erection guidelines of various instruments | Urvesh and Chiragbhai visited site for installation guidelines of all supplied analysers. Refer attached MOM. | Urvesh and Chiragbhai visited site for installation guidelines of all supplied analysers. Refer attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2407506277 | 2024-08-22 | 2024-08-25 | Cust 477 | AHMEDABAD | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 2nd AMC | Ajay visited the site for AMC, Refer to the attached MOM. | Ajay visited site for AMC, Refer attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFAD2407574278 | 2024-08-23 | 2024-08-23 | Cust 539 | AHMEDABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration done | Calibration done | Calibration done | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFAD2407581279 | 2024-08-23 | 2024-08-23 | Cust 545 | AHMEDABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration Done & working found ok. | Calibration Done & working found ok. | Calibration Done &amp;amp; working found ok. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFAD2407866280 | 2024-08-28 | 2024-09-02 | Cust 470 | AHMEDABAD | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 1st AMC | Urvesh alongwith Jigar visited site for AMC, Refer attached MOM. | Urvesh alongwith Jigar visited site for AMC, Refer attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFAD2407896281 | 2024-08-29 | 2024-08-30 | Cust 520 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Data not updating on LOCAL+CPCB Server. | There was CPU port side issue hardware end now data live on LOCAL+CPCB+RSPCB server. | There was CPU port side issue hardware end now data live on LOCAL+CPCB+RSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2407917282 | 2024-08-29 | 2024-09-02 | Cust 544 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Smartpro-con) | Commissioning | Chiragbhai and Riteshbhai visited site for commissioning of analysers, Refer atatched MOM. | Chiragbhai and Riteshbhai visited site for commissioning of analysers, Refer atatched MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2407917282 | 2024-08-29 | 2024-09-02 | Cust 544 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Smartpro-con) | Commissioning | Chiragbhai and Riteshbhai visited site for commissioning of analysers, Refer atatched MOM. | Chiragbhai and Riteshbhai visited site for commissioning of analysers, Refer atatched MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2407917282 | 2024-08-29 | 2024-09-02 | Cust 544 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Smartpro-con) | Commissioning | Chiragbhai and Riteshbhai visited site for commissioning of analysers, Refer atatched MOM. | Chiragbhai and Riteshbhai visited site for commissioning of analysers, Refer atatched MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2407917282 | 2024-08-29 | 2024-09-02 | Cust 544 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Smartpro-con) | Commissioning | Chiragbhai and Riteshbhai visited site for commissioning of analysers, Refer atatched MOM. | Chiragbhai and Riteshbhai visited site for commissioning of analysers, Refer atatched MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2407917282 | 2024-08-29 | 2024-09-02 | Cust 544 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Smartpro-con) | Commissioning | Chiragbhai and Riteshbhai visited site for commissioning of analysers, Refer atatched MOM. | Chiragbhai and Riteshbhai visited site for commissioning of analysers, Refer atatched MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2407917282 | 2024-08-29 | 2024-09-02 | Cust 544 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Smartpro-con) | Commissioning | Chiragbhai and Riteshbhai visited site for commissioning of analysers, Refer atatched MOM. | Chiragbhai and Riteshbhai visited site for commissioning of analysers, Refer atatched MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2407917282 | 2024-08-29 | 2024-09-02 | Cust 544 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Smartpro-con) | Commissioning | Chiragbhai and Riteshbhai visited site for commissioning of analysers, Refer atatched MOM. | Chiragbhai and Riteshbhai visited site for commissioning of analysers, Refer atatched MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2407917282 | 2024-08-29 | 2024-09-02 | Cust 544 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Smartpro-con) | Commissioning | Chiragbhai and Riteshbhai visited site for commissioning of analysers, Refer atatched MOM. | Chiragbhai and Riteshbhai visited site for commissioning of analysers, Refer atatched MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2407917282 | 2024-08-29 | 2024-09-02 | Cust 544 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Smartpro-con) | Commissioning | Chiragbhai and Riteshbhai visited site for commissioning of analysers, Refer atatched MOM. | Chiragbhai and Riteshbhai visited site for commissioning of analysers, Refer atatched MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2407917282 | 2024-08-29 | 2024-09-02 | Cust 544 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Commissioning | Chiragbhai and Riteshbhai visited site for commissioning of analysers, Refer atatched MOM. | Chiragbhai and Riteshbhai visited site for commissioning of analysers, Refer atatched MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2407917282 | 2024-08-29 | 2024-09-02 | Cust 544 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Commissioning | Chiragbhai and Riteshbhai visited site for commissioning of analysers, Refer atatched MOM. | Chiragbhai and Riteshbhai visited site for commissioning of analysers, Refer atatched MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2407917282 | 2024-08-29 | 2024-09-02 | Cust 544 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Commissioning | Chiragbhai and Riteshbhai visited site for commissioning of analysers, Refer atatched MOM. | Chiragbhai and Riteshbhai visited site for commissioning of analysers, Refer atatched MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2407917282 | 2024-08-29 | 2024-09-02 | Cust 544 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | TOC ANALYSER | FM 101 TOC Analyser(FM 101 TOC) | SWAN make LINETOC Analyser pre-commissioning | Chiragbhai and Riteshbhai visited the site for the commissioning of analysers, Refer attached MOM. | Chiragbhai and Riteshbhai visited site for commissioning of analysers, Refer atatched MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2407918283 | 2024-08-29 | 2024-10-05 | Cust 470 | AHMEDABAD | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | MULTIPARAMETER ANALYSER | Xenon Lamp Power supply for Multiparameter(AWA-XLAMPPS) | No output of power supply. | New Xenon lamp power supply sent to customer, the same was installed and found analyser working ok, Refer attached MOM. | New Xenon lamp power supply sent to customer, the same was installed and found analyser working ok, Refer attached MOM. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2407952284 | 2024-08-30 | 2024-08-30 | Cust 470 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB + RSPCB data upload issue. | we took access of your system and changed Device configuration from RS485 AWA card to Scanner, so now the issue has been resolved and data is live on CPCB + RSPCB server. | we took access of your system and changed Device configuration from RS485 AWA card to Scanner, so now the issue has been resolved and data is live on CPCB + RSPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD2407974285 | 2024-08-30 | 2024-08-30 | Cust 546 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB+LOCAL Server. | There was loose connection from hardware end once connection were checked data started uploading on LOCAL+CPCB+RSPCB server. | There was loose connection from hardware end once connection were checked data started uploading on LOCAL+CPCB+RSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2407992286 | 2024-08-30 | 2024-09-02 | Cust 523 | AHMEDABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Smartpro-con) | Calibration of Conductivity Transmitter | Urvesh and Jigar visited site for calibration of conductivity analysers, Refer attached MOM. | Urvesh and Jigar visited site for calibration of conductivity analysers, Refer attached MOM. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFAD2408345287 | 2024-09-04 | 2024-12-28 | Cust 477 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | pH measurement issue | Urvesh and Alpeshbhai visited site for troubleshooting of pH sensor, refer attached MOM. | It was already resolved and closed. The complaint was reopened by mistake so closing the complaint. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2408368288 | 2024-09-04 | 2024-09-06 | Cust 464 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | New flowmeter has been installed of Endress+Hauser need to do configuration of CPCB Server. | Configuration on New Flowmeter completed and now data is live on CPCB+RSPCB+LOCAL+RTDMS Server. | Configuration on New Flowmeter completed and now data is live on CPCB+RSPCB+LOCAL+RTDMS Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2408398289 | 2024-09-05 | 2024-09-09 | Cust 547 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | TOC ANALYSER | 1950plus, 1 Yr. Spare Parts Kit(200123) | Problem: want to understand the wiring as we have supplied DI to MODBUS JB. | Ashish Sutariya and Manoj Vasagade discussed with Customer Mr. Gaurav over the phone and provided required details for Modbus cabling, so closing the complaint for now. | Ashish Sutariya and Manoj Vasagade discussed with Customer Mr. Gaurav over the phone and provided required details for Modbus cabling, so closing the complaint for now. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2408429290 | 2024-09-05 | 2024-09-09 | Cust 548 | AHMEDABAD | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Visit to provide installation guidelines for 03 Nos. Analyser | Chiragbhai and Ajay visited site for installation guidelines, Refer attached mail. | Chiragbhai and Ajay visited site for installation/commissioning, Refer attached mail. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2408429290 | 2024-09-05 | 2024-09-09 | Cust 548 | AHMEDABAD | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | Aqua4Trans | Aqua4Trans(1) | TDS Transmitters installation guidelines | Chiragbhai and Ajay visited site for installation/commissioning, Refer attached mail. | Chiragbhai and Ajay visited site for installation/commissioning, Refer attached mail. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2408429290 | 2024-09-05 | 2024-09-09 | Cust 548 | AHMEDABAD | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | Aqua4Trans | Aqua4Trans(1) | TDS Transmitters installation guidelines | Chiragbhai and Ajay visited site for installation/commissioning, Refer attached mail. | Chiragbhai and Ajay visited site for installation/commissioning, Refer attached mail. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2408429290 | 2024-09-05 | 2024-09-09 | Cust 548 | AHMEDABAD | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | Aqua4Trans | Aqua4Trans(1) | TDS Transmitters installation guidelines | Chiragbhai and Ajay visited site for installation/commissioning, Refer attached mail. | Chiragbhai and Ajay visited site for installation/commissioning, Refer attached mail. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2408430291 | 2024-09-05 | 2024-09-06 | Cust 524 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Aqua4Trans | Aqua4Trans(1) | Sensor Failure error | Urvesh guided customer over phone and resolved the issue, refer attached mail. | Urvesh guided customer over phone and resolved the issue, refer attached mail. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2408576292 | 2024-09-09 | 2024-09-13 | Cust 464 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Wrong flow data updating on CPCB and RSPCB Server. | We took access to the system and now Flow data is updating on CPCB and RSPCB server. | We took access to the system and now Flow data is updating on CPCB and RSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2408692293 | 2024-09-11 | 2024-09-17 | Cust 540 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Reading issue, suspected issue with sensor. | Urvesh guided over phone for troubleshooting, however observed slope error. Suggested to replace the sensor with new one. Refer attached Mail. | Urvesh guided over phone for troubleshooting, however observed slope error. Suggested to replace the sensor with new one. Refer attached Mail. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2408795294 | 2024-09-13 | 2024-09-13 | Cust 470 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not updating to cpcb server | As of now we are closing this complaint. We will reopen this once we get access to your system. | As of now we are closing this complaint. We will reopen this once we get access to your system. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2408881295 | 2024-09-16 | 2024-09-19 | Cust 527 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on RSPCB server. | commercially not cleared. | commercially not cleared. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2408923296 | 2024-09-16 | 2024-09-17 | Cust 465 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | display not powering up | Chiragbhai guided over phone and suspected issue with Transmitter, suggested to send the same to Pune for further checking, Refer to the attached Mail. | Chiragbhai guided over phone and suspected issue with Transmitter, suggested to send the same to Pune for further checking, Refer to the attached Mail. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2408924297 | 2024-09-16 | 2024-09-17 | Cust 536 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | mA output checking | Urvesh guided over phone and observed no issue with mA output, Refer attached Mail. | Urvesh guided over phone and observed no issue with mA output, Refer attached Mail. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2408924297 | 2024-09-16 | 2024-09-17 | Cust 536 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | mA output checking | Urvesh guided over phone and observed no issue with mA output, Refer attached Mail. | Urvesh guided over phone and observed no issue with mA output, Refer attached Mail. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2409029298 | 2024-09-18 | 2024-09-20 | Cust 464 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera is not playing on CPCB Server. | We took access to your system and camera is now live on CPCB Server. | We took access to your system and camera is now live on CPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2409100299 | 2024-09-20 | 2024-09-23 | Cust 512 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on RSPCB server. | Camera is not playing out of network. So as of now we are closing this complaint. | Camera is not playing out of network. So as of now we are closing this complaint. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2409308300 | 2024-09-24 | 2024-09-27 | Cust 470 | AHMEDABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Analyser calibration after xenon power installation | Alpeshbhai and Riteshbhai visited the site for calibration after the xenon lamp power supply change, Refer attached MOM. | Alpeshbhai and Riteshbhai visited site for calibration, Refer attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFAD2409365301 | 2024-09-25 | 2024-09-30 | Cust 547 | AHMEDABAD | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | TOC ANALYSER | FM 101 TOC Analyser(FM 101 TOC) | Pre-commissioning visit for installation erection checking | Alpeshbhai and Riteshbhai visited site for pre-commissioning checking of Supplied Analysers, Refer attached MOM. | Alpeshbhai and Riteshbhai visited site for pre-commissioning checking of Supplied Analysers, Refer attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2409623302 | 2024-09-30 | 2024-10-03 | Cust 464 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera Streaming issue on the CPCB server. | We checked and haven't found any issue as the Camera streaming is live on the CPDMS ? Forbes Marshall site. | We checked and haven't found any issue as the Camera streaming is live on the CPDMS ? Forbes Marshall site. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD2409654303 | 2024-09-30 | 2024-10-05 | Cust 464 | AHMEDABAD | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC Visit | Riteshbhai and Jigar visited site for AMC, Refer attached MOM. | Riteshbhai and Jigar visited site for AMC, Refer attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFAD2409655304 | 2024-09-30 | 2024-10-05 | Cust 478 | AHMEDABAD | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 1st AMC Visit | Ajay visited the site for AMC, Refer attached MOM. | Ajay visited site for AMC, Refer attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFAD2409768305 | 2024-10-01 | 2024-10-04 | Cust 507 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Reinstallation of software and Camera configuration as PC crashed. | shared software with customer will reopen complaint once installation completed. | shared software with customer will reopen complaint once installation completed. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2409846306 | 2024-10-03 | 2024-10-05 | Cust 549 | AHMEDABAD | Codel Products - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Data upload issue on CPCB server | Following our discussion with the customer, they will provide access to the Blackbox on Monday. Therefore, we are closing this complaint for now. We will reopen it once we have access to the Blackbox. | Following our discussion with the customer, they will provide access to the Blackbox on Monday. Therefore, we are closing this complaint for now. We will reopen it once we have access to the Blackbox. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2409909307 | 2024-10-05 | 2024-10-08 | Cust 523 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Reading over range | Urvesh guided over phone and suspected issue with Transmitter, Suggested to send the same to Pune for further check / repair. Refer to the attached Mail. | Urvesh guided over phone and suspected issue with Transmitter, Suggested to send the same to Pune for further check / repair. Refer to the attached Mail. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2410015308 | 2024-10-08 | 2024-10-11 | Cust 462 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to the CPCB and RSPCB server. | As of now, we are closing this complaint. We will reopen it once we get access to your system. | As of now, we are closing this complaint. We will reopen it once we get access to your system. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2410038309 | 2024-10-08 | 2024-10-11 | Cust 507 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Reinstallation and configuration of software. | As of now, we are closing this complaint. We will reopen it once we get access to your system. | As of now, we are closing this complaint. We will reopen it once we get access to your system. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2410061310 | 2024-10-09 | 2024-10-25 | Cust 550 | AHMEDABAD | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | Aqua4Trans | Aqua4Trans(1) | 24VDC were not coming from Aqua4Trans for Chlorine Sensor. Due to this no response of Chlorine sensor observed. Refer photographs for more clarity. Site is in NEPAL . | Replacement Transmitter sent as per attached docket details. So closing the complaint for now. | Replacement Transmitter sent as per attached docket details. So closing the complaint for now. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2410118311 | 2024-10-09 | 2024-10-14 | Cust 504 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | ORP transmitter(1) | Troubleshooting of ORP Analyser | Alpeshbhai visited the site for troubleshooting, Refer attached Mail. | Alpeshbhai visited the site of troubleshooting, Refer attached Mail. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2410345312 | 2024-10-15 | 2024-10-18 | Cust 462 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to RSPCB server also camera is not live on RSPCB server. | Data is now live on the RSPCB server. We have created the RTSP link for camera which is not playing out of customer's network. Hence as of now we are closing this complaint. We will reopen this once the issue is resolved from customer end. | We have created the RTSP link which is not playing out of customer's network. Hence as of now we are closing this complaint. We will reopen this once the issue is resolved from customer end. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2410378313 | 2024-10-15 | 2024-10-15 | Cust 473 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on RSPCB server. | Now data is live on the RSPCB server. | Now data is live on the RSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2410401314 | 2024-10-16 | 2024-10-18 | Cust 543 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Camera not playing on RSPCB server. | As of now we are closing this complaint we will reopen it once the issue resolved from your end. | As of now we are closing this complaint we will reopen it once the issue resolved from your end. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2410694315 | 2024-10-22 | 2024-10-25 | Cust 543 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Camera not playing on RSPCB server. | As of now we are closing this complaint and we will reopen this once the camera login issue is resolved from your end. | As of now we are closing this complaint and we will reopen this once the camera login issue is resolved from your end. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2410710316 | 2024-10-22 | 2024-10-28 | Cust 534 | AHMEDABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | AQUADO Analyzer | AQUAMAX DOT(1) | Facing the issues regarding the instruments against PO No-1247/16 & PO No-1247/13 also attached the name plate of instruments . Please send the service person at below site address at the earliest for rectification. Unilever India Exports Ltd Kandla Exports Factory Plot No 253-257, Sector IV Kandla, 24- 370230 Contact person : Sandesh Dhule Mob No ? 8591603900. | Chirag Shah guided over phone and suggested to check the sensor with Cap. Awaiting the feedback from Customer so closing the complaint as per attached mail. | Chirag Shah guided over phone and suggested to check the sensor with Cap. Awaiting the feedback from Customer so closing the complaint as per attached mail. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2410756317 | 2024-10-23 | 2024-10-25 | Cust 464 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Data not updating on CPCB and RSPCB server. | Server is accessible and data is live on the CPCB server. | Server is accessible and data is live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2410823318 | 2024-10-24 | 2024-10-25 | Cust 541 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | No data on Eversense cloud server | Chiragbhai and Riteshbhai visited the site for Communication troubleshooting, Refer attached mail. | Chiragbhai and Riteshbhai visited site for Communication troubleshooting, Refer atatched mail. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2410898319 | 2024-10-26 | 2024-10-29 | Cust 536 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | ORP transmitter(1) | 4-20 mA output configuration and troubleshooting | Urvesh visited the site for Troubleshooting and configured the mA output properly. Refer attached MOM. | Urvesh visited site for Troubleshooting, Refer attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2410898319 | 2024-10-26 | 2024-10-29 | Cust 536 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | 4-20 mA output configuration and troubleshooting | Urvesh visited the site for Troubleshooting and configured the mA output properly. Refer attached MOM. | Urvesh visited site for Troubleshooting, Refer attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2410898319 | 2024-10-26 | 2024-10-29 | Cust 536 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | 4-20 mA output configuration and troubleshooting | Urvesh visited the site for Troubleshooting and configured the mA output properly. Refer attached MOM. | Urvesh visited site for Troubleshooting, Refer attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2410898319 | 2024-10-26 | 2024-10-29 | Cust 536 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | 4-20 mA output configuration and troubleshooting | Urvesh visited the site for Troubleshooting and configured the mA output properly. Refer attached MOM. | Urvesh visited site for Troubleshooting, Refer attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2410898319 | 2024-10-26 | 2024-10-29 | Cust 536 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Turbidity Input card of Aqua4Trans(1) | 4-20 mA output configuration and troubleshooting | Urvesh visited the site for Troubleshooting and configured the mA output properly. Refer attached MOM. | Urvesh visited site for Troubleshooting, Refer attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2410898319 | 2024-10-26 | 2024-10-29 | Cust 536 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Turbidity Input card of Aqua4Trans(1) | 4-20 mA output configuration and troubleshooting | Urvesh visited the site for Troubleshooting and configured the mA output properly. Refer attached MOM. | Urvesh visited site for Troubleshooting, Refer attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2411039320 | 2024-10-29 | 2024-10-30 | Cust 543 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | RSPCB camera streaming issue. | We have created an RTSP link and sent to customer. | We have created an RTSP link and sent to customer. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2411229321 | 2024-11-04 | 2024-11-07 | Cust 464 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on RSPCB server. | We took access to the system and now data is live on RSPCB Server. | We took access to the system and now data is live on RSPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2411397322 | 2024-11-07 | 2024-11-09 | Cust 512 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on RSPCB server. | As of now closing the complaint as the static IP provided by customer is not playing out of network. | As of now closing the complaint as the static IP provided by customer is not playing out of network. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2411420323 | 2024-11-07 | 2024-11-11 | Cust 477 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | COD Reading issue | Ajay guided over phone and resolved the issue, Refefr attached mail. | Ajay guided over phone and resolved the issue, Refefr attached mail. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2411481324 | 2024-11-09 | 2024-11-14 | Cust 493 | AHMEDABAD | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC visit | Ajay and Jigar visited site for AMC. Refer attached MOM | Ajay and Jigar visited the site for AMC. Refer attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFAD2411576325 | 2024-11-11 | 2024-11-14 | Cust 512 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on RSPCB server. | As of now we are closing this complaint. We will re-open the complaint once we get Static IP from customer end. | As of now we are closing this complaint. We will re-open the complaint once we get Static IP from customer end. | YES | No Trip | Process | Y |  | Analytical - CHAKAN
AFAD2411578326 | 2024-11-11 | 2024-11-14 | Cust 486 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | online monitoring data is not showing properly ( in percentage) on the RSPCB server. | This is common issue and we have taken up the issue with RSPCB Team and mail communication is shared with customer. | This is common issue and we have taken up the issue with RSPCB Team and mail communication is shared with customer. | YES | No Trip | Process | Y |  | Analytical - CHAKAN
AFAD2411584327 | 2024-11-11 | 2024-11-14 | Cust 484 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data breaking on RSPCB server. | This is common issue and we have taken up the issue with RSPCB Team and mail communication is shared with customer. | This is common issue and we have taken up the issue with RSPCB Team and mail communication is shared with customer. | YES | No Trip | Process | Y |  | Analytical - CHAKAN
AFAD2411607328 | 2024-11-12 | 2024-11-16 | Cust 534 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | AQUAMAX DOT(1) | Commissioning of DO sensor supplied at below site. Site Address : Unilever India Exports Ltd, Kandla Exports Factory, Plot No 253-257, Sector IV, Kandla, 24- 370230 | Urvesh along with Ajay and Jigar visited site for troubleshooting and commissioning, Refer to the attached Mail. | Urvesh along with Ajay and Jigar visited site for troubleshooting and commissioning, Refer to the attached Mail. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2411607328 | 2024-11-12 | 2024-11-16 | Cust 534 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | AQUAMAX DOT(1) | Commissioning of DO sensor supplied at below site. Site Address : Unilever India Exports Ltd, Kandla Exports Factory, Plot No 253-257, Sector IV, Kandla, 24- 370230 | Urvesh along with Ajay and Jigar visited site for troubleshooting and commissioning, Refer to the attached Mail. | Urvesh along with Ajay and Jigar visited site for troubleshooting and commissioning, Refer to the attached Mail. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2411607328 | 2024-11-12 | 2024-11-16 | Cust 534 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | AQUAMAX DOT(1) | Commissioning of DO sensor supplied at below site. Site Address : Unilever India Exports Ltd, Kandla Exports Factory, Plot No 253-257, Sector IV, Kandla, 24- 370230 | Urvesh along with Ajay and Jigar visited site for troubleshooting and commissioning, Refer to the attached Mail. | Urvesh along with Ajay and Jigar visited site for troubleshooting and commissioning, Refer to the attached Mail. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2411694329 | 2024-11-13 | 2024-11-16 | Cust 551 | AHMEDABAD | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 1st AMC Visit | Urvesh visited the site with Jigar for AMC, Refer to the attached MOM. | Urvesh visited the site with Jigar for AMC, Refer attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFAD2411775330 | 2024-11-14 | 2024-11-22 | Cust 551 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software re-installation | now the data is live on the Local Server. | now the data is live on the Local Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2411781331 | 2024-11-14 | 2025-02-10 | Cust 552 | AHMEDABAD | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | The transmitter not powering up. Alpeshbhai checked over the phone and found on current high. Requested Customer to send the transmitter to Pune for further check / repair. | The transmitter was received at Pune and found signal conditioning card was faulty. The replacement was provided which was received by the customer as per the attached mail. | The transmitter was received at Pune and found signal conditioning card was faulty. The replacement was provided which was received by the customer as per the attached mail. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2412013332 | 2024-11-19 | 2024-11-19 | Cust 464 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on CPCB Server. | We took access to the system and now camera is live on CPCB Server. | We took access to the system and now camera is live on CPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2412088333 | 2024-11-21 | 2024-11-23 | Cust 544 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | ORP transmitter(1) | Commissioning of Analysers | Urvesh visited the Site for the commissioning of analysers. Refer attached MOM | Urvesh visited site for Commissioning, Refer attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2412088333 | 2024-11-21 | 2024-11-23 | Cust 544 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | ORP transmitter(1) | Commissioning of Analysers | Urvesh visited the Site for the commissioning of analysers. Refer attached MOM | Urvesh visited site for Commissioning, Refer attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2412088333 | 2024-11-21 | 2024-11-23 | Cust 544 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | ORP transmitter(1) | Commissioning of Analysers | Urvesh visited the Site for the commissioning of analysers. Refer attached MOM | Urvesh visited site for Commissioning, Refer attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2412088333 | 2024-11-21 | 2024-11-23 | Cust 544 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Smartpro-con) | Commissioning of Analysers | Urvesh visited the Site for the commissioning of analysers. Refer attached MOM | Urvesh visited site for Commissioning, Refer attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2412088333 | 2024-11-21 | 2024-11-23 | Cust 544 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Commissioning of Analysers | Urvesh visited the Site for the commissioning of analysers. Refer attached MOM | Urvesh visited site for Commissioning, Refer attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2412088333 | 2024-11-21 | 2024-11-23 | Cust 544 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Commissioning of Analysers | Urvesh visited the Site for the commissioning of analysers. Refer attached MOM | Urvesh visited site for Commissioning, Refer attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2412088333 | 2024-11-21 | 2024-11-23 | Cust 544 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Commissioning of Analysers | Urvesh visited the Site for the commissioning of analysers. Refer attached MOM | Urvesh visited site for Commissioning, Refer attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2412088333 | 2024-11-21 | 2024-11-23 | Cust 544 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Commissioning of Analysers | Urvesh visited the Site for the commissioning of analysers. Refer attached MOM | Urvesh visited site for Commissioning, Refer attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2412088333 | 2024-11-21 | 2024-11-23 | Cust 544 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Commissioning of Analysers | Urvesh visited the Site for the commissioning of analysers. Refer attached MOM | Urvesh visited site for Commissioning, Refer attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2412113334 | 2024-11-21 | 2024-11-22 | Cust 464 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on CPCB server. | We took access to the system and now camera is live on CPCB server. | We took access to the system and now camera is live on CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2412294335 | 2024-11-25 | 2024-11-26 | Cust 464 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on RSPCB Server. | This is the common issue for all industries we have takenup the with RSPCB Team and informed the same to the customer. | This is the common issue for all industries we have takenup the with RSPCB Team and informed the same to the customer. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2412357336 | 2024-11-26 | 2024-11-28 | Cust 553 | AHMEDABAD | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Pre-commissioning guidelines | Urvesh visited the site and provided installation guidelines, Refer attached MOM. | Urvesh visited site for installation guidelines, refer attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2412511337 | 2024-11-28 | 2024-12-02 | Cust 462 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB and RSPCB data upload issue | The COM port is not getting detected as the devices are not connected to the PC. hence as of now we are closing this complaint. we will reopen this once the hardware side issue is resolved. | The COM port is not getting detected as the devices are not connected to the PC. hence as of now we are closing this complaint. we will reopen this once the hardware side issue is resolved. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2412888338 | 2024-12-03 | 2024-12-10 | Cust 544 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | THERMO SERIES | AV88 transmitter(1) | Flouride analyser 2109XP commissioning | Riteshbhai and Chiragbhai visited site for commissooning, However commisioning could not complete due to sample issue. Refer attached MOM | Riteshbhai and Chiragbhai visited site for commissooning, Refer attached MOM | EXTENDED | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2413149339 | 2024-12-09 | 2024-12-11 | Cust 490 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Conductivity Transmitter Aquacon | Aquacon 2000(AQCN 2000 p) | Not Powering up | Chiragbhai visited site for Troubleshooting and found issue with PLC supply, Refer attached Mail. | Chiragbhai visited site for Troubleshooting, Refer attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2413364340 | 2024-12-13 | 2024-12-16 | Cust 477 | AHMEDABAD | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 3rd AMC visit | Urvesh visited site for 3rd AMC, Refer attached MOM. | Urvesh visited site for 3rd AMC, Refer attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFAD2413457341 | 2024-12-16 | 2024-12-19 | Cust 464 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | contact details need to be updated on CPCB server and camera not playing on CPCB Server. | We took access to the system and updated the contact details as per RTDMS and camera streaming window was closed. Camera is live on CPCB server. | We took access to the system and updated the contact details as per RTDMS and camera streaming window was closed. Camera is live on CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2413460342 | 2024-12-16 | 2024-12-23 | Cust 554 | AHMEDABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | AQUADO Analyzer | Aquamax system(1) | Customer facing problem with the DO System. Site Address:- Subh laxmi process Pvt Ltd Railway station, hamirgarh Chittorgarh to Ajmer Road, opposite Hamirgarh, Bhilwara, Rajasthan 311025. Contact: Ashu Tembhare - 8698559524 | VISITED SITE AND CHECK SENSOR AND CONTROLLER WORKING WITH AIR AND SOLLUTION READING COMING OK. | VISITED SITE AND CHECK SENSOR AND CONTROLLER WORKING WITH AIR AND SOLLUTION READING COMING OK. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFAD2413615343 | 2024-12-19 | 2024-12-23 | Cust 555 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning | Urvesh and Suraj visited the site and commissioned the analysers, Refer attached MOM. | Urvesh and Suraj visited the site for commissioning, Refer to the attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2413615343 | 2024-12-19 | 2024-12-23 | Cust 555 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | ORP transmitter(1) | Commissioning | Urvesh and Suraj visited the site and commissioned the analysers, Refer attached MOM. | Urvesh and Suraj visited the site for commissioning, Refer to the attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2413615343 | 2024-12-19 | 2024-12-23 | Cust 555 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Turbidity Meter | Optex turbidity sensor(1) | Commissioning | Urvesh and Suraj visited the site and commissioned the analysers, Refer attached MOM. | Urvesh and Suraj visited the site for commissioning, Refer to the attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2413615343 | 2024-12-19 | 2024-12-23 | Cust 555 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Chlorine Analyzer-MAC 100 | MAC 100(1) | Commissioning | Urvesh and Suraj visited the site and commissioned the analysers, Refer attached MOM. | Urvesh and Suraj visited the site for commissioning, Refer to the attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2413615343 | 2024-12-19 | 2024-12-23 | Cust 555 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Commissioning | Urvesh and Suraj visited the site and commissioned the analysers, Refer attached MOM. | Urvesh and Suraj visited the site for commissioning, Refer to the attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2413615343 | 2024-12-19 | 2024-12-23 | Cust 555 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | SULPHIDE ANALYSER | SULPHIDE 8810 Single Channel ANALYSER(8810 Sulphide) | Commissioning of AWA Make Sulphate Analyser CL1000-822 | Urvesh and Suraj visited the site and commissioned the analysers, Refer attached MOM. | Urvesh and Suraj visited the site for commissioning, Refer to the attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2413615343 | 2024-12-19 | 2024-12-23 | Cust 555 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Conductivity Input card of Aqua4trans(1) | Commissioning of conductivity Analyser | Urvesh and Suraj visited the site and commissioned the analysers, Refer attached MOM. | Urvesh and Suraj visited the site for commissioning, Refer to the attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2413630344 | 2024-12-19 | 2024-12-23 | Cust 462 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to CPCB and RSPCB | As of now we are closing this complaint. We will reopen this once the issue is resolved from the hardware side. | As of now we are closing this complaint. We will reopen this once the issue is resolved from the hardware side. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2413865345 | 2024-12-23 | 2024-12-26 | Cust 462 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | camera streaming issue on RSPCB server. | As of now we are closing this complaint. We will reopen this once the camera issue is resolved from your side. | As of now we are closing this complaint. We will reopen this once the camera issue is resolved from your side. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2413945346 | 2024-12-25 | 2024-12-27 | Cust 556 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | pH Transmitter Aquamon | Ph Transmitter Aquamon 2000(PH2000) | commissioning | Chiragbhai guided customer over phone and commissioned the Analyser, Refer attached Mail. | Chiragbhai guided customer over phone and commissioned the Analyser, Refer attached Mail. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2413970347 | 2024-12-26 | 2024-12-28 | Cust 470 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on CPCB + Local + RSPCB servers. | As discussed during call, we accessed the system and observed that the data is not communicating with the local system. We checked the data using MODSCAN, So this issue is from the hardware side...so we are closing this compliant and we will reopen it once the hardware issue is resolved. | As discussed during call, we accessed the system and observed that the data is not communicating with the local system. We checked the data using MODSCAN, So this issue is from the hardware side.. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2414001348 | 2024-12-26 | 2024-12-28 | Cust 519 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Camera not playing on RSPCB server. | As of now closing this complaint as there is hardware side issue from customer end. | As of now closing this complaint as there is hardware side issue from customer end. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2414091349 | 2024-12-28 | 2024-12-28 | Cust 462 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB and RSPCB data uploading issue | Data is now continuously live on the Local, CPCB and RSPCB server after changing the converter from customer end. | Data is now continuously live on the Local, CPCB and RSPCB server after changing the converter from customer end. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2414214350 | 2024-12-30 | 2025-01-03 | Cust 478 | AHMEDABAD | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 2nd AMC Visit | Urvesh visited site for AMC, Refer attached MOM | Urvesh visited site for AMC, Refer attached MOM | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFAD2414255351 | 2024-12-30 | 2025-01-03 | Cust 557 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Aqua4Trans | Aqua4 trans-A4TW033000(1) | Reading issue of pH & TDS measurement. | Riteshbhai guided customer over the phone and Carried out calibration and commissioned the Analyser. Refer attached mail. | Riteshbhai guided customer over the phone and Carried out calibration and commissioned the Analyser. Refer attached mail. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2414269352 | 2024-12-30 | 2025-01-03 | Cust 464 | AHMEDABAD | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 2nd AMC Visit | Urvesh visited the site for AMC, Refer to the attached MOM | Urvesh visited site for AMC, Refer attached MOM | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFAD2414581353 | 2025-01-03 | 2025-01-06 | Cust 541 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Installation of ISOLATOR for PCB communication | Chiragbhai and Alpeshbhai visited site for installation of RS 485 Isolator, Refer attached Mail. | Chiragbhai and Alpeshbhai visited site for installation of Splitter, Refer attached Mail. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2414776354 | 2025-01-08 | 2025-01-10 | Cust 465 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | cpcb connectivity issue. | we took access of the black box and completed CPCB connectivity for Flow meter, now flow data started uploading on CPCB server. | we took access of the black box and completed CPCB connectivity for Flow meter, now flow data started uploading on CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD2414904355 | 2025-01-10 | 2025-01-10 | Cust 465 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | cpcb connectivity needs to be done. | we took access of the black box and completed CPCB connectivity for Flow meter, now flow data started uploading on CPCB server. | we took access of the black box and completed CPCB connectivity for Flow meter, now flow data started uploading on CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD2415011356 | 2025-01-13 | 2025-01-14 | Cust 465 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera connectivity needs to be done. | we took access to your system and completed Camera configuration, now Camera is streaming live on CPCB server. | we took access to your system and completed Camera configuration, now Camera is streaming live on CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD2415015357 | 2025-01-13 | 2025-01-16 | Cust 548 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning of Analyser. The site is at Sysmex India Pvt Ltd, Sanand | Chiragbhai and Ajay visited the site for commissioning, Refer to the attached MOM. | Chiragbhai and Ajay visited site for commissioning, Refer attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2415015357 | 2025-01-13 | 2025-01-16 | Cust 548 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning of Analyser. The site is at Sysmex India Pvt Ltd, Sanand | Chiragbhai and Ajay visited the site for commissioning, Refer to the attached MOM. | Chiragbhai and Ajay visited site for commissioning, Refer attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2415015357 | 2025-01-13 | 2025-01-16 | Cust 548 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning of Analyser. The site is at Sysmex India Pvt Ltd, Sanand | Chiragbhai and Ajay visited the site for commissioning, and observed issue with Motherboard. Refer to the attached MOM. | Chiragbhai and Ajay visited site for commissioning, Refer attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2415042358 | 2025-01-14 | 2025-01-16 | Cust 465 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | GPCB connectivity needs yo be done. | we took access of the Black Box (IP : 192.168.1.111), and completed GPCB connectivity now data is live on the GPCB server. | we took access of the Black Box (IP : 192.168.1.111), and completed GPCB connectivity now data is live on the GPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD2415066359 | 2025-01-14 | 2025-01-27 | Cust 471 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data availability issue | Now the data availabilty is increase on the RSPCB server. | Now the data availabilty is increase on the RSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2415071360 | 2025-01-14 | 2025-01-17 | Cust 519 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | eversense Software re-installation and re-configuration | We have completed CPCB and RSPCB configuration and data is now live on the same servers. | We have completed CPCB and RSPCB configuration and data is now live on the same servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2415187361 | 2025-01-16 | 2025-01-19 | Cust 458 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Not powering up | Alpeshbhai visited the site for troubleshooting, observed analyser working ok. Refer attached mail. | Alpeshbhai visited site for troubleshooting, Refer attached mail. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2415187361 | 2025-01-16 | 2025-01-19 | Cust 458 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Aqua4Trans | Aqua4Trans(1) | Turbidity Analyser not powering up | Alpeshbhai visited the site for troubleshooting, observed analyser power pcb burnt due to power spike. Suggested to send for repair. Refer attached mail. | Alpeshbhai visited site for troubleshooting, Refer attached mail. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2415187361 | 2025-01-16 | 2025-01-19 | Cust 458 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | DO Transmitters | optex series(optex series) | DPI 201 not powering up | Alpeshbhai visited the site for troubleshooting, observed analyser 24V SMPS burnt due to power spike. Suggested to send for repair. Refer attached mail. | Alpeshbhai visited site for troubleshooting, Refer attached mail. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2415194362 | 2025-01-16 | 2025-01-17 | Cust 537 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Turbidity Meter | Optex(FMSC-P3) | The display not showing. | Urvesh guided over phone and observed issue with sensor and Analyser. Suggested to replace the same with new one being an obsolete model. Refer to the attached Mail. | Urvesh guided over phone and observed issue with sensor and Analyser. Suggested to replace the same with new one being an obsolete model. Refer to the attached Mail. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2415221363 | 2025-01-17 | 2025-01-19 | Cust 544 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | TOC ANALYSER | TOC 1950 Single StreamANALYSER ANATEL(1950 TOC) | SWAN TOC analyser Calibration and commissioning | Chiragbhai and Riteshbhai visited the site for calibration of Analyser after receipt of standard solution, refer atached MOM | Chiragbhai and Riteshbhai visited the site for calibration of Analyser, refer atached MOM | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2415229364 | 2025-01-17 | 2025-01-20 | Cust 493 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | DATCON | DATCON(DATCON) | CPCB and GPCB data uploaded issue. | As of now we are closing this complaint and will reopen it once we get DATCon access. | As of now we are closing this complaint and will reopen it once we get DATCon access. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD2415549365 | 2025-01-24 | 2025-01-27 | Cust 464 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | camera configuration needs to be done. | We took access and changed the configuration from .ogg to web.m and now the Camera streming is running. | We took access and changed the configuration from .ogg to web.m and now the Camera streming is running. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD2415552366 | 2025-01-24 | 2025-01-31 | Cust 474 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Re-installation of software and configuration needs to be done. | NEW PC IN RE INSTALLED AND CONFIGURATION SYTEM AFTER WORKING OK | NEW PC IN RE INSTALLED AND CONFIGURATION SYTEM AFTER WORKING OK | YES | No Trip | Process | Y |  | Analytical - CHAKAN
AFAD2415561367 | 2025-01-24 | 2025-01-27 | Cust 548 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | CX2000 commissioning | Chiragbhai and Ajay visited the site for the replacement of Motherboard and all other analysers connectivity at FMBB, Refer to the attached MOM. | Chiragbhai and Ajay visited site for commissioning activities, Refer attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2415585368 | 2025-01-25 | 2025-01-27 | Cust 548 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Blackbox Commissioning for Sysmex Pvt Ltd site. | we completed the configuration. Data was also coming to the BB, | we completed the configuration. Data was also coming to the BB, | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD2415732369 | 2025-01-27 | 2025-01-31 | Cust 548 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DIGITAL CONEECTIVITY | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Connectivity of CX2000 | Chiragbhai and Amol Sathe visited the site for Connectivity and upgradation, Refer to the attached MOM. | Chiragbhai and Amol Sathe visited the site for Connectivity and upgradation, Refer to the attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2415788370 | 2025-01-28 | 2025-02-03 | Cust 494 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | Basic checking of old supplied analysers as per customer's request. | Chiragbhai and Amol Sathe visited site for healthiness checking of analysers, assessed 15 AWA analysers during the visit, refer attached mail. | Chiragbhai and Amol Sathe visited site for healthiness checking of analysers, refer attached mail. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2415801371 | 2025-01-28 | 2025-02-04 | Cust 558 | AHMEDABAD | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | Model: CX1000-6022, Refer following issue detail and attached videos/photos. Issue:- 1) Touch Screen not working at all. 2) Carried out hardware reset procedure and observed RS 232 Card slot was showing "WRONG". Ignored it carried out configuration but observed display stuck at Cleaning Cycle. Even no countdown of Cleaning pump On cycle happening. 3) Removed RS 232 card and reset - reconfiguraton procedure again. Observed the same issue as mentioned above. 4) Removed CMOS battery and plug it back, did resetting- reconfiguration again but same problem observed. | Ajay and debanjan visited site and replaced faulty display and PCX1 card with spare received from Pune. After replacement, found analyser working ok. Refer attached mail. | Ajay and debanjan visited site and replaced faulty display with PCX1 card. Refer attached mail. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2415893372 | 2025-01-29 | 2025-02-01 | Cust 464 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | RSPCB data uploaded issue. | we identified a Java-related problem, which has now been resolved and the data is now live on both sites (Local and RSPCB | we identified a Java-related problem, which has now been resolved and the data is now live on both sites (Local and RSPCB | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD2415942373 | 2025-01-29 | 2025-01-31 | Cust 501 | AHMEDABAD | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 1st AMC visit | Ajay visited site for AMC of Analysers, Refer attached MOM. | Ajay visited site for AMC of Analysers, Refer attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFAD2415942373 | 2025-01-29 | 2025-01-31 | Cust 501 | AHMEDABAD | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 1st AMC visit | Ajay visited site for AMC of Analysers, Refer attached MOM. | Ajay visited site for AMC of Analysers, Refer attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFAD2415983374 | 2025-01-30 | 2025-01-31 | Cust 553 | AHMEDABAD | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Pre-commissioning visit | Urvesh and Riteshbhai visited site for installation guidelines, Refer attached MOM. | Urvesh and Riteshbhai visited site for installation guidelines, Refer attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2416019375 | 2025-01-30 | 2025-01-31 | Cust 458 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | pH Transmitter Aquamon | Ph Transmitter Aquamon 2000(PH-Trans Aqu-2000) | Calibration and Commissioning | Alpeshbhai and Chiragbhai guided over the phone for commissioning and Calibration, Refer to the attached mail. | Alpeshbhai and Chiragbhai guided over the phone for commissioning and Calibration, Refer to the attached mail. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2416019375 | 2025-01-30 | 2025-01-31 | Cust 458 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | pH Transmitter Aquamon | Ph Transmitter Aquamon 2000(PH-Trans Aqu-2000) | Calibration and Commissioning | Alpeshbhai and Chiragbhai guided over the phone for commissioning and Calibration, Refer to the attached mail. | Alpeshbhai and Chiragbhai guided over the phone for commissioning and Calibration, Refer to the attached mail. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2416019375 | 2025-01-30 | 2025-01-31 | Cust 458 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Transmitter Aquacon | Conductivity Transmitter Aquamon 2000(COND-Trans Aqu-2000) | Commissioning | Alpeshbhai and Chiragbhai guided over the phone for commissioning and Calibration, Refer to the attached mail. | Alpeshbhai and Chiragbhai guided over the phone for commissioning and Calibration, Refer to the attached mail. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2416219376 | 2025-02-01 | 2025-02-07 | Cust 464 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera streaming issue. | There was Java error issue which has been resolved and now camera is live on the CPCB server. | There was Java error issue which has been resolved and now camera is live on the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD2416391377 | 2025-02-05 | 2025-02-07 | Cust 464 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | ETP data not updating on CPCB server. | We took access to the system and now data is updating on CPCB server. | We took access to the system and now data is updating on CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2416411378 | 2025-02-06 | 2025-02-08 | Cust 497 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data uploaded issue. | Updated the BB with the new patch. | Updated the BB with the new patch. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD2416547379 | 2025-02-10 | 2025-02-13 | Cust 559 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on CPCB server. | We took access to the system and observed the streaming window was closed. We have started the streaming window and now camera is live on CPCB server. | We took access to the system and observed the streaming window was closed. We have started the streaming window and now camera is live on CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2416585380 | 2025-02-10 | 2025-02-12 | Cust 551 | AHMEDABAD | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 2nd AMC Visit | Chiragbhai visited the site for 2nd AMC of Multiparameter Analyser, Refer attached MOM. | Chiragbhai visited the site for AMC, Refer attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFAD2416601381 | 2025-02-11 | 2025-02-14 | Cust 553 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning of CX2000-3922 Analyser | Urvesh visited the site for commissioning, however, found the Xenon lamp power supply faulty. Need to provide a replacement for further commissioning. Refer attached MOM. | Urvesh visited the site for commissioning, Refer attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2416604382 | 2025-02-11 | 2025-02-18 | Cust 553 | AHMEDABAD | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | MULTIPARAMETER ANALYSER | Xenon Lamp Power supply for Multiparameter(AWA-XLAMPPS) | Xenon lamp Power supply found not working during commissioning. No output voltage and Light indication observed in Power supply. 12 V input found ok. | Replacement spares were sent as per the attached mail, closing the complaint for now. | Replacement spares were sent as per the attached mail, closing the complaint for now. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2416611383 | 2025-02-11 | 2025-02-14 | Cust 462 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Service Visit | Ajay visited the site for an Analyser service. Refer attached MOM. | Ajay visited the site for Analyser service. Refer attached MOM. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFAD2416614384 | 2025-02-11 | 2025-02-18 | Cust 553 | AHMEDABAD | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Blackbox ping issue of both ethernet ports which was supplied with CX2000 | Replacement spares were sent as per the attached mail, closing the complaint for now. | Replacement spares were sent as per the attached mail, closing the complaint for now. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2416670385 | 2025-02-12 | 2025-02-14 | Cust 520 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Reinstallation and configuration of software. | We have completed the configuration of software and now data is live on LOCAL+CPCB+RSPCB server. | We have completed the configuration of software and now data is live on LOCAL+CPCB+RSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2416671386 | 2025-02-12 | 2025-02-14 | Cust 464 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | RSPCB camera streaming issue. | We have created the rtsp link which is playing out of customers network and shared same with the customer. | We have created the rtsp link which is playing out of customers network and shared same with the customer. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2416768387 | 2025-02-14 | 2025-02-18 | Cust 501 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Reading issue | Ajay visited the site for troubleshooting and rectified the issue, Refer attached MOM. | Ajay visited the site for troubleshooting, Refer attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFAD2416768387 | 2025-02-14 | 2025-02-18 | Cust 501 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Reading issue | Ajay visited the site for troubleshooting and rectified the issue, Refer attached MOM. | Ajay visited the site for troubleshooting, Refer attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFAD2416919388 | 2025-02-18 | 2025-05-09 | Cust 497 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data uploaded issue and BB hanging issue. | After Troubleshooting online by Pune team, suspected issue with FMBB. So suggested customer to send FMBB to Pune for further checking / repair. Refer attached Mail. | After Troubleshooting online by Pune team, suspected issue with FMBB. So suggested customer to send FMBB to Pune for further checking / repair. Refer attached Mail. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD2416927389 | 2025-02-18 | 2025-02-20 | Cust 560 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | reading issue | Guided over phone and resolved the issue by proper wiring, Refer attached mail. | Guided over phone and resolved the issue by proper wiring, Refer attached mail. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2417010390 | 2025-02-20 | 2025-02-23 | Cust 477 | AHMEDABAD | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 4th AMC visit | Urvesh visited the site for 4th AMC work, Refer to the attached MOM. | Urvesh visited the site for 4th AMC work, Refer to the attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFAD2417010390 | 2025-02-20 | 2025-02-23 | Cust 477 | AHMEDABAD | Analytical - CHAKAN | CLOSED | AMC | Turbidity Meter | Optex(FMSC-P3) | 4th AMC of SS Analyser FMSC-T4 | Urvesh visited the site for 4th AMC work, Refer to the attached MOM. | Urvesh visited the site for 4th AMC work, Refer to the attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFAD2417018391 | 2025-02-20 | 2025-02-23 | Cust 553 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning of Analyser | Chiragbhai visited the site for commissioning, replaced the faulty power supply and FMBB and commissioned the Analyser. Refer to the attached MOM. | Chiragbhai visited the site for commissioning, Refer to the attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2417459392 | 2025-02-28 | 2025-03-03 | Cust 561 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | Commissioning of resistivity meter | Urvesh visited the site for checking the flow through issue, suggested to send the flow through chamber to Pune for further change/cleaning, Refer attached mail. | Urvesh visited the site for checking the flow through issue, Refer attached mail. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2417459392 | 2025-02-28 | 2025-03-03 | Cust 561 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | Commissioning of resistivity meter | Urvesh visited site and suggested for reinstallation of flow through chamber. Refer attached mail. | Urvesh visited the site for checking the flow through issue, Refer attached mail. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2417470393 | 2025-02-28 | 2025-03-05 | Cust 501 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | TSS Reading Issue | Urvesh and Pune team guided for troubleshooting and resolved the issue as the sample line choking was observed which was creating the issue, refer attached mail. | Urvesh and Pune team guided for troubleshooting and resolved the issue as the sample line choking was observed which was creating the issue, refer attached mail. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2417687394 | 2025-03-04 | 2025-03-07 | Cust 555 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning of CX2000-3912 | Urvesh and Chirag Shah visited the site for commissioning, refer to the attached MOM. | Urvesh and Chirag Shah visited the site for commissioning, Refer attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2417687394 | 2025-03-04 | 2025-03-07 | Cust 555 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Turbidity Meter | Optex(FMSC-P3) | Calibration and commissioning | Urvesh and Chirag Shah visited the site for commissioning, Refer to the attached MOM. | Urvesh and Chirag Shah visited the site for commissioning, Refer attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2417687394 | 2025-03-04 | 2025-03-07 | Cust 555 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Turbidity Meter | Optex(FMSC-P3) | Calibration and commissioning | Urvesh and Chirag Shah visited the site for commissioning, Refer to the attached MOM. | Urvesh and Chirag Shah visited the site for commissioning, Refer attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2417687394 | 2025-03-04 | 2025-03-07 | Cust 555 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | SULPHIDE ANALYSER | SULPHIDE 8810 Single Channel ANALYSER(8810 Sulphide) | Commissioning and calibration of Sulphate Analyser | Urvesh and Chirag Shah visited the site for commissioning, Refer to the attached MOM. | Urvesh and Chirag Shah visited the site for commissioning, Refer attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2417687394 | 2025-03-04 | 2025-03-07 | Cust 555 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | Commissioning of TDS meter. | Urvesh and Chirag Shah visited the site for commissioning, Refer attached MOM. | Urvesh and Chirag Shah visited the site for commissioning, Refer attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2417687394 | 2025-03-04 | 2025-03-07 | Cust 555 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | Commissioning of TDS meter. | Urvesh and Chirag Shah visited the site for commissioning, Refer attached MOM. | Urvesh and Chirag Shah visited the site for commissioning, Refer attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2417780395 | 2025-03-05 | 2025-03-06 | Cust 562 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Conductivity Transmitter Aquacon | Aquacon 2000(AQCN 2000 p) | Reading fluctuation reported | Urvesh checked the transmitter at office and observed no issue in the transmitter, Refer attached mail. | Urvesh checked the transmitter at office and observed no issue in the transmitter, Refer attached mail. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2418041396 | 2025-03-10 | 2025-03-22 | Cust 501 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 0 reading of TSS reported, Suggested for proper zero calibration. | Urvesh guided over the phone and carried out proper zero calibration of TSS, later observed the reading for three days and found non-zero. Refer attached mail. | Urvesh guided over the phone and carried out proper zero calibration of TSS, later observed the reading for three days and found non-zero. Refer attached mail. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2418127397 | 2025-03-12 | 2025-03-14 | Cust 477 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Localhost credentails are not accessible. | We shared the login ID and Password with the customer and now localhost is accessible. | We shared the login ID and Password with the customer and now localhost is accessible. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2418182398 | 2025-03-13 | 2025-03-16 | Cust 563 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Aqua4Trans | Power supply card of Aqua4Trans(1) | Power supply short circuit | Chiragbhai guided over phone and observed issue with transmitter. Suggested to send the faulty transmitter to pune. Refer attached mail. | Chiragbhai guided over phone and observed issue with transmitter. Suggested to send the faulty transmitter to pune. Refer attached mail. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2418376399 | 2025-03-18 | 2025-03-21 | Cust 555 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Turbidity Meter | Optex(FMSC-P3) | mA output not working | Urvesh discussed with the customer regarding the issue and suggested sending the transmitter to Pune for further checking/repair. Refer to the attached Mail. | Urvesh discussed with the customer regarding the issue and suggested sending the transmitter to Pune for further checking/repair. Refer to the attached Mail. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2418441400 | 2025-03-20 | 2025-03-22 | Cust 478 | AHMEDABAD | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 3rd AMC visit | Chiragbhai visited the site for Maintenance, Refer attached MOM. | Chiragbhai visited the site for Maintenance, Refer attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFAD2418473401 | 2025-03-21 | 2025-03-22 | Cust 557 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Aqua4Trans | Aqua4Trans(1) | TDS meter reading issue | Urvesh visited the site for troubleshooting, Calibrated the Analyser with a standard solution and found ok. Refer to the attached Mail. | Urvesh visited the site for troubleshooting, Refer to the attached MOM. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2418473401 | 2025-03-21 | 2025-03-22 | Cust 557 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Aqua4Trans | Aqua4Trans(1) | TDS meter reading issue | Urvesh visited the site for troubleshooting, Observed an issue with the TDS sensor Refer to the attached Mail. | Urvesh visited the site for troubleshooting, Refer to the attached MOM. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2418476402 | 2025-03-21 | 2025-03-25 | Cust 464 | AHMEDABAD | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 3rd AMC visit | Chiragbhai visited the site for 3rd AMC visit, Refer attached MOM. | Chiragbhai visited the site for 3rd AMC visit, Refer attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFAD2418633403 | 2025-03-25 | 2025-03-28 | Cust 498 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Spare installation and servicing | Chiragbhai visited the site for Analyser servicing, Refer attached MOM. | Chiragbhai visited the site for spare installation and chargeable service, Refer attached MOM. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFAD2418817404 | 2025-03-28 | 2025-03-29 | Cust 464 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | LOCALHOST not accessible | We took access to the system ad installed java and started camera streaming window after which localhost + camera is data live. | We took access to the system ad installed java and started camera streaming window after which localhost + camera is data live. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2418870405 | 2025-03-29 | 2025-03-29 | Cust 519 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Configuration of Eversense software. | We took access to the system and completed the configuration of software and now data is live on server. | We took access to the system and completed the configuration of software and now data is live on server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2418919406 | 2025-03-30 | 2025-03-30 | Cust 545 | AHMEDABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | calibration job | calibration done | calibration done | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFAD2418985407 | 2025-03-31 | 2025-04-03 | Cust 464 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on RSPCB server. | We have created the rtsp link and shard with the customer for RSPCB camera connectivity. | We have created the rtsp link and shard with the customer for RSPCB camera connectivity. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2418989408 | 2025-03-31 | 2025-04-07 | Cust 471 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to RSPCB server. | Now the issue has been resolved and data is live on the RSPCB server. | Now the issue has been resolved and data is live on the RSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2500006409 | 2025-04-01 | 2025-04-04 | Cust 519 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on RSPCB server. | We took access to your system and done the necessary changes in software and now data availability has increased. | We took access to your system and done the necessary changes in software and now data availability has increased. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2500134410 | 2025-04-04 | 2025-04-05 | Cust 564 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Smartpro-con) | Zero Reading issue | Urvesh guided over phone and suspected issue with Transmitter, Suggested to send the same to Pune for further check/repair. Refer attached Mail. | Urvesh guided over phone and suspected issue with Transmitter, Suggested to send the same to Pune for further check/repair. Refer attached Mail. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2500170411 | 2025-04-06 | 2025-04-07 | Cust 523 | AHMEDABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | Calibration | Urvesh visited the site for Calibration of Analysers, Refer attached MOM. | Urvesh visited the site for Calibration of Analysers, Refer attached MOM. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFAD2500203412 | 2025-04-07 | 2025-04-08 | Cust 565 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Transmitter Aquacon | Aquacon 2000(AQCN 2000 p) | Calibration. | Alpeshbhai guided over phone for calibration and commissioned the Analyser, Refer attached mail. | Alpeshbhai guided over phone for calibration and commissioned the Analyser, Refer attached mail. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2500203412 | 2025-04-07 | 2025-04-08 | Cust 565 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Transmitter Aquacon | Aquacon 2000(AQCN 2000 p) | Calibration. | Alpeshbhai guided over phone for calibration and commissioned the Analyser, Refer attached mail. | Alpeshbhai guided over phone for calibration and commissioned the Analyser, Refer attached mail. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2500207413 | 2025-04-07 | 2025-04-08 | Cust 566 | AHMEDABAD | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Erection and Installation guideline visit | Chiragbhai visited the site for installation/erection guidelines, Refer attached Mail. | Chiragbhai visited the site for installation/erection guidelines, Refer attached Mail. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2500328414 | 2025-04-09 | 2025-04-11 | Cust 470 | AHMEDABAD | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 2nd AMC visit | Jigar visited the site for 2nd AMC of Multiparameter Analyser, Refer attached MOM. | Jigar visited the site for 2nd AMC of Multiparameter Analyser, Refer attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFAD2500392415 | 2025-04-10 | 2025-04-11 | Cust 567 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | AQUAMAX DOT(1) | Troubleshooting and Commissioning | Urvesh guided customer over phone, Corrected the wiring and resolve the issue. Found Analyser working ok. Refer attached mail. | Urvesh guided customer over phone, Corrected the wiring and resolve the issue. Found Analyser working ok. Refer attached mail. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2500394416 | 2025-04-10 | 2025-04-24 | Cust 548 | AHMEDABAD | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | TSS reading very high, Low PD1 and PD2 reading. Observed Xenon lamp power supply faulty. | Based on the remote checking, observed issue with xenon power supply. Ajay and Jigar installed new Power supply received and resolved the issue. Refer attached MOM. | Ajay and Jigar visited the site for troubleshooting, Refer attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2500397417 | 2025-04-10 | 2025-04-11 | Cust 553 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB connectivity needs to be done. | we have completed the CPCB connectivity, and data is live on the server. | we have completed the CPCB connectivity, and data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD2500401418 | 2025-04-10 | 2025-04-11 | Cust 458 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | AQUAMAX DOT(1) | Zero Reading of Analyser | Urvesh guided over phone and suspected issue with sensor/cleaning. Suggested to remove the sensor from tank. It will take time to empty the tank and remove the sensor. so closing the complaint for now. Refer atatched mail. | Urvesh guided over phone and suspected issue with sensor/cleaning. Suggested to remove the sensor from tank. It will take time to empty the tank and remove the sensor. Refer atatched mail. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2500446419 | 2025-04-11 | 2025-04-12 | Cust 484 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on LOCAL+CPCB+RSPCB server. | Services were stopped after restarting data is live now on LOCAL+CPCB+RSPCB server. | Services were stopped after restarting data is live now on LOCAL+CPCB+RSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2500484420 | 2025-04-13 | 2025-04-16 | Cust 568 | AHMEDABAD | Steam System Division | CLOSED | QUALITY PROBLEM | CRS | CONDENSATE RECOVERY METER - MODEL CRM-485(CRM-485R) | CRM IN TEMP READING NOT COMING | Reference C3G RJ OA Number 5526200014 for M/s. AMRUTVA FINE FOODS LLP The item CRM 485R = Qty01No | Reference C3G RJ OA Number 5526200014 for M/s. AMRUTVA FINE FOODS LLP The item CRM 485R = Qty01No | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFAD2500558421 | 2025-04-15 | 2025-04-17 | Cust 470 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to Local, CPCB and RSPCB server. Camera streaming not live on RSPCB. | Data is live on the CPCB and RSPCB servers. We have created and sent an RTSP link which is playing out of customers network. | Data is live on the CPCB and RSPCB servers. We have created and sent an RTSP link which is playing out of customers network. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2500568422 | 2025-04-15 | 2025-04-16 | Cust 569 | AHMEDABAD | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(Ph-Trans-Smartpro) | Installation guidelines | Chiragbhai has guided customer and provided installation guidelines for PH sensor. Also suggested to install with DIP type system in case of installation in deaeration tank. Refer attached Mail. | Chiragbhai has guided customer and provided installation guidelines for PH sensor. Also suggested to install with DIP type system in case of installation in deaeration tank. Refer attached Mail. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2500809423 | 2025-04-22 | 2025-04-24 | Cust 478 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data availability is very less. | We took access to the system and observed data availability has increased on RSPCB server. | We took access to the system and observed data availability has increased on RSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2500917424 | 2025-04-24 | 2025-04-26 | Cust 570 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Installation | Chiragbhai guided over phone for installation, refer attached mail. | Chiragbhai guided over phone for installation, refer attached mail. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2501162425 | 2025-04-29 | 2025-05-02 | Cust 471 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | software re-installation and Local + CPCB + RSPCB connectivity | We have re-installed the software with an upgraded version and completed CPCB, Local and RSPCB connectivity. Now the data is live on the same servers. | We have re-installed the software with an upgraded version and completed CPCB, Local and RSPCB connectivity. Now the data is live on the same servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2501165426 | 2025-04-29 | 2025-05-02 | Cust 523 | AHMEDABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | ORP transmitter(1) | Calibration | Chiragbhai visited the site for Calibration, Refer attached MOM. | Chiragbhai visited the site for Calibration, Refer attached MOM. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFAD2501165426 | 2025-04-29 | 2025-05-02 | Cust 523 | AHMEDABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | ORP transmitter(1) | Calibration | Chiragbhai visited the site for Calibration, Refer attached MOM. | Chiragbhai visited the site for Calibration, Refer attached MOM. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFAD2501165426 | 2025-04-29 | 2025-05-02 | Cust 523 | AHMEDABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Calibration | Chiragbhai visited the site for Calibration, Refer attached MOM. | Chiragbhai visited the site for Calibration, Refer attached MOM. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFAD2501165426 | 2025-04-29 | 2025-05-02 | Cust 523 | AHMEDABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Calibration | Chiragbhai visited the site for Calibration, Refer attached MOM. | Chiragbhai visited the site for Calibration, Refer attached MOM. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFAD2501165426 | 2025-04-29 | 2025-05-02 | Cust 523 | AHMEDABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Calibration | Chiragbhai visited the site for Calibration, Refer attached MOM. | Chiragbhai visited the site for Calibration, Refer attached MOM. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFAD2501165426 | 2025-04-29 | 2025-05-02 | Cust 523 | AHMEDABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Calibration | Chiragbhai visited the site for Calibration, Refer attached MOM. | Chiragbhai visited the site for Calibration, Refer attached MOM. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFAD2501165426 | 2025-04-29 | 2025-05-02 | Cust 523 | AHMEDABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Calibration | Chiragbhai visited the site for Calibration, Refer attached MOM. | Chiragbhai visited the site for Calibration, Refer attached MOM. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFAD2501165426 | 2025-04-29 | 2025-05-02 | Cust 523 | AHMEDABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Calibration | Chiragbhai visited the site for Calibration, Refer attached MOM. | Chiragbhai visited the site for Calibration, Refer attached MOM. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFAD2501165426 | 2025-04-29 | 2025-05-02 | Cust 523 | AHMEDABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Calibration | Chiragbhai visited the site for Calibration, Refer attached MOM. | Chiragbhai visited the site for Calibration, Refer attached MOM. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFAD2501165426 | 2025-04-29 | 2025-05-02 | Cust 523 | AHMEDABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Calibration | Chiragbhai visited the site for Calibration, Refer attached MOM. | Chiragbhai visited the site for Calibration, Refer attached MOM. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFAD2501205427 | 2025-04-29 | 2025-05-16 | Cust 548 | AHMEDABAD | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Xenon lamp power supply issue | Chiragbhai visited the site for checking the issue. Observed issue with Xenon Power supply. Carried back faulty power supply and sent to Pune for further checking/replacement. On 14th May, 2025, Urvesh and Amol Sathe visited the site and replaced the Faulty xenon lamp power supply, Main board and HMI. Calibrated the Analyser and found working ok. Refer attached mail. | Urvesh and Amol Sathe visited the site and replaced the Faulty xenon lamp power supply, Main board and HMI. Refer attached mail. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2501491428 | 2025-05-02 | 2025-05-05 | Cust 571 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | pH reading issue | Urvesh visited the site for Troubleshooting and resolved issue by Transmitter calibration, Refer attached mail. | Urvesh visited the site for Troubleshooting, Refer attached mail. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2501585429 | 2025-05-05 | 2025-05-07 | Cust 572 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Installation/commissioning of analyser | Chiragbhai provided installation guidelines and suggested to confirm once the Analyser is ready for commissioning. Refer attached mail. Meanwhile closing the complaint for now. | Chiragbhai provided installation guidelines and suggested to confirm once the Analyser is ready for commissioning. Refer attached mail. Meanwhile closing the complaint for now. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2501767430 | 2025-05-09 | 2025-05-12 | Cust 470 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB+RSPCB server. | As of now we are closing this complaint, we will reopen the complaint once we get access to your system. | As of now we are closing this complaint, we will reopen the complaint once we get access to your system. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2501769431 | 2025-05-09 | 2025-05-12 | Cust 462 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB+RSPCB server. | As of now we are closing this complaint, we will reopen the complaint once we get access to your system. | As of now we are closing this complaint, we will reopen the complaint once we get access to your system. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2501785432 | 2025-05-09 | 2025-05-10 | Cust 573 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4Trans(1) | MA output wiring and troubleshooting | Urvesh guided over phone and configured the mA outputs, Refer attached Mail. | Urvesh guided over phone and configured the mA outputs, Refer attached Mail. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2501808433 | 2025-05-10 | 2025-05-12 | Cust 561 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Smartpro-con) | Reading issue | Urvesh visited the site and observed issues with Transmitters, suggested to send the transmitters to Pune for further repair, Refer attached mail. | Urvesh visited the site and observed issues with Transmitters, Refer attached mail. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2501808433 | 2025-05-10 | 2025-05-12 | Cust 561 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Smartpro-con) | Reading issue | Urvesh visited the site and observed issues with Transmitters, suggested to send the transmitters to Pune for further repair, Refer attached mail. | Urvesh visited the site and observed issues with Transmitters, Refer attached mail. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2501854434 | 2025-05-12 | 2025-05-15 | Cust 473 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on RSPCB server. | This is a common issue for RSPCB server. | This is a common issue for RSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2501871435 | 2025-05-12 | 2025-05-15 | Cust 551 | AHMEDABAD | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 3rd AMC Visit | Chiragbhai visited the site for Maintenance of Analyser, Refer attached MOM. | Chiragbhai visited the site for Maintenance of Analyser, Refer attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFAD2502000436 | 2025-05-15 |  | Cust 484 | AHMEDABAD | Analytical - CHAKAN | ASSIGNED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not reflecting to the RSPCB server. |  |  | YES | No Trip | Process | Y |  | Analytical - CHAKAN
AFAD2502016437 | 2025-05-15 | 2025-05-20 | Cust 497 | AHMEDABAD | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | Aqua4Trans | Aqua4Trans(1) | Installation Guideline | Chiragbhai visited the site for installation guidelines. Refer attached mail. | Chiragbhai visited the site for installation guidelines, Refer attached Mail. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2502042438 | 2025-05-16 | 2025-05-19 | Cust 574 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Conductivity Electrode | Forbes Marshall Cell Model 8312, K=1(FM -8312) | Conductivity cable has issue. client replace with spare cable and made transmitter functional. | Alpeshbhai guided over phone and resolved the issue, Refer attached mail. | Alpeshbhai guided over phone and resolved the issue, Refer attached mail. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2502263439 | 2025-05-21 | 2025-05-23 | Cust 496 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Data not updating on CPCB server. | There is issue from Hardware end confirmed by customer once resolve we will take access and check the issue. | There is issue from Hardware end confirmed by customer once resolve we will take access and check the issue. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2502333440 | 2025-05-22 | 2025-05-23 | Cust 462 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB server. | We took access to the system and observed data is not coming till modscan and camera IP not playing on google. | We took access to the system and observed data is not coming till modscan and camera IP not playing on google. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2502350441 | 2025-05-22 | 2025-05-26 | Cust 575 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | Commissioning | Urvesh guided over phone and visited the site for commissioning, Refer attached Mail. | Urvesh guided over phone and visited the site for commissioning, Refer attached Mail. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2502351442 | 2025-05-22 | 2025-05-23 | Cust 576 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Reading stuck | Urvesh guided over phone and resolved the issue, Refer atatched Mail. | Urvesh guided over phone and resolved the issue, Refer atatched Mail. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2502388443 | 2025-05-23 | 2025-05-26 | Cust 542 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | TSS Reading issue | Urvesh guided over the phone and resolved the issue by zero calibration. Refer attached Mail. | Urvesh guided over the phone and resolved the issue by zero calibration. Refer attached Mail. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2502463444 | 2025-05-25 | 2025-05-25 | Cust 577 | AHMEDABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Electrode | Polilyte HT 120(HT238 431) | calibration work | calibration done | calibration done | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFAD2502464445 | 2025-05-25 | 2025-05-25 | Cust 577 | AHMEDABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Electrode | Chemotrode ORP 120(HT238 740) | calibration work | calibration done | calibration done | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFAD2502491446 | 2025-05-26 | 2025-06-10 | Cust 578 | AHMEDABAD | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | SWAS & Analysers | Steam & Water Analysis System(SWAS & Analysers) | 3 Nos. pH sensor Cables Found faulty. Also in One Aqua4trans, Enter Key not Working. | Replacement material send by MARK EXPRESS PRIVATE LIMITED vide LR NO.-210145129 Dt- 09-06-25. | Replacement material send by MARK EXPRESS PRIVATE LIMITED vide LR NO.-210145129 Dt- 09-06-25. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2502509447 | 2025-05-26 | 2025-06-10 | Cust 574 | AHMEDABAD | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | Conductivity Tranmitter Smartpro | Conductivity Electrode Cable(Cond-Elec-cbl) | Issue with cable. When checked at office, observed discontinuity sometimes. | Complaint log in depot new generated complaint No.- CCS25D0369. | Complaint log in depot new generated complaint No.- CCS25D0369. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2502515448 | 2025-05-26 | 2025-05-29 | Cust 572 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning of Analyser at below site. Site address- M/s Maruti Suzuki India limited. Guj, BLOCK No.334, 335- OLD 158, 293,, VILLAGE HANSALPUR, NEAR BECHARAJI,,MANDAL, Gujarat | Chiragbhai and Jigar visited the site for commissioning of Analyser, Refer atatched MOM. | Chiragbhai and Jigar visited the site for commissioning of Analyser, Refer atatched MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFAD2502521449 | 2025-05-26 | 2025-06-10 | Cust 572 | AHMEDABAD | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | in FM Blackbox, Data not coming on RS485 | Complaint logged in Depot new generated complaint no.-CCS25D0368 NAME(Maruti Suzuki) | Complaint logged in Depot new generated complaint no.-CCS25D0368 NAME(Maruti Suzuki) | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2502558450 | 2025-05-27 | 2025-05-30 | Cust 496 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Dat not updating on Eversense +RSPCB+CPCB server. | We took access to the system and now camera and data is live on CPCB server.For RSPCB we have takenup with development team. | We took access to the system and now camera and data is live on CPCB server.For RSPCB we have takenup with development team. | YES | No Trip | Process | Y |  | Analytical - CHAKAN
AFAD2502563451 | 2025-05-27 | 2025-05-31 | Cust 501 | AHMEDABAD | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC visit. | Jigar visited the site for AMC of Analysers, Refer attached MOM. | Jigar visited the site for AMC of Analysers, Refer attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFAD2502563451 | 2025-05-27 | 2025-05-31 | Cust 501 | AHMEDABAD | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC visit. CX2000 Model. | Jigar visited the site for AMC of Analysers, Refer attached MOM. | Jigar visited the site for AMC of Analysers, Refer attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFAD2502644452 | 2025-05-28 | 2025-05-29 | Cust 478 | AHMEDABAD | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | CPU Display card for Multiparameter(AWA-CPU1) | 4th AMC visit | Chiragbhai visited the site for 4th AMC of Analyser, Refer attached MOM. | Chiragbhai visited the site for AMC, Refer attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFAD2502702453 | 2025-05-29 | 2025-05-30 | Cust 496 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on CPCB server. | We took access to the system and now camera is live on CPCB server. | We took access to the system and now camera is live on CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2502791454 | 2025-05-30 | 2025-06-02 | Cust 497 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Blackbox Commissioning | we took access of the Blackbox and completed the configuration, so now the Data is communicating till the FMBB, | we took access of the Blackbox and completed the configuration, so now the Data is communicating till the FMBB, | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFAD2503028455 | 2025-06-02 | 2025-06-04 | Cust 523 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Smartpro-con) | troubleshooting | Urvesh guided over phone and suspected issue with transmitter. Suggested to Send the transmitter to Pune for necessary checking / repair. | Urvesh guided over phone and suspected issue with transmitter. Suggested to Send the transmitter to Pune for necessary checking / repair. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2503029456 | 2025-06-02 | 2025-06-04 | Cust 576 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | pH reading issue | Urvesh guided over phone and resolved the issue by changing the sensor and cable. Refer attached mail. | Urvesh guided over phone and resolved the issue by changing the sensor and cable. Refer attached mail. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFAD2503092457 | 2025-06-04 | 2025-06-07 | Cust 470 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB server. | As of now, we are closing this complaint and will reopen it once we receive remote access to the system from coustomer end. | As of now, we are closing this complaint and will reopen it once we receive remote access to the system from coustomer end. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2503191458 | 2025-06-06 | 2025-06-08 | Cust 579 | AHMEDABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC and service of Analyser | Ajay and Jigar visited the site for AMC, Refer attached MOM. | Ajay and Jigar visited the site for AMC, Refer attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFAD2503271459 | 2025-06-09 | 2025-06-10 | Cust 464 | AHMEDABAD | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 4th AMC Visit | Urvesh visited the site for 4th AMC of Analyser, Refer attached MOM. | Urvesh visited the site for 4th AMC of Analyser, Refer attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFAD2503298460 | 2025-06-10 | 2025-06-13 | Cust 470 | AHMEDABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB server. | As of now w are closing this complaint we will reopen once we receive access from the customer. | As of now w are closing this complaint we will reopen once we receive access from the customer. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFAD2503315461 | 2025-06-10 | 2025-06-11 | Cust 477 | AHMEDABAD | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 1st AMC Visit | Ajay and Jigar visited the site for AMC of analyser, Refer attached MOM. | Ajay and Jigar visited the site for AMC of analyser, Refer attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFAD2503470462 | 2025-06-13 | 2025-06-17 | Cust 580 | AHMEDABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning of Analyser CX2000 | Urvesh visited the site for Commissioning, Refer to the attached MOM. | Urvesh visited the site for Commissioning, Refer to the attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBD2303754463 | 2023-06-11 | 2023-06-26 | Cust 581 | BANGLADESH | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | OXYSENS(OXYSENS) | Commissioning-calibration support request for supplied Do-sensor_Smartpro Tx(ETP plant) | Installed sensor-calibrated and commissioned the unit. | Visited on time, installed new Sensor, calibrated, and made mom but missed to close it. Will atrach Snap attd, MOM. | YES | Warranty (Free) | Intops | N |  | Analytical - CHAKAN
AFBD2303756464 | 2023-06-11 | 2023-06-15 | Cust 582 | BANGLADESH | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Silica Analyser | Silica 9210(9210) | SWAN Silica-AMI2 -not reading flow, customer requested for FM support. | Found 6-way valve internally rusted, replaced by spare unit, now unit working fine. | Found 6-way valve internally rusted, replaced by spare unit, now unit working fine. | NO | No Trip | Intops | N |  | Analytical - CHAKAN
AFBD2307551465 | 2023-08-20 | 2023-08-27 | Cust 583 | BANGLADESH | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Smartpro Electronics Assembly(Smartpro-Elc-cbl) | PH Analyzer/Electrode_ETP not responding, requested for FM support. End user: Sector Seven Int. Ltd(Wss) Nilphamari. | Checked &amp; found PH sensor FM1106 faulty, Tx ok. FM will send offer the for sensor soon. | Checked &amp;amp;amp; found PH sensor FM1106 faulty, Tx ok. FM will send offer the for sensor soon. | NO | No Trip | Intops | N |  | Analytical - CHAKAN
AFBD2310536466 | 2023-10-16 | 2023-10-22 | Cust 584 | BANGLADESH | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | Commissioning/Calibration of DO Tx and also to fix AO-PLC signal issue of supplied IFC, end user: Coca-Cola(Biotech-India), direct supply | Visited at site and somved AO issue and did calibration of DO. | Visited at site and somved AO issue and did calibration of DO. | YES | Warranty (Free) | Intops | N |  | Analytical - CHAKAN
AFBD2314300467 | 2023-12-27 | 2023-12-31 | Cust 585 | BANGLADESH | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(Ph-Trans-Smartpro) | Krohne Analytical project (PH,DO)_commissioning call for installed analyzers. | Visited and commissioned 3 sets of analyzer and three mag flow. | Visited and commissioned 3 sets of analyzer and three mag flow. | YES | Warranty (Free) | Intops | N |  | Analytical - CHAKAN
AFBD2314595468 | 2023-12-31 | 2024-01-05 | Cust 585 | BANGLADESH | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | optex series(optex series) | Commissioning support requested for balance PH/DO and (PH,DO & Ultrasonic LT-Open channel) | Run COND ANALYZER | Run COND ANALYZER | YES | Warranty (Free) | Intops | N |  | Analytical - CHAKAN
AFBD2315746469 | 2024-01-22 | 2024-01-28 | Cust 585 | BANGLADESH | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | optex series(optex series) | ADO-Imported, PH-Smartmak 400 & Open Channel-Rader Level Commissioning call. | DO & PH commissioning done, cond also run,cal is pending | DO & PH commissioning done, cond also run,cal is pending | YES | Warranty (Free) | Intops | N |  | Analytical - CHAKAN
AFBD2316267470 | 2024-01-31 | 2024-02-02 | Cust 586 | BANGLADESH | Analytical - CHAKAN | CLOSED | COMMISSIONING | TOC ANALYSER | BLOCK, INPUT, SAMPLE/DI, 1950 PLUS(11589A) | We need our Service Engineer for installation of Swan make TOC , procured from you as per PO No - KI/PO/23-24/334 Dated 09/06/23 Site address is as mentioned in attachment Kindly note our Project Engineer is Mr.Nimish.Chandel Mobile No - 8850406499 / 9619172525 | Dear Ankit, tried reaching you but the number registered on google is wrong. Kindly contact me on mtank@forbesmarshall.com or 8308861545. We will discuss first | Dear Ankit, tried reaching you but the number registered on google is wrong. Kindly contact me on mtank@forbesmarshall.com or 8308861545. We will discuss first | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFBD2316536471 | 2024-02-04 | 2024-02-08 | Cust 587 | BANGLADESH | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | TOC ANALYSER | FM 101 TOC Analyser(FM 101 TOC) | Erection-commissioning support request for TWO SETS of SWAN TOC Analyzers through OA: 5524102825, OEM: Komal Industries-India, end user: Reckitt Bencher-Chittagonj, Bangladesh. | Installation and Sample inlet TUBING is completed at site for the customer will call for commissioning shortly when the process will start. | Installation and Sample inlet TUBING is completed at site for TOC-customer will call for commissioning shortly when process will start. | YES | Warranty (Free) | Intops | N |  | Analytical - CHAKAN
AFBD2317614472 | 2024-02-25 | 2024-02-29 | Cust 584 | BANGLADESH | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | End user: Bio tech/Coca-cola BD: Troubleshooting/Calibration work of Do Analyzer on a chargeable basis | Solved the issue and done calibration at site. | Solved the issue and done calibration at site. | NO | Paid | Intops | N |  | Analytical - CHAKAN
AFBD2318518473 | 2024-03-13 | 2024-03-18 | Cust 587 | BANGLADESH | Analytical - CHAKAN | CLOSED | COMMISSIONING | TOC ANALYSER | TOC 1950 Single StreamANALYSER ANATEL(1950 TOC) | Commissioning request of two sets of SWAN make- TOCAnalyzers( One-Direct/OEM supplied & another by FMBD) | Visited and commissioned both Analyzers, mom attd. | Visited and commissioned both Analyzers, mom attd. | YES | Warranty (Free) | Intops | N |  | Analytical - CHAKAN
AFBD2318804474 | 2024-03-18 | 2024-03-24 | Cust 585 | BANGLADESH | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Transmitter Aquacon | Conductivity Transmitter Aquamon 2000(COND-Trans Aqu-2000) | The customer arranged PLC -Panel to connect our loop Powered IMPORTED Analyzers. Request through WSS. | Visited -connected one Analyzer with PLC and shown loop power connection to WSS engr for rest units at PLC. | Visited -connected one Analyzer with PLC and shown loop power connection to WSS engr for rest units at PLC. | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFBD2400969475 | 2024-04-22 | 2024-04-25 | Cust 588 | BANGLADESH | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | Turbidity Meter | Krohne turbidity transmitter(1) | CCDS-Cooler-turbidity-FCV-DPI installation support. | Visited and guided for installation at site. | Visited and guided for installation at site. | YES | Warranty (Free) | Intops | N |  | Analytical - CHAKAN
AFBD2401291476 | 2024-04-28 | 2024-05-04 | Cust 589 | BANGLADESH | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Commissioning support request for PH analyzer and also to do CAL of Orifice meter/DPT | commissioned DO Anzlyzer and also fix dram level DPT ORIFICE flow issue | commissioned DO Anzlyzer and also fix dram level DPT ORIFICE flow issue | YES | Warranty (Free) | Intops | N |  | Analytical - CHAKAN
AFBD2401293477 | 2024-04-28 | 2024-05-04 | Cust 590 | BANGLADESH | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | Installation/Erection work of two sets of SWAN DO Analyzer by modifying existing panel through PPCL. End user-NWPGCL, Khulna. | Guided to complete installation /weldingvwork by end user throgh PPCL physical site visit. | Guided to complete installation /weldingvwork by end user throgh PPCL physical site visit. | NO | No Trip | Intops | N |  | Analytical - CHAKAN
AFBD2407190478 | 2024-08-16 | 2024-08-19 | Cust 591 | BANGLADESH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | KSPCB data mismatch. | As of now we are closing this complaint and we will reopen this complaint once we get access to the system. | As of now we are closing this complaint and we will reopen this complaint once we get access to the system. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFBD2411207479 | 2024-11-03 | 2024-11-07 | Cust 588 | BANGLADESH | Analytical - CHAKAN | CLOSED | COMMISSIONING | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | CCDS with Krohne Turbidity/CV-PPPU, order through SSD/Boiler Commissioning support was requested from site. | Pre commissioning of the system done and commissioned CRS | Pre commissioning of the system done and commissioned CRS | YES | Warranty (Free) | Intops | N |  | Analytical - CHAKAN
AFBD2416248480 | 2025-02-03 | 2025-02-07 | Cust 588 | BANGLADESH | Analytical - CHAKAN | CLOSED | COMMISSIONING | Turbidity Meter | Krohne turbidity transmitter(1) | Sample lines completed by customer & requested for commissioning/cal support of supplied CCDS. | Commissioned CCDS system and handed over after calibration /validation by standard. | Commissioned CCDS system and handed over after calibration /validation by standard. | YES | Warranty (Free) | Intops | N |  | Analytical - CHAKAN
AFBG2303645481 | 2023-06-07 | 2023-06-12 | Cust 463 | BANGALORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Smartpro Electronics Assembly(Smartpro-Elc-cbl) | Found Milli Amps output value is low. When we Test the milliamps from 4-20 ma , If we test at 20 MA found 14.5 MA output only for both Transmitters, Found electronics faulty. | Electronices to be replace. | Electronices to be replace. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFBG2303645481 | 2023-06-07 | 2023-06-12 | Cust 463 | BANGALORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Smartpro Electronics Assembly(Smartpro-Elc-cbl) | Found Milli Amps output value is low. When we Test the milliamps from 4-20 ma , If we test at 20 MA found 14.5 MA output only for both Transmitters, Found electronics faulty. | Electronices to be replace. | Electronices to be replace. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFBG2303648482 | 2023-06-07 | 2023-06-09 | Cust 592 | BANGALORE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC visit to be done. | AMC visit done. | AMC visit done | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2303882483 | 2023-06-13 | 2023-06-15 | Cust 593 | BANGALORE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | data upload issue on kspcb /cpcb server. | zscalar security issue was tehre .right now internet is given through dongle.data is live on cpcb/kspec server. | zscalar security issue was tehre .right now internet is given through dongle.data is live on cpcb/kspec server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFBG2303923484 | 2023-06-13 | 2023-06-16 | Cust 594 | BANGALORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | calibration to be done. | Telephonic support provided for the calibration. | Telephonic support provided for the calibration. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFBG2304624485 | 2023-06-26 | 2023-06-29 | Cust 593 | BANGALORE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB & KSPCB data upload issue. | kspcb das updated to 4.6. | kspcb das updated to 4.6. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFBG2305016486 | 2023-07-03 | 2023-07-09 | Cust 595 | BANGALORE | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | AQUADO Analyzer | DPI-Digital process indicator(1) | DPI indicator not working. | Spare DPI is sent against the RF for the case. | Spare DPI is sent against the RF for the case. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFBG2305266487 | 2023-07-08 | 2023-07-09 | Cust 596 | BANGALORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | Commissioning support to be provided. | Telephonic support provided for the commissioning. | Telephonic support provided for the commissioning. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFBG2305775488 | 2023-07-17 | 2023-07-22 | Cust 597 | BANGALORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | AQUADO Analyzer | AQUAMAX DOT(1) | DO analyzer not working. | DO analyzer attended and solution convey to the customer. | Problem attended and solution convey to the customer. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2305909489 | 2023-07-20 | 2023-07-21 | Cust 598 | BANGALORE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | data nit updating on cpcb server | did not get access. | did not get access. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFBG2305938490 | 2023-07-20 | 2023-07-23 | Cust 593 | BANGALORE | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | PROACTIVE visit for MPA AMC to be done. | proactive visit done. | proactive visit done | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2306418491 | 2023-07-29 | 2023-07-31 | Cust 599 | BANGALORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | DO transmitter to be commissioned. | Telephonic support provided for the commissioning. | Telephonic support provided for the commissioning. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFBG2306703492 | 2023-08-01 | 2023-08-03 | Cust 600 | BANGALORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | commissioning Multiparameter analyser | Multiparameter analyser commissioning done. | Multiparameter analyser commissioning done. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFBG2306718493 | 2023-08-02 | 2023-08-05 | Cust 595 | BANGALORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | DO transmitter with new Display to be commissioned. | New Controller commisioned. | commissioning done successfully, | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2306720494 | 2023-08-02 | 2023-08-05 | Cust 600 | BANGALORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | MPA calibration and verification to be done. | MPA calibration done. | Problem attended and solution convey to the customer. | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2306787495 | 2023-08-03 | 2023-08-05 | Cust 601 | BANGALORE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MIL card of Multiparameter(1) | amc visit to be done. | AMC visit done. | AMC visit done | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2307266496 | 2023-08-14 | 2023-09-29 | Cust 602 | BANGALORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MUVX Card of Multiparameter(1) | Issue : MPA CX1000- RS485 MODBUS card burnt, new card to be supplied. | Replacement given to the customer for the same. Due to some reason it is not installed at the site. | Replacement given to the customer for the same. Due to some reason it is not installed at the site. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFBG2307685497 | 2023-08-21 | 2023-08-26 | Cust 592 | BANGALORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning of the repair PCX1 to be done. | Commissioning done and issue convey to the customer. | Problem attended and solution convey to the customer. | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2308067498 | 2023-08-28 | 2023-08-30 | Cust 592 | BANGALORE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation needs to be done. | We took the access to the system and successfully did software reinstallation + local + CPCB configuration, and now data is live on both the servers. | We took the access to the system and successfully did software reinstallation + local + CPCB configuration, and now data is live on both the servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFBG2308312499 | 2023-09-01 | 2023-09-04 | Cust 603 | BANGALORE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | KSPCB data upload issue. | We took the access to the customers PC, and update the KSPCB DAS, after theta now data is live on KSPCB server. | We took the access to the customers PC, and update the KSPCB DAS, after theta now data is live on KSPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFBG2308413500 | 2023-09-04 | 2023-09-11 | Cust 604 | BANGALORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | ph Electronicss to be commissioned. | Commissioning done. | commissioning done successfully, | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2308759501 | 2023-09-12 | 2023-09-16 | Cust 605 | BANGALORE | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Proactive visit to be done for analyzer. | Proactive visit done for the special analyzer. | proactive visit done | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2308760502 | 2023-09-12 | 2023-09-16 | Cust 593 | BANGALORE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC visit to be done. | AMC visit done. | AMC visit done | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2308761503 | 2023-09-12 | 2023-09-16 | Cust 606 | BANGALORE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(Ph-Trans-Smartpro) | Calibration to be done. | Calibration visit done. | calibration done successfully. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2309078504 | 2023-09-19 | 2023-09-23 | Cust 607 | BANGALORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | Commissioning to be done. | Commissioning done. | commissioning done successfully, | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2309080505 | 2023-09-19 | 2023-09-23 | Cust 608 | BANGALORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | MPA commissioning support to be provided. | Telephonic support provided for the troubleshooting. | Telephonic support provided for the troubleshooting. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFBG2309132506 | 2023-09-20 | 2023-09-23 | Cust 609 | BANGALORE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration to be done. | Calibration done. | calibration done successfully. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2309338507 | 2023-09-25 | 2023-09-29 | Cust 610 | BANGALORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Commissioning to be done. | Telephonic support provided for the commissioning. | Telephonic support provided for the commissioning. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFBG2309341508 | 2023-09-25 | 2023-09-29 | Cust 611 | BANGALORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration support to be provided. | Telephonic support provided for the calibration. | Telephonic support provided for the calibration. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFBG2309380509 | 2023-09-26 | 2023-09-26 | Cust 612 | BANGALORE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data in between is missing when we open the collective data of a month. | Took access of the system and successfully updated RTDAs now data is live on CPCB Server. | Took access of the system and successfully updated RTDAs now data is live on CPCB Server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFBG2309456510 | 2023-09-27 | 2023-09-29 | Cust 613 | BANGALORE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | APPCB data upload issue. Ruchika | We took the access to the system and found that the data is not communicating with the local system, then check data communication cable after that the data is live on local and APPCB server. | We took the access to the system and found that the data is not communicating with the local system, then check data communication cable after that the data is live on local and APPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFBG2309488511 | 2023-09-28 | 2023-09-30 | Cust 614 | BANGALORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Aqua4Trans | Display card of Aqua4Trans(1) | Display got blurr of two Instruments pH and TSS. | REPLACEMENT is arranging for the same and same will be courier at site. | REPLACEMENT is arranging for the same. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFBG2309498512 | 2023-09-28 | 2023-09-30 | Cust 595 | BANGALORE | Analytical - CHAKAN | CLOSED | PRODUCT UPGRADE | Aqua4Trans | Aqua4Trans(1) | Wash timer limit is not suitable as per the application. | REPLACEMENT is arranging for the same and same will be courier at site. | REPLACEMENT is arranging for the same and same will be courier at site. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFBG2309572513 | 2023-09-28 | 2023-09-30 | Cust 615 | BANGALORE | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Proactive visit for the MPA to be done. | PROACTIVE VISIT DONE. | proactive visit done | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2309752514 | 2023-09-30 | 2023-09-30 | Cust 600 | BANGALORE | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | MULTIPARAMETER ANALYSER | PH1 detector(1) | MPA showing error 2 and error8 . To check on the performance of the system | Visited the site and found the MPA operation satisfactory in colorless samples. Observed the ETP water is changing in every cycles and affecting the operation. Customer to avoid more dark colored sample solution to prevent such issues in future. | . Customer to avoid more dark colored sample solution to prevent such issues in future. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFBG2310265515 | 2023-10-10 | 2023-10-16 | Cust 614 | BANGALORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | CX2000 MPA to be commissioned. | Commissioning done. | commissioning done successfully, | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2310266516 | 2023-10-10 | 2023-10-16 | Cust 614 | BANGALORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Display card of Aqua4Trans(1) | new display card to be fixed. | Commissioning done. | commissioning done successfully, | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2310267517 | 2023-10-10 | 2023-10-16 | Cust 595 | BANGALORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Turbidity Input card of Aqua4Trans(1) | new instrument with firmware to be comissioned, | Commissioning done. | commissioning done successfully, | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2310586518 | 2023-10-17 | 2023-10-21 | Cust 616 | BANGALORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | commissioning to be done. | Commissioning done. | commissioning done successfully, | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2310759519 | 2023-10-21 | 2023-10-25 | Cust 615 | BANGALORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | MPA commissioning and Integration with connectivity to be done. | Telephonic support provided for the commissioning. | Telephonic support provided for the commissioning. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFBG2310895520 | 2023-10-25 | 2023-10-30 | Cust 617 | BANGALORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Commissioning of the pH system to be done. | Commissioning done. | commissioning done successfully, | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2311066521 | 2023-10-28 | 2023-10-30 | Cust 610 | BANGALORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(Ph-Trans-Smartpro) | Commissioning of the Ph Analyzer to be done. | Telephonic support provided for the commissioning. | Telephonic support provided for the commissioning. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFBG2311242522 | 2023-10-30 | 2023-11-05 | Cust 604 | BANGALORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Turbidity Meter | Optex turbidity sensor(1) | Turbidity Sensor not working. | Problem attended and found issue with the sensor. | Problem attended and solution convey to the customer. | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2311452523 | 2023-11-01 | 2023-11-03 | Cust 615 | BANGALORE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB and SPCB. | As of now we are closing this complaint and will reopen once we get access of the system from the customer end. | As of now we are closing this complaint and will reopen once we get access of the system from the customer end. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFBG2311599524 | 2023-11-05 | 2023-11-09 | Cust 610 | BANGALORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | pH Transmitter to be commissioned. | Telephonic support provided and arrange the RF. | Telephonic support provided and arrange the RF. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFBG2311600525 | 2023-11-05 | 2023-11-09 | Cust 591 | BANGALORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | MPA not working. | Telephonic support provided for the troubleshooting. | Telephonic support provided for the troubleshooting. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFBG2311836526 | 2023-11-09 | 2023-11-11 | Cust 615 | BANGALORE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation | We took access to the system and completed the software reinstallation and Local + CPCB connectivity. Now data is live on both servers. | We took access to the system and completed the software reinstallation and Local + CPCB connectivity. Now data is live on both servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFBG2311842527 | 2023-11-09 | 2023-11-13 | Cust 615 | BANGALORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning of the MPA and its connectivity to be done. | Commissioning done. | commissioning done successfully, | NO | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2312152528 | 2023-11-19 | 2023-11-19 | Cust 610 | BANGALORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | ORP transmitter(1) | ORP instrument to be commissioned. | Telephonic support provided for the commissioning. | Telephonic support provided for the commissioning. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFBG2312176529 | 2023-11-19 | 2023-11-26 | Cust 610 | BANGALORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | SS FTC not working at site. Site is in Surat Gujarat. | RF raised for the Replacement. | RF raised for the Replacement. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFBG2312251530 | 2023-11-21 | 2023-11-26 | Cust 614 | BANGALORE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration to be done. | calibration done. | calibration done successfully. | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2312443531 | 2023-11-25 | 2023-11-27 | Cust 618 | BANGALORE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not updating on CPCB and MPCB servers. | Completed CPCB and MPCB connectivity. Data is now live on CPCB and MPCB servers. | Data is now live on CPCB and MPCB servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFBG2312627532 | 2023-11-27 | 2023-12-02 | Cust 619 | BANGALORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Ph Transmitter not working. | Problem attended and resolve the issue. | problem attended in resolved the issue. | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2312721533 | 2023-11-28 | 2023-12-02 | Cust 614 | BANGALORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | pH readings are not okay. | pH working fine | commissioning done successfully, | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2312722534 | 2023-11-28 | 2023-12-02 | Cust 604 | BANGALORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Turbidity Input card of Aqua4Trans(1) | Turbidity Sensor to be fixed. | Problem attended and solution convey to the customer. | Problem attended and solution convey to the customer. | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2313087535 | 2023-12-04 | 2023-12-06 | Cust 618 | BANGALORE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue | data is now live on CPCB and MPCB servers. | data is now live on CPCB and MPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFBG2313130536 | 2023-12-04 | 2023-12-10 | Cust 620 | BANGALORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Turbidity Meter | Krohne turbidity transmitter(1) | Commissioning of the Turbidity meter tobe done. | Commissioning done. | commissioning done successfully, | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2313183537 | 2023-12-05 | 2023-12-09 | Cust 621 | BANGALORE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Calibration of the Analytical instruments to be done. | Calibration done. | calibration done successfully. | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2313336538 | 2023-12-09 | 2024-02-13 | Cust 614 | BANGALORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Analyzer pH reading is not matching with Check manu or actual values . | Site visited and resolve the issue by software update. | Site visited and resolve the issue by software update. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFBG2313356539 | 2023-12-10 | 2023-12-10 | Cust 610 | BANGALORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Commissioning of the instrument to be done. | Telephonic support provided for the commissioning. | Telephonic support provided for the commissioning. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFBG2313461540 | 2023-12-12 | 2023-12-15 | Cust 601 | BANGALORE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC to be done. | AMC visit done. | AMC visit done | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2313487541 | 2023-12-12 | 2023-12-15 | Cust 622 | BANGALORE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Calibration to be done. | Calibration done. | calibration done successfully. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2313489542 | 2023-12-12 | 2023-12-15 | Cust 571 | BANGALORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4 trans-A4TW033000(1) | Commissioning of DO meter to be done. | Telephonic support provided for addressing the issue. | Telephonic support provided for addressing the issue. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFBG2313493543 | 2023-12-12 | 2023-12-15 | Cust 592 | BANGALORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | MPA sample pump is not working. | Telephonic support provided for addressing the issue. | Telephonic support provided for addressing the issue. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFBG2313572544 | 2023-12-14 | 2023-12-19 | Cust 534 | BANGALORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | As per customer email the DO meter (PO no.1238/10 dtd.28.11.2022 Sr. No. 23SDG033) and are not giving readings. Please attend as per below Site Details- Jubilant Foodworks, Plot No. 58, Adinarayana Hosahalli, Doddaballapur, Bengaluru Rural, Karnataka, 561203 Site person contact details: Mr. Arun Kasar- 9819714376 | Site visited and solution convey to the customer. As there is problem with the hardware we suggest them to send the unit to pune for testing and checking. | Problem attended and solution convey to the customer. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2313610545 | 2023-12-14 | 2023-12-14 | Cust 623 | BANGALORE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB connectivity needs to be done. | We did CPCB connectivity, and now data is live on CPCB server. | We did CPCB connectivity, and now data is live on CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFBG2314120546 | 2023-12-24 | 2023-12-25 | Cust 614 | BANGALORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | BMS integration to be done with BIAL system. | Telephonic support provided for the modbus integration to BIAL and their team over the phone. | Telephonic support provided for the modbus integration to BIAL and their team over the phone. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFBG2314253547 | 2023-12-26 | 2023-12-31 | Cust 620 | BANGALORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Turbidity Input card of Aqua4Trans(1) | commissioning to be done. | Commissioning done. | commissioning done successfully, | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2314255548 | 2023-12-26 | 2023-12-31 | Cust 624 | BANGALORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Turbidity Input card of Aqua4Trans(1) | commissioning to be done. | commissioning done. | commissioning done successfully, | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2314401549 | 2023-12-28 | 2023-12-29 | Cust 625 | BANGALORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | ph system not working. | Problem attended and Solution convey to the customer. | Problem attended and solution convey to the customer. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2314944550 | 2024-01-06 | 2024-01-06 | Cust 616 | BANGALORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | MPA not working. | Problem attended and resolve the issue. | problem attended in resolved the issue. | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2315090551 | 2024-01-09 | 2024-01-12 | Cust 626 | BANGALORE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue. | we took access to the system and now data is live on the CPCB server. | we took access to the system and now data is live on the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFBG2315311552 | 2024-01-13 | 2024-01-20 | Cust 610 | BANGALORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Electrode Cable(Ph-Cab) | Ph Electrode cable is not working. | RF is arranging for the cable. | RF is arranging for the cable. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFBG2315494553 | 2024-01-16 | 2024-01-19 | Cust 603 | BANGALORE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB connectivity needs to be done. | we took access to your system and did CPCB configuration for Spent wash to Inceneration boiler & 65_2 TPH incenerations boilers and now data is live on the CPCB server. | we took access to your system and did CPCB configuration for Spent wash to Inceneration boiler &amp;amp; 65_2 TPH incenerations boilers and now data is live on the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFBG2315530554 | 2024-01-17 | 2024-01-19 | Cust 627 | BANGALORE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Reconfiguration of software in new pc. | We took access to the system and successfully completed the installation and reconfiguration of software now data is live on all servers. | We took access to the system and successfully completed the installation and reconfiguration of software now data is live on all servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFBG2315554555 | 2024-01-18 | 2024-01-22 | Cust 534 | BANGALORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | DO transmiter to be commissioned. | Commisioning done. | problem attended in resolved the issue. | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2315860556 | 2024-01-24 | 2024-01-25 | Cust 628 | BANGALORE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB + KSPCB data upload issue | Now data is live on CPCB + KSPCB server. | Now data is live on CPCB + KSPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFBG2316090557 | 2024-01-30 | 2024-02-03 | Cust 614 | BANGALORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | pH system not working. | Software updated and now its working fine. | problem attended in resolved the issue. | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2316440558 | 2024-02-01 | 2024-03-25 | Cust 463 | BANGALORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Electrode | FM1106(1) | FM1106- 2 sensors of sl no:- FM1106052305036, FM110605230637 failed within warranty period after installing in 10 days. | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210091208 LR Dt - 04-MAR-2024 | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210091208 LR Dt - 04-MAR-2024 | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFBG2316440558 | 2024-02-01 | 2024-03-25 | Cust 463 | BANGALORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Electrode | FM1106(1) | FM1106- 2 sensors of sl no:- FM1106052305036, FM110605230637 failed within warranty period after installing in 10 days. | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210091208 LR Dt - 04-MAR-2024 | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210091208 LR Dt - 04-MAR-2024 | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFBG2316490559 | 2024-02-02 | 2024-02-03 | Cust 624 | BANGALORE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Display board configuration needs to be done. | We took the access to your system and did Display board configuration. | We took the access to your system and did Display board configuration. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFBG2316516560 | 2024-02-03 | 2024-02-10 | Cust 629 | BANGALORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | Commissioning to be done. | Telephonic support provided for the commissioning. | Telephonic support provided for the commissioning. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFBG2316518561 | 2024-02-03 | 2024-02-10 | Cust 630 | BANGALORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Conductivity Electrode | Forbes Marshall Cell Model 8311, K=0.1(FM -8311) | Conductivity measurement is not withing the range. | Telephonic support provided for the calibration. | Telephonic support provided for the calibration. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFBG2316635562 | 2024-02-06 | 2024-02-13 | Cust 631 | BANGALORE | Analytical - CHAKAN | CLOSED | PRODUCT UPGRADE | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | Issue : Replace conductivity sensor model FM8312 and supply FM8311 model due to end users change in process range at site. Cust Contact details : Name : Mr.Rajendra Kannan Mob : 9964735567 Email : rajendra@jananiwater.com Address : same as OA. Shipping Instructions : Air ways , To pay basis. | Replacement provided and its reached to the customer. | Replacement provided and its reached to the customer. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFBG2316655563 | 2024-02-06 | 2024-02-09 | Cust 632 | BANGALORE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data mismatch issue. | Now data is matching on all the servers | Now data is matching on all the servers | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFBG2316658564 | 2024-02-06 | 2024-02-11 | Cust 620 | BANGALORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Turbidity Meter | Optex turbidity sensor(1) | commissioning to be done. | Commissioning done. | commissioning done successfully, | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2316735565 | 2024-02-08 | 2024-02-11 | Cust 603 | BANGALORE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data mismatch issue. | Now data is matching to CPCB server with actual readings. | Now data is matching to CPCB server with actual readings. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFBG2316909566 | 2024-02-12 | 2024-02-15 | Cust 603 | BANGALORE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data mismatch issue. | we took the access to your system and did some changes from configuration end and now flow meter data is matching to the actual reading. | we took the access to your system and did some changes from configuration end and now flow meter data is matching to the actual reading. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFBG2316983567 | 2024-02-13 | 2024-02-18 | Cust 610 | BANGALORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Aqua4Trans | Aqua4 trans-A4TW033000(1) | ph Measurement page is not appearing. | Telephonic support provided and asked them to send the unit at our works. | Telephonic support provided and asked them to send the unit at our works. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFBG2317325568 | 2024-02-20 | 2024-02-24 | Cust 593 | BANGALORE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | MAC visit to be done. | AMC Visit done. | AMC visit done | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2317476569 | 2024-02-22 | 2024-02-24 | Cust 633 | BANGALORE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Calibration of the Ph and other instruments to be done, | Calibration done. | calibration done successfully. | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2317610570 | 2024-02-25 | 2024-02-27 | Cust 614 | BANGALORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | TSS readings are not stable. | Problem attended and solution convey to the customer. | Problem attended and solution convey to the customer. | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2318115571 | 2024-03-01 | 2024-03-05 | Cust 634 | BANGALORE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration of the analyzer to be done. | Calibration done. | calibration done successfully. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2318167572 | 2024-03-04 | 2024-04-02 | Cust 463 | BANGALORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Electrode | FM1106(1) | We have supplied FM1106-FG of sl no:-A2110293856, A2110293857 is failed within 2 days of installation and the material is within warranty period. hence we need to supply 2 number of FM1106-P on urgent basis. | Replacment material has been booked through TCI EXPRESS vide LR No - 110007585144 LR Dt - 26-MAR-2024 | Replacment material has been booked through TCI EXPRESS vide LR No - 110007585144 LR Dt - 26-MAR-2024 | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFBG2318170573 | 2024-03-04 | 2024-03-07 | Cust 635 | BANGALORE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | CAMERA STREAMING ISSUE. | We have checked and currently the camera is streaming live on CPCB server. | We have checked and currently the camera is streaming live on CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFBG2318187574 | 2024-03-04 | 2024-03-07 | Cust 592 | BANGALORE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | some of additional/extra unwanted parameters/devices added to cpcb server. | unwanted devices which were added to site are removed. | unwanted devices which were added to site are removed. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFBG2318444575 | 2024-03-11 | 2024-03-15 | Cust 601 | BANGALORE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC visit to be done. | MAC visit done. | AMC visit done | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2318850576 | 2024-03-18 | 2024-03-23 | Cust 636 | BANGALORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | MPA analyzer issues to be addressed. | Problem attended and solution convey to the customer. | Problem attended and solution convey to the customer. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2319566577 | 2024-03-29 | 2024-03-30 | Cust 614 | BANGALORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Multiparameter readings are not satisfactory. | Problem attended and solution convey to the customer. | problem attended in resolved the issue. | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2400100578 | 2024-04-02 | 2024-04-03 | Cust 624 | BANGALORE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB + KSPCB configuration needs to be done. | we took the access to your system and completed CPCB + KSPCB connectivity. And now data is live on both the servers for Stack_2_Boiler_50TPH. | we took the access to your system and completed CPCB + KSPCB connectivity. And now data is live on both the servers for Stack_2_Boiler_50TPH. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFBG2400532579 | 2024-04-11 | 2024-04-16 | Cust 592 | BANGALORE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | pH input card-FM of Multiparameter(1) | AMC visit to be done. | A/MC visit done. | AMC visit done | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2400661580 | 2024-04-15 | 2024-04-19 | Cust 637 | BANGALORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Calibration to be done and repeatability to be shown. | Calibration done and repeatability shown. | calibration done successfully. | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2400776581 | 2024-04-18 | 2024-04-23 | Cust 617 | BANGALORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Turbidity Meter | Optex turbidity sensor(1) | Commissioning to be done. | Commissioning done. | commissioning done successfully, | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2401033582 | 2024-04-23 | 2024-04-28 | Cust 636 | BANGALORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4 trans-A4TW033000(1) | Commissioning of the aqua4trans turbdity meter to be done. | Commissioning of the Turbidity meter done. | commissioning done successfully, | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2401674583 | 2024-05-03 | 2024-05-06 | Cust 592 | BANGALORE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue. | We took access to your system and now the issue has been resolved and data is live on CPCB + KSPCB server. | We took access to your system and now the issue has been resolved and data is live on CPCB + KSPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFBG2401741584 | 2024-05-06 | 2024-05-08 | Cust 624 | BANGALORE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Display board configuration needs to be done. | We took the access to your system and completed Display board configuration. And display board files are generated with a frequency of 1 minute on C drive: /DisplayBoard folder with name VSPL | We took the access to your system and completed Display board configuration. And display board files are generated with a frequency of 1 minute on C drive: /DisplayBoard folder with name VSPL | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFBG2401918585 | 2024-05-09 | 2024-05-12 | Cust 614 | BANGALORE | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Proactive visit for the digitalization to be done. | Proactive visit for the MPA done. | proactive visit done | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2401999586 | 2024-05-12 | 2024-05-18 | Cust 610 | BANGALORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Smartpro-con) | Commissioning of the conductivity mteer to be done. | Telephonic support provided for the commissioning of the conductivity meter. | Telephonic support provided for the commissioning of the conductivity meter. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFBG2402080587 | 2024-05-14 | 2024-05-20 | Cust 638 | BANGALORE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Please note that all analysers need to be calibrated and cross verify with the systems. Some of pH and conductivity transmitters are showing erratic reading on display. Please note OAs' - 5524104944 ; 5524103103 Site address: Tata Electronics Pvt Ltd Thimjepalle Village Kelamaglam Royakottai Road Hosur Krishnagiri Districts Tamil Nadu-635113 Mr. Mohan - +91 9524562894 Mr. Ankit - +91 8879989641 | Checked with Site Engineer and ask them to arrange the Buffer solution for calibration and replace the pH Electrode. | Checked with Site Engineer and ask them to arrange the Buffer solution for calibration. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFBG2402484588 | 2024-05-23 | 2024-05-26 | Cust 639 | BANGALORE | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | Proactive visit to be done. | Proactive visit done. | proactive visit done | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2402492589 | 2024-05-23 | 2024-05-24 | Cust 640 | BANGALORE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Central CPCB data upload issue. | Please check now the data is live on the CPCB server. | Please check now the data is live on the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFBG2402519590 | 2024-05-23 | 2024-05-30 | Cust 638 | BANGALORE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Electrode | FM1104(1) | In pH they are getting difference of pH and it seems issue with pH electrode we have asked them to replace it as pH electrode is consumable item. Site & Contact Details - Tata Electronics - Hosur, Banglore. Site Engg: Ankit Singh - (+91 8879989641) Mr. Mohan Kumaar - (+91 9524562894) | Ph and conductivity analyzers Calibration done with buffers and found okay. | calibration done successfully. | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2403133591 | 2024-05-31 | 2024-06-04 | Cust 641 | BANGALORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(Ph-Trans-Smartpro) | Commissioning to be done. | Commissioning done. | commissioning done successfully, | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2403373592 | 2024-06-06 | 2024-06-09 | Cust 614 | BANGALORE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration of the MPA to be done. | Calibration done. | calibration done successfully. | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2403940593 | 2024-06-18 | 2024-06-23 | Cust 610 | BANGALORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4 trans-A4TW033000(1) | Commissioning support to be provided for the Aqua4trans. | Telephonic support provided for the commissioning. | Telephonic support provided for the commissioning. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFBG2403941594 | 2024-06-18 | 2024-06-23 | Cust 610 | BANGALORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(Ph-Trans-Smartpro) | Commissioning support for the pH to be provided. | Telephonic support provided for the commissioning. | Telephonic support provided for the commissioning. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFBG2404015595 | 2024-06-20 | 2024-06-23 | Cust 614 | BANGALORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | MPA not working. | Calibration done and found working okay. | problem attended in resolved the issue. | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2404020596 | 2024-06-20 | 2024-06-23 | Cust 601 | BANGALORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Smartpro-con) | Conductivity analyzer not working. | Telephonic support provided for the troubleshooting. | Telephonic support provided for the troubleshooting. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFBG2404375597 | 2024-06-27 | 2024-06-28 | Cust 642 | BANGALORE | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Proactive visit to be done. | Proactive visit done FOR CX2000. | proactive visit done | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2404376598 | 2024-06-27 | 2024-06-28 | Cust 643 | BANGALORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Electrode | Ph Electrode(FM1104(4C)) | pH Electrode to be checked for its performance. | Problem attended and solution convey to the customer. | Problem attended and solution convey to the customer. | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2405290599 | 2024-07-12 | 2024-07-15 | Cust 644 | BANGALORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | pH transmitters not working. CHEMPLAST SANMAR LIMITED, Custom Manufactured Chemicals Divn., , 44 ,Theertham Road, Berigai, 635105 HOSUR TN | Problem attended and solution ocnvey to the customer. | Problem attended and solution convey to the customer. | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2405317600 | 2024-07-13 | 2024-07-15 | Cust 645 | BANGALORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Ph System not working. | Telephonic support provided and resolve the issue. | Telephonic support provided and resolve the issue. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFBG2405421601 | 2024-07-15 | 2024-07-19 | Cust 636 | BANGALORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | MPA analyzer not working. | Problem attended and resolve the issue. | problem attended in resolved the issue. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2405422602 | 2024-07-15 | 2024-07-19 | Cust 592 | BANGALORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning of the repaired analyzer to be done. | Commissioning done. | commissioning done successfully, | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2405776603 | 2024-07-22 | 2024-07-27 | Cust 636 | BANGALORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning of the MPA spares to be done at STP. | Commissioning done. | commissioning done successfully, | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2406359604 | 2024-07-30 | 2024-08-05 | Cust 615 | BANGALORE | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Connectivity related proactive visit to be done. | Proactive visit done. | proactive visit done | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2406437605 | 2024-07-31 | 2024-08-03 | Cust 646 | BANGALORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | pH input card-FM of Multiparameter(1) | commissioning required - 552104868, ref attached mail | Discussion had with the customer and asked them installed the analyzer for the commissioning. Analyzer not installed and sample line not provided. UPS supply to be given. | Discussion had with the customer and asked them installed the analyzer for the commissioning. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFBG2406579606 | 2024-08-02 | 2024-08-03 | Cust 636 | BANGALORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | MPA analyzer not working. | Commissioning done. | commissioning done successfully, | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2406746607 | 2024-08-06 | 2024-08-08 | Cust 592 | BANGALORE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data upload issue. | we took access of the black box and made some modifications from configuration end, now data is live on the CPCB server. | we took access of the black box and made some modifications from configuration end, now data is live on the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFBG2406867608 | 2024-08-08 | 2024-08-10 | Cust 614 | BANGALORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Readings were not matching with actual values. | Problem attended and resolve the issue. | problem attended in resolved the issue. | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2406970609 | 2024-08-12 | 2024-08-14 | Cust 603 | BANGALORE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue. | for now we are closing this complaint. | for now we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFBG2406975610 | 2024-08-12 | 2024-08-14 | Cust 603 | BANGALORE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB + KSPCB data upload issue. | for now we are closing tbis complaint. | for now we are closing tbis complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFBG2407016611 | 2024-08-12 | 2024-08-16 | Cust 647 | BANGALORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Chlorine Analyzer-MAC 100 | MAC 100(1) | commissioning support is required for newly supplied analyser | telephonic support provided and observed that the its sensor membrane got damaged. Communication shared with the customer. | telephonic support provided and observed that the its sensor membrane got damaged. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFBG2407137612 | 2024-08-15 | 2024-08-15 | Cust 648 | BANGALORE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | Calibration support to be provided for the conductivity analyzer. | Problem attended and guided over the phone. | Problem attended and guided over the phone. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFBG2407500613 | 2024-08-22 | 2024-08-24 | Cust 634 | BANGALORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | Service visit required against Oa no: 5523108748 | Telephonic support provided for the same and asked them to send the unit for the repair. | Telephonic support provided for the same and asked them to send the unit for the repair. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFBG2407553614 | 2024-08-23 | 2024-08-23 | Cust 591 | BANGALORE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | data mismatch issue on KSPCB. | Now the KSPCB data upload issue has been resolved. | Now the KSPCB data upload issue has been resolved. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFBG2407730615 | 2024-08-26 | 2024-08-29 | Cust 617 | BANGALORE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | Calibration of the conductivity meter to be done. | Calibration support provided. | Calibration support provided. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFBG2407753616 | 2024-08-27 | 2024-08-29 | Cust 638 | BANGALORE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Calibration to be done. | Calibration done. | calibration done successfully. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2407844617 | 2024-08-28 | 2024-09-04 | Cust 649 | BANGALORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | rectify the issue of DO transmitter | Problem attended and solution convey to the customer. | Problem attended and solution convey to the customer. | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2407927618 | 2024-08-29 | 2024-08-31 | Cust 592 | BANGALORE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC visit to be done. | AMC visit for the instrument is done. | AMC visit done | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2407948619 | 2024-08-30 | 2024-08-31 | Cust 634 | BANGALORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Not powering on | Found RS485 card faulty creating load in amps removed the same and checked found working ok | Found RS485 card Found faulty removed and checked found working ok | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2407979620 | 2024-08-30 | 2024-09-02 | Cust 650 | BANGALORE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Transmitter Aquacon | Conductivity Transmitter Aquamon 2000(COND-Trans Aqu-2000) | Calibration support to be provided for the conductivity analyzer. | Telephonic support provided for the calibration. | Telephonic support provided for the calibration. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFBG2408060621 | 2024-08-31 | 2024-09-02 | Cust 610 | BANGALORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | Commissioning support to be provided. | Telephonic support provided for the commissioning. | Telephonic support provided for the commissioning. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFBG2408272622 | 2024-09-02 | 2024-09-05 | Cust 592 | BANGALORE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue. | Now data is live on CPCB. | Now data is live on CPCB. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFBG2408281623 | 2024-09-02 | 2024-09-07 | Cust 591 | BANGALORE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration of the MPA to be done. | Calibration done. | calibration done successfully. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2408376624 | 2024-09-04 | 2024-09-07 | Cust 601 | BANGALORE | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | proactive visit to be done. | AMC visit done. | AMC visit done | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2408673625 | 2024-09-10 | 2024-09-14 | Cust 646 | BANGALORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning to be done. | Site is not ready hence closing this complaint. | Site is not ready hence closing this complaint. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFBG2409043626 | 2024-09-18 | 2024-09-21 | Cust 646 | BANGALORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning to be done. | Commissioning done. | commissioning done successfully, | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2409045627 | 2024-09-18 | 2024-09-21 | Cust 625 | BANGALORE | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Proactive visit to be done. | Proactive visit done. | proactive visit done | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2409047628 | 2024-09-18 | 2024-09-21 | Cust 608 | BANGALORE | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Proactive viist to be done | Proactive visit done. | proactive visit done | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2409049629 | 2024-09-18 | 2024-09-21 | Cust 610 | BANGALORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Commissioning to be done. | Telephonic Support provided for the commissioning. | Telephonic Support provided for the commissioning. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFBG2409050630 | 2024-09-18 | 2024-09-21 | Cust 631 | BANGALORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4Trans(1) | Commissioning support to be provided. | Telephonic Support provided for the commissioning. | Telephonic Support provided for the commissioning. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFBG2409051631 | 2024-09-18 | 2024-09-21 | Cust 564 | BANGALORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | ORP transmitter(1) | Commissioning support to be provided. | Telephonic Support provided for the commissioning. | Telephonic Support provided for the commissioning. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFBG2409212632 | 2024-09-23 | 2024-09-26 | Cust 592 | BANGALORE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Customer requested for the CPCB URL. | Shared the CPCB link and procedure to change the URL with the customer. | Shared the CPCB link and procedure to change the URL with the customer. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFBG2409634633 | 2024-09-30 | 2024-10-03 | Cust 592 | BANGALORE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB Central Portal Data upload issue. | Data is live on our server hence we are asked customer to check this issue with the CPCB team directly. | Data is live on our server hence we are asked customer to check this issue with the CPCB team directly. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFBG2409804634 | 2024-10-02 | 2024-10-06 | Cust 647 | BANGALORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Chlorine sensor | CHLORINE SENSOR(1) | Chlorine sensor to be commissioned. | Problem attended and solution convey to the customer. | Problem attended and solution convey to the customer. | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2409964635 | 2024-10-07 | 2024-10-10 | Cust 624 | BANGALORE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB + KSPCB + Local server data upload issue. | We took access to the system and now the issue has been resolved and the local server is accessible and data is live on CPCB + Local + KSPCB server. | We took access to the system and now the issue has been resolved and the local server is accessible and data is live on CPCB + Local + KSPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFBG2410074636 | 2024-10-09 | 2024-10-09 | Cust 543 | BANGALORE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Data upload issue on KSPCB server. | Now the issue has been resolved and data is live on the KSPCB server. | Now the issue has been resolved and data is live on the KSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFBG2410711637 | 2024-10-22 | 2024-10-26 | Cust 631 | BANGALORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Aqua4Trans | Display card of Aqua4Trans(1) | Display Not working. | RF RAISED for the dispatched the display card. | RF RAISED for the dispatched the display card. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFBG2410818638 | 2024-10-24 | 2024-10-25 | Cust 651 | BANGALORE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | KSPCB configuration needs to be done. | We have completed KSPCB configuration and now data is live on KSPCB server. | We have completed KSPCB configuration and now data is live on KSPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFBG2410821639 | 2024-10-24 | 2024-10-25 | Cust 593 | BANGALORE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | data upload issue on CPCB server | Data is live on the Local and CPCB server. | Data is live on the Local and CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFBG2410984640 | 2024-10-28 | 2024-10-30 | Cust 593 | BANGALORE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | KSPCB data uploading issue | Data is live on the KSPCB server. | Data is live on the KSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFBG2411263641 | 2024-11-04 | 2024-11-06 | Cust 593 | BANGALORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | Flow cell for Multiparameter(AWA-FLOWCELL) | MPA not working. | Telephonic support provided for troubleshooting the MPA. | Telephonic support provided for troubleshooting the MPA. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFBG2411298642 | 2024-11-05 | 2024-11-06 | Cust 652 | BANGALORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Aqua4Trans | Turbidity Input card of Aqua4Trans(1) | Aqua4Trans Turbidity not working. | Support provided for the troubleshooting. | Support provided for the troubleshooting. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFBG2411594643 | 2024-11-11 | 2024-11-16 | Cust 624 | BANGALORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Turbidity Meter | Optex(FMSC-P3) | Commissioning to be done. | Proactive visit for the Turbidity meter done. | proactive visit done | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2411724644 | 2024-11-13 | 2024-11-19 | Cust 653 | BANGALORE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Aqua4Trans | Aqua4Trans(1) | We purchased on TDS cum pH Transmitter for one of our project in puttur, Kurnataka agains Tax Invoice No. 5555005587 Dated 14/10/2024. Details of Transmitter are as follows; 1. Model No.- Aqua 4Trans-4T Power Input -240VAC 4-20 mA Output -2 Nos Relay Output -4 Nos TDS Sensor -FM8934,1.5 " TC pH Sensor - FM 1104 Channel No. B for TDS is working fine. We installed the same and faces the problems are as below; 1. Transmtter get Overheated . 2. Channel No. A for pH Sensor not shown the appropraite reading. 3. Channel No. A connectors are loose. To rectify this we tried to calibrate the senso with Standard solution but calibration gets failed. We tried to call your service personnel Mr. Chandrashekhar and Mr. Nilesh Gawade, but they didn?t responded. We have pH Sensor & cable with us. The Address of site is as follows; Megha Fruit Processing Pvt. Ltd., Narimogeru,Puttur -574202, D.K. Karnataka. Contact Person Name - Mr. Suresh Bhat, Contact Details- +91988088176/ +91 8147789341 | I contacted Mr. Bhat regarding the site issue and he told that the site is not handed over and not aware of the issues. Site engineer is not at site and nothing cleared about the site issue neither contact details are available. Hence closing the site for now and log once again if the details are available. | Incomplete details and site engineer is not at site hence closing this. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFBG2411942645 | 2024-11-18 | 2024-11-20 | Cust 612 | BANGALORE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data uploading issue on CPCB | Telephonic support provided and resolve the issue in SPM. | Telephonic support provided and resolve the issue in SPM. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFBG2412026646 | 2024-11-19 | 2024-11-24 | Cust 615 | BANGALORE | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | proactive visit to be done. | Proactive visit done. | proactive visit done | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2412027647 | 2024-11-19 | 2024-11-24 | Cust 628 | BANGALORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Troubleshooting to be done of MPA. | Problem attended and solution convey to the customer. | Problem attended and solution convey to the customer. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2412760648 | 2024-12-01 | 2024-12-06 | Cust 601 | BANGALORE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC visit to be done. | AMC VISIT DONE | AMC visit done | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2412761649 | 2024-12-01 | 2024-12-06 | Cust 608 | BANGALORE | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Proactive visit to be done. | PROACTIVE VISIT DONE | proactive visit done | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2412908650 | 2024-12-03 | 2024-12-06 | Cust 654 | BANGALORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | commissioning to be done. | commissioning done. | commissioning done successfully, | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2412997651 | 2024-12-05 | 2024-12-06 | Cust 655 | BANGALORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | commissioning to be done. | commissioning done. | commissioning done successfully, | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2412998652 | 2024-12-05 | 2024-12-07 | Cust 653 | BANGALORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | PH input card(1) | PH TRANSMITTER TO BE COMMISSIONED. | Telephonic support provided for the commissioning of the Aqua4trans with Ph. Customer is MEGHA FOOD. | Telephonic support provided for the commissioning of the Aqua4trans with Ph. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFBG2413317653 | 2024-12-12 | 2024-12-15 | Cust 593 | BANGALORE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software re-installation | We are closing this complaint for now and will reopen it once we have access to the system. | We are closing this complaint for now and will reopen it once we have access to the system. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFBG2413560654 | 2024-12-18 | 2024-12-21 | Cust 628 | BANGALORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data upload issue. | for now we are closing this complaint. | for now we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFBG2413603655 | 2024-12-18 | 2024-12-22 | Cust 615 | BANGALORE | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Proactive visit to be done. | Proactive visit done. | proactive visit done | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2413605656 | 2024-12-18 | 2024-12-22 | Cust 628 | BANGALORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning to be done. | Commissioning to be done. | commissioning done successfully, | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2413820657 | 2024-12-23 | 2024-12-23 | Cust 628 | BANGALORE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software installation + CPCB + KSPCB configuration needs to be done. | We have completed software installation + CPCB + KSPCB configuration now data is live on both servers. | We have completed software installation + CPCB + KSPCB configuration now data is live on both servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFBG2413900658 | 2024-12-24 | 2024-12-24 | Cust 656 | BANGALORE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Cameras not playing on CPCB server. | We took access to the system and now camera is live on CPCB server. | We took access to the system and now camera is live on CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFBG2413950659 | 2024-12-25 | 2024-12-28 | Cust 593 | BANGALORE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC visit to be done. | AMC visit done. | AMC visit done | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2414468660 | 2024-12-31 | 2024-12-31 | Cust 636 | BANGALORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | MPA not working. | Problem attended and solution convey to the customer. | Problem attended and solution convey to the customer. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2414565661 | 2025-01-02 | 2025-01-04 | Cust 657 | BANGALORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Electrode | Polilyte Plus 120(HT242 431) | pH Electrode not working. | Problem attended and Solution convey to the solution. | Problem attended and solution convey to the customer. | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2414568662 | 2025-01-02 | 2025-01-04 | Cust 658 | BANGALORE | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | AQUADO Analyzer | AQUAMAX DOT(1) | Proactive visit for the DO to be done. | Proactive visit done. | proactive visit done | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2414583663 | 2025-01-03 | 2025-01-03 | Cust 659 | BANGALORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Chlorine sensor | CHLORINE SENSOR(1) | Chlorine analyser always showing Zero PPM at higher dosing. | Telephonic support provided and resolve the issue. | Telephonic support provided and resolve the issue. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFBG2414719664 | 2025-01-07 | 2025-01-10 | Cust 592 | BANGALORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | MPA not working. | Telephonic support provided for the solution. | Telephonic support provided for the solution. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFBG2415158665 | 2025-01-16 | 2025-01-17 | Cust 603 | BANGALORE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera connectivity needs to be done. | as of now we are closing this complaint and will reopen it once the camera ip is playing in custiomers network | as of now we are closing this complaint and will reopen it once the camera ip is playing in custiomers network | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFBG2415250666 | 2025-01-17 | 2025-01-19 | Cust 592 | BANGALORE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC visit to be done. | AMC visit done. | AMC visit done | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2415257667 | 2025-01-17 | 2025-01-19 | Cust 660 | BANGALORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | Commissioning support to be provided. | Telephonic support provided over the phone for the commissioning of the instrument. | Telephonic support provided over the phone for the commissioning of the instrument. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFBG2415269668 | 2025-01-18 | 2025-01-19 | Cust 608 | BANGALORE | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Proactive Visit for the MPA to be done. | Proactive Visit for the MPA done. | proactive visit done | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2415547669 | 2025-01-24 | 2025-01-27 | Cust 592 | BANGALORE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB DATA UPLOADED ISSUE | Found internet isuse and same has been resolved by the coutomer | Found internet isuse and same has been resolved by the coutomer | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFBG2415932670 | 2025-01-29 | 2025-01-29 | Cust 602 | BANGALORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning support to be provided for the MPA. | Telephonic support provided for the commissioning of the MPA. | Telephonic support provided for the commissioning of the MPA. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFBG2416095671 | 2025-01-31 | 2025-02-03 | Cust 592 | BANGALORE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | ETP PARAMETERS not updating | there was a network issue from the customer side and the same has been resolved and data is now live. | there was a network issue from the customer side and the same has been resolved and data is now live. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFBG2416105672 | 2025-01-31 | 2025-02-10 | Cust 595 | BANGALORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | FM make pH card(1) | Ph Card of the MPA is not working.Need replacement. | pH card replaced and now its working fine. | commissioning done successfully, | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2416163673 | 2025-01-31 | 2025-01-31 | Cust 600 | BANGALORE | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Installation guidelines of multiparameter analyser. | Installation guidelines of multiparameter analyser done. | Installation guidelines of multiparameter analyser done. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFBG2416350674 | 2025-02-04 | 2025-02-10 | Cust 595 | BANGALORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | MPA analyzer pH card not working. | Ph card replaced and now its working fine. | problem attended in resolved the issue. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2416392675 | 2025-02-05 | 2025-02-07 | Cust 593 | BANGALORE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software re-installation + Local, CPCB and KSPCB configuration | We will reopen this complaint once we get access to the system. | We will reopen this complaint once we get access to the system. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFBG2416545676 | 2025-02-10 | 2025-02-13 | Cust 593 | BANGALORE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software re-installation and Local, CPCB, KSPCB configuration | Software re-installation and configuration has been completed. Data is now live on the Local, CPCB and KSPCB servers. | Software re-installation and configuration has been completed. Data is now live on the Local, CPCB and KSPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFBG2416556677 | 2025-02-10 | 2025-02-14 | Cust 601 | BANGALORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | MPA not working. | Telephonic support provided for the Troubleshooting. | Telephonic support provided for the Troubleshooting. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFBG2416624678 | 2025-02-11 | 2025-02-14 | Cust 621 | BANGALORE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Calibration to be done. | Calibration done. | calibration done successfully. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2417505679 | 2025-02-28 | 2025-03-03 | Cust 661 | BANGALORE | Steam System Division | CLOSED | QUALITY PROBLEM | PRV - MEASURETROL | FM MEASURETROL(SIZE 100NB) | Measuretrol malfunctioning | Found PD1 Sensor failed. Change the PD1. | Reference site visit by R&amp;amp;D Engineer and replacement of components | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFBG2417506680 | 2025-02-28 | 2025-03-03 | Cust 661 | BANGALORE | Steam System Division | CLOSED | QUALITY PROBLEM | PRV - MEASURETROL | FM MEASURETROL(SIZE 080NB) | Measuretrol malfunctioning | Pilot valve block response was slow while exhausting the air. Changed the pilot valve block assembly. | Reference site visit by R&amp;amp;D Engineer and replacement of components | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFBG2417508681 | 2025-02-28 | 2025-03-03 | Cust 661 | BANGALORE | Steam System Division | CLOSED | QUALITY PROBLEM | PRV - MEASURETROL | FM MEASURETROL(SIZE 050NB) | Measuretrol malfunctioning | Found P1 section was not working on MTrol PCB. Changed the PCB. | Reference site visit by R&amp;amp;D Engineer and replacement of components | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFBG2417509682 | 2025-02-28 | 2025-03-03 | Cust 661 | BANGALORE | Steam System Division | CLOSED | QUALITY PROBLEM | PRV - MEASURETROL | FM MEASURETROL(SIZE 100NB) | measuretrol malfunctioing | Found PD1 Sensor failed. Change the PD1. | Reference site visit by R&amp;amp;D Engineer and replacement of components | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFBG2417580683 | 2025-02-28 | 2025-03-01 | Cust 610 | BANGALORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | PH input card(1) | Commissioning support to be provided. | Telephonic support provided for the commissioning. | Telephonic support provided for the commissioning. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFBG2418201684 | 2025-03-14 | 2025-03-17 | Cust 650 | BANGALORE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Local server is not accessible. | as of now we are closing this complaint and will reopen it once we get system access. | as of now we are closing this complaint and will reopen it once we get system access. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFBG2418209685 | 2025-03-14 | 2025-03-17 | Cust 593 | BANGALORE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to CPCB server. | Data is now showing live on the CPCB and KSPCB servers. Kindly check from your end once. | Data is now showing live on the CPCB and KSPCB servers. Kindly check from your end once. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFBG2418274686 | 2025-03-17 | 2025-03-25 | Cust 659 | BANGALORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Chlorine sensor | CHLORINE SENSOR(1) | Chlorine sensor FMRCS X1 reads always minimum value, not responding to dosing. Sensor swapped and confirmed issue with the sensor. Other sensor reads 6.5PPM. | Replacement will be arranged as both the sensor not responding. | Problem attended and solution convey to the customer. | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2418620687 | 2025-03-24 | 2025-03-28 | Cust 601 | BANGALORE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC Visit to be done. | AMC visit done. | AMC visit done | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2419095688 | 2025-03-31 | 2025-03-31 | Cust 662 | BANGALORE | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | INSTALLATION GUIDELINES FOR MULTI PARAMETER ANALYSER SUPPLIED FOR ETP | VISITED TH SITE AND PROVIDED THE INSTALLATION GUIDELINES | CLOSED | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFBG2500304689 | 2025-04-08 | 2025-04-11 | Cust 659 | BANGALORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Chlorine sensor | CHLORINE SENSOR(1) | Chlorine sensor not working. | Chlorine sensor replaced but still readings are not matching with the customer portable analyzer. Hence DPD kit is called for verification of it. | Problem attended and solution convey to the customer. | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2500305690 | 2025-04-08 | 2025-04-14 | Cust 615 | BANGALORE | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | MPA Proactive visit to be done. | Proactive visit done. | proactive visit done | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2500524691 | 2025-04-14 | 2025-04-18 | Cust 639 | BANGALORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Commissioning support to be provided. | Commissioning done. | commissioning done successfully, | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2501388692 | 2025-04-30 | 2025-04-30 | Cust 662 | BANGALORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | commissioning of cx2000 multiparameter installed at ETP outlet | visited the site and completed the commissioning and found the values inline with lab values | closed | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFBG2501392693 | 2025-04-30 | 2025-04-30 | Cust 600 | BANGALORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | commissioning of newly supplied multiparameter cx2000 installed at etp outlet line | visited the site and completed the commissioning | closed | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFBG2501492694 | 2025-05-02 | 2025-05-03 | Cust 625 | BANGALORE | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Proactive visit to be done at site for the Spares procurement. | Proactive visit done. | proactive visit done | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2501559695 | 2025-05-05 | 2025-05-07 | Cust 659 | BANGALORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Chlorine sensor | CHLORINE SENSOR(1) | Chlorine sensor not working. | Problem attended and its sensor not working. | Problem attended and solution convey to the customer. | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2501711696 | 2025-05-07 | 2025-05-10 | Cust 609 | BANGALORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | Commissioning to be done. | Commissioning done. | problem attended in resolved the issue. | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2501739697 | 2025-05-08 | 2025-05-09 | Cust 593 | BANGALORE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB Server login issue. | Now the login issue is resolved. | Now the login issue is resolved. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFBG2502018698 | 2025-05-15 | 2025-05-17 | Cust 625 | BANGALORE | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Proactive visit to be done. | proactive visit done. | proactive visit done | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2503233699 | 2025-06-08 | 2025-06-12 | Cust 609 | BANGALORE | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | OIL ANALYSER | UV Lamp box(1) | Oil In water analyzer performance to be check. | Proactive visit for the same done. | proactive visit done | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFBG2503606700 | 2025-06-17 |  | Cust 663 | BANGALORE | Analytical - CHAKAN | ASSIGNED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | MPA commissioning |  |  | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFCB2303656701 | 2023-06-08 | 2023-06-08 | Cust 664 | COIMBATORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4Trans(1) | Chlorine analyser commissioning | Chlorine analyser commissioned and found matching with lab | Chlorine analyser commissioned and found matching with lab | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2304307702 | 2023-06-21 | 2023-06-23 | Cust 665 | COIMBATORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning | commissioning of the analyser done and also verified the readings against STD solutions found matching. | commissioning of the analyser done and also verified the readings against STD solutions found matching. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2304673703 | 2023-06-27 | 2023-06-27 | Cust 665 | COIMBATORE | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Pre commissioning | Located the location of the analyser and necessary instruction given for installation. | Located the location of the analyser and necessary instruction given for installation. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2305143704 | 2023-07-05 | 2023-07-07 | Cust 666 | COIMBATORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | AQUADO Analyzer | DO2 Arc Sensors(DO2 Arc Sensors) | Error current | Found error in electronic suggested to send it to pune | Found error in electronic suggested to send it to pune | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2305188705 | 2023-07-06 | 2023-07-24 | Cust 667 | COIMBATORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Display not working | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210054399 LR Dt - 12-JUL-2023 | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210054399 LR Dt - 12-JUL-2023 | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFCB2306295706 | 2023-07-27 | 2023-07-31 | Cust 668 | COIMBATORE | Analytical - CHAKAN | CLOSED | DIGITAL CONEECTIVITY | Aqua4Trans | RS485 card- of Aqua4Trans(1) | Installation/commissioning of RS 485 card for the pH, TDS and Turbidity sensor | Completed the commissioning of RS 485 card | Completed the commissioning of RS 485 card | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFCB2306321707 | 2023-07-28 | 2023-07-28 | Cust 665 | COIMBATORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning | Commissioning of the analyser done and found working Ok | Commissioning of the analyser done and found working Ok | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2306769708 | 2023-08-03 | 2023-08-04 | Cust 669 | COIMBATORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Electrode | FM1106(1) | commissioning | Commissioning of pH transmitter done and verified it against std found ok | Commissioning of pH transmitter done and verified it against std found ok | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2307080709 | 2023-08-09 | 2023-08-10 | Cust 670 | COIMBATORE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 2nd AMC visit | Did calibration of analyser and Found ok | Did calibration of analyser and Found ok | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFCB2308407710 | 2023-09-04 | 2023-09-10 | Cust 671 | COIMBATORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | TOC ANALYSER | Main PCB(017?016) | AMC visit of TOC analysers | TOC analysers checked and found some of the analyser Display faulty | TOC analysers checked and found some of the analyser Display faulty | NO | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2308819711 | 2023-09-13 | 2023-09-15 | Cust 672 | COIMBATORE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration of MPA | Calibration of Forbes Marshall instruments done and found within accepted tolerance | Calibration of Forbes Marshall instruments done and found within accepted tolerance | NO | Paid | Core | N |  | Analytical - CHAKAN
AFCB2309324712 | 2023-09-25 | 2023-09-29 | Cust 671 | COIMBATORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | TOC ANALYSER | FM 101 TOC Analyser(FM 101 TOC) | Display Blank | Checked and found keypad faullty | Checked and found keypad faullty | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2310317713 | 2023-10-12 | 2023-10-14 | Cust 665 | COIMBATORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning of MPA CX2000 | Analyser commissioned, working ok. | MPA commissioned | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2311209714 | 2023-10-30 | 2023-11-02 | Cust 673 | COIMBATORE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 3rd AMC Visit | AMC of the analyser carried out and also calibration done found working within accepted tolerance. | AMC of the analyser carried out and also calibration done found working within accepted tolerance. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFCB2311431715 | 2023-11-01 | 2023-11-02 | Cust 674 | COIMBATORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | DO Transmitters | VISI pro DO(VISI pro DO) | Healthiness Checking | Checked the analyser in Na2so3 and found 0.2ppm also checked the anlayser in air found 5.3ppm.carried out air calibration. Now analyser reading matching with lab reading | Checked the analyser in Na2so3 and found 0.2ppm also checked the anlayser in air found 5.3ppm.carried out air calibration. Now analyser reading matching with lab reading | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2311678716 | 2023-11-07 | 2023-11-07 | Cust 675 | COIMBATORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Turbidity Meter | Optex turbidity sensor(1) | FMTC-100 MI-100 reading always zero | Offset found programmed in the transmitter, same was removed, cross checked with different samples and found ok | Issues resolved | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2313149717 | 2023-12-05 | 2023-12-08 | Cust 676 | COIMBATORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Electrode | FM1104(1) | Not responding to any buffer | pH electrode found contaminated replaced the electrolyte and tried regeneration still the electrode is not responding | pH electrode found contaminated replaced the electrolyte and tried regeneration still the electrode is not responding | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2313476718 | 2023-12-12 | 2023-12-15 | Cust 673 | COIMBATORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Display blank | Found PCX1 and display board faulty need to be replaced | Found PCX1 and display board faulty need to be replaced | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFCB2313645719 | 2023-12-15 | 2023-12-15 | Cust 677 | COIMBATORE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Smartpro Electronics Assembly(Smartpro-Elc-cbl) | Annual calibration of pH meters | one electrode not responding. Another electrode ok. Same was used for both the pH transmitters, calibrated and found ok. | pH analysers calibrated | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFCB2314167720 | 2023-12-25 | 2023-12-27 | Cust 673 | COIMBATORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | LCD of Multiparameter Display(1) | Commissioning of the Spare PCX1 with LCD screen | Commissioning of the Spare PCX1 with LCD screen done and configuration carried out ,also 4th AMC visit completer. Found analyzer working Ok | Commissioning of the Spare PCX1 with LCD screen done and configuration carried out ,also 4th AMC visit completer. Found analyzer working Ok | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFCB2314481721 | 2023-12-30 | 2024-01-02 | Cust 678 | COIMBATORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Aqua4Trans | Aqua4Trans(1) | ORP sensor not responding to std | Checked and found sluggish response of sensor suggested them to replace the sensor. Transmitter found ok | Checked and found sluggish response of sensor suggested them to replace the sensor. Transmitter found ok | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2315041722 | 2024-01-08 | 2024-01-11 | Cust 644 | COIMBATORE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Preparation of standards and calibration of MPA | Calibrations of the MPA done found ok | Calibrations of the MPA done found ok | NO | Paid | Core | N |  | Analytical - CHAKAN
AFCB2315897723 | 2024-01-24 | 2024-01-27 | Cust 636 | COIMBATORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Chlorine Analyzer-MAC 100 | MAC 100(1) | Commissioning of chlorine sensor | Commissioning of the sensor done and found matching with lab | Commissioning of the sensor done and found matching with lab | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2316248724 | 2024-01-31 | 2024-02-01 | Cust 666 | COIMBATORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4Trans(1) | Commissioning of TSS & PH sensor | Mr. Sasidharan guided over phone, commissioning done, analysers working okay. | Mr. Sasidharan guided over phone, commissioning done, analysers working okay. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFCB2316705725 | 2024-02-07 | 2024-02-09 | Cust 679 | COIMBATORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | commissioning | Mr. Sasidharan guided the customer over phone, commissioned the analysers, flow meter, RTU installed in Abu Dhabi, found working ok | Mr. Sasidharan guided the customer over phone, commissioned the analysers, flow meter, RTU installed in Abu Dhabi, found working ok | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFCB2316789726 | 2024-02-09 | 2024-02-10 | Cust 617 | COIMBATORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | THERMO SERIES | ADAPTER(1) | Commissioning of spare Thermo scientific Hardness Orion 2120XP analyser | Commissioning of the analyser done and calibrated it with STD found ok | Commissioning of the analyser done and calibrated it with STD found ok | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2316979727 | 2024-02-13 | 2024-02-14 | Cust 680 | COIMBATORE | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Smartpro-con) | INSTALLATION GUIDELINE | DISCUSSED ABOUT LOCATION AND ELECTRICAL WIRING | DISCUSSED ABOUT LOCATION AND ELECTRICAL WIRING | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFCB2317430728 | 2024-02-22 | 2024-02-23 | Cust 681 | COIMBATORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | Flowcell disc of Multiparameter(1) | Not responding | Did alignment of the detector and carried out zero calibration and STD calibration found analyser working ok | Did alignment of the detector and carried out zero calibration and STD calibration found analyser working ok | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2317506729 | 2024-02-23 | 2024-02-26 | Cust 673 | COIMBATORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning of PCX1 after service | PCX1 board changed and did configuration of the MPA now analyser working OK | PCX1 board changed and did configuration of the MPA now analyser working OK | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2318207730 | 2024-03-05 | 2024-03-06 | Cust 665 | COIMBATORE | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning of pH dipstick | PH diptisk installed and checked the response with std solutions found ok | PH diptisk installed and checked the response with std solutions found ok | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2318840731 | 2024-03-18 | 2024-03-22 | Cust 682 | COIMBATORE | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | Conductivity Transmitter Aquacon | Conductivity Transmitter Aquamon 4000(COND-Trans Aqu-4000) | Guidelines for the installation , checking of transmitter as well and the configuration of relay outputs. | Given the Guidelines for the installation , also checked the transmitter as well and done the configuration of relay outputs. | Given the Guidelines for the installation , also checked the transmitter as well and done the configuration of relay outputs. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFCB2400123732 | 2024-04-03 | 2024-04-04 | Cust 673 | COIMBATORE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC VISIT | First AMC visit completed, analyser working fine. | First Visit AMC completed | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFCB2400182733 | 2024-04-04 | 2024-04-06 | Cust 679 | COIMBATORE | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Pre-commissioning and training on CX2000 MPA - 2nos. | Training on cx2000 given to customer | Training on cx2000 given to customer | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2400182733 | 2024-04-04 | 2024-04-06 | Cust 679 | COIMBATORE | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Pre-commissioning and training on CX2000 MPA - 2nos. | Training on cx2000 given to customer | Training on cx2000 given to customer | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2400373734 | 2024-04-09 | 2024-04-10 | Cust 683 | COIMBATORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4Trans(1) | Commissioning | Commissioning of the TDS meter done and calibrated against the STD found ok | Commissioning of the TDS meter done and calibrated against the STD found ok | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2402295735 | 2024-05-20 | 2024-05-24 | Cust 684 | COIMBATORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | VISI pro DO(VISI pro DO) | Commissioning | DO analyser fixed and checked the readings found working ok | DO analyser fixed and checked the readings found working ok | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2402354736 | 2024-05-20 | 2024-06-06 | Cust 685 | COIMBATORE | Analytical - CHAKAN | CLOSED | PRODUCT DESIGN FAILURE | Aqua4Trans | Communication card of Aqua4Trans(1) | Modbus output hanging | material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210100632 LR Dt - 30-MAY-2024 | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210100632 LR Dt - 30-MAY-2024 | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFCB2402385737 | 2024-05-20 | 2024-05-24 | Cust 646 | COIMBATORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4Trans(1) | pH transmitter commissioning | pH transmitter commissioning done and verified the pH analyser against std found ok | pH transmitter commissioning done and verified the pH analyser against std found ok | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2403223738 | 2024-06-03 | 2024-06-05 | Cust 671 | COIMBATORE | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | Aqua4Trans | Aqua4Trans(1) | Discussion with contractor for erection of analyser | discussed with contractor and fixed the analyser and also guided them to mount four more TSS analyser. Sample tubing and drain syytem confirmed. Also discussed power supply cable termination | discussed with contractor and fixed the analyser and also guided them to mount four more TSS analyser | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2403271739 | 2024-06-04 | 2024-06-07 | Cust 646 | COIMBATORE | Swas - CHAKAN | CLOSED | AMC | SWAS & Analysers | Steam & Water Analysis System(SWAS & Analysers) | 2nd AMC visit | 2md AMC visit completed successfully | AMC completed | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFCB2403342740 | 2024-06-06 | 2024-06-07 | Cust 686 | COIMBATORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Conductivity Tranmitter Smartpro | Smartpro Electronics Assembly(Smartpro-ele-con) | PH not responding | pH transmitter housing found faulty suggested to send to Pune for necessary repairs | pH transmitter housing found faulty suggested to send to Pune for necessary repairs | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2403833741 | 2024-06-17 | 2024-06-23 | Cust 671 | COIMBATORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | Commissioning of OIW, Turbidity,H2S and Ammonia analyser | Commissioning of the analyser is going on verification and SAT has to be done. | Commissioning of the analyser is going on verification and SAT has to be done. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2404114742 | 2024-06-23 | 2024-06-28 | Cust 671 | COIMBATORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | commissioning of oil in water H2S Ammonia analyser | Sulphide Analyser - 2nos., Ammonia Analyser - 1no., Oil in Water analyser - 1no., TSS Analyser - 6nos. and pH -1no., commissioned, working ok. | ETP V Analysers commissioned | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2404719743 | 2024-07-01 | 2024-07-03 | Cust 673 | COIMBATORE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC Visit | AMC visit completed and calibrated the analyser with std found ok | AMC visit completed and calibrated the analyser with std found ok | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFCB2405057744 | 2024-07-08 | 2024-07-10 | Cust 671 | COIMBATORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | Training of the analysers to End user and completing SAT Procedure | Hands on Training given to Engineers from BPCL and also found analyser working ok | Hands on Training given to Engineers from BPCL and also found analyser working ok | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2405452745 | 2024-07-16 | 2024-07-19 | Cust 687 | COIMBATORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | VISI pro DO(VISI pro DO) | Commissioning of DO analyser | Guided customer over phone, analyser commissioned. | Guided customer over phone, analyser commissioned | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFCB2405453746 | 2024-07-16 | 2024-07-19 | Cust 665 | COIMBATORE | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Healthiness checking of MPA | Powered On the analyser after shutdown and checked the healthiness of the anlayser found ok | Powered On the analyser after shutdown and checked the healthiness of the anlayser found ok | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2405914747 | 2024-07-25 | 2024-07-26 | Cust 688 | COIMBATORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Smartpro-con) | Commissioning of TDS analyser after repairs | Sasidharan guided Mr. Ramesh, commissioning done, working ok. | Sasidharan guided Mr. Ramesh, commissioning done, working ok. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFCB2406148748 | 2024-07-28 | 2024-07-29 | Cust 689 | COIMBATORE | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | THERMO SERIES | AV38 TRANSMITTER(1) | Fluoride analyser - Spares management and maintenance discussions | Spares/consumables required for Fluoride Analyser Thermo Orion 2109XP, FM Fluoride 1604440 has been identified and minuted with customer | Issues identified | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2406402749 | 2024-07-31 | 2024-08-01 | Cust 690 | COIMBATORE | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | Conductivity Transmitter Aquacon | Conductivity Transmitter Aquamon 4000(Cond-Trans-4000) | TDS value mismatch | Done calibration found working fine. | Done calibration found working fine. | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFCB2406678750 | 2024-08-05 | 2024-08-07 | Cust 691 | COIMBATORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4Trans(1) | Commissioning of the pH and TDS analyser | PH and TDS analyser commissioning done and found working ok | PH and TDS analyser commissioning done and found working ok | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2406788751 | 2024-08-07 | 2024-08-11 | Cust 671 | COIMBATORE | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | OIL ANALYSER | Flow cell(1) | Monthly visit for checking the healthiness of ETP5 analysers | Checked the healthiness of analyser and found ok | Checked the healthiness of analyser and found ok | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2407271752 | 2024-08-19 | 2024-08-20 | Cust 644 | COIMBATORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Only cleaning cycle happening continuously, key pad not working | Found analyser hanging .Factory reset done and configured the analyser back .Checked the performance of analyser for 3 hour and found ok | Found analyser hanging .Factory reset done and configured the analyser back .Checked the performance of analyser for 3 hour and found ok | NO | Paid | Core | N |  | Analytical - CHAKAN
AFCB2407431753 | 2024-08-21 | 2024-08-24 | Cust 659 | COIMBATORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Chlorine sensor | CHLORINE SENSOR(1) | commissioning | Commissioning of analyser done and found ok | Commissioning of analyser done and found ok | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2407643754 | 2024-08-25 | 2024-08-31 | Cust 692 | COIMBATORE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Chlorine sensor | Chlorine sensor-Krohne(1) | sensor showing high value | Sensor found damage need to be change that same. | Sensor found damage need to be change that same. | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFCB2407644755 | 2024-08-25 | 2024-08-31 | Cust 692 | COIMBATORE | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | Aqua4Trans | Aqua4 trans-A4TW033000(1) | DO sensor not responding | DO sensor cap damage need to change that same. | DO sensor cap damage need to change that same. | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFCB2407847756 | 2024-08-28 | 2024-08-30 | Cust 693 | COIMBATORE | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | Aqua4Trans | Aqua4Trans(1) | Pre-commissioning activities for pH, Cond., Cl2, ORP at Pillur III project | Pre-commissioning visit made, suggestions given | Pre commissioning and training given | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2408406757 | 2024-09-05 | 2024-11-04 | Cust 694 | COIMBATORE | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | Chlorine sensor | CHLORINE SENSOR(1) | COMMISSIONING OF CHLROINE ANALYSER TRANSMITTER NOT RESPONDING | Replacement of transmitter and sensor sent to site. Issue is resolved | Replacement of transmitter and sensor sent to site. Issue is resolved | YES | No Trip | Core | Y | Product / equipment failure due to which customers plant or part of plant is stopped which leaads to production loss. | Analytical - CHAKAN
AFCB2408540758 | 2024-09-09 | 2024-09-11 | Cust 671 | COIMBATORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | OIL ANALYSER | Flow cell(1) | Monthly visit | Monthly visit done and checked the healthiness of all FM analyser and found sample pump of oil in water is faullty | Monthly visit done and checked the healthiness of all FM analyser and found sample pump of oil in water is faullty | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2409186759 | 2024-09-23 | 2024-09-24 | Cust 671 | COIMBATORE | Analytical - CHAKAN | CLOSED | APPLICATION ISSUE | SULPHIDE ANALYSER | Metering pump : 24 V 50 Hz(08810=A=7150) | Sample pump not working | Poor quality sample - issues discussed with customer | Issues identified | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2409275760 | 2024-09-24 | 2024-09-26 | Cust 691 | COIMBATORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4Trans(1) | Commissioning | Commissioning of the analysers done and found the reading of TDS and pH were matching the lab | Commissioning of the analysers done and found the reading of TDS and pH were matching the lab | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2409296761 | 2024-09-24 | 2024-09-27 | Cust 695 | COIMBATORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Blurred display | Backlight gone, analyser remains in reset initial mode. Issue found with R1 and R2 resistance value which was around 200Kohms instead of 470Kohms in PCX1. Also memory found erased automatically, if the power supply is interrupted. 3V battery in PCX1 is ok. Suggested to repair the PCX1 board or can run the analyser without power interruptions. Reconfigurations done, analyser working okay now. | Issues identified | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2409320762 | 2024-09-25 | 2024-09-26 | Cust 670 | COIMBATORE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | Flow cell for Multiparameter(AWA-FLOWCELL) | AMC visit | AMC visit done and calibrated the analyser aganist the std found ok | AMC visit done and calibrated the analyser aganist the std found ok | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFCB2409328763 | 2024-09-25 | 2024-09-27 | Cust 676 | COIMBATORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Turbidity Meter | Optex turbidity sensor(1) | Not responding | As per the images received from the site, scratches found in the optics leading to higher values even with distilled water. Suggested to send the analyser to HO for repairs/replacements. | Suggested to send to HO for necessary repairs | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFCB2409838764 | 2024-10-03 | 2024-10-05 | Cust 646 | COIMBATORE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration and functional checking of MPA | Cleaning of flow cell carried out physically and with 5% sulphuric acid cleaning solution. Sample pump neoprene tube reversed. Acid pump neoprene tube reversed. Zero calibration done with cleaning solution. Acid and sample tubes found physically damaged/chocked at many places, which was replaced with available tubes provided by M/s. CSL Response of the analyser checked for different samples and with DM water, found ok. Later analyser checked with standard solution of 100PPM & 500PPM COD solutions and calibrated. Reading after calibration found to be 89PPM & 500PPM. Present detector values with cleaning acid are PH1 ? 752 and PH2 ? 2890. Since PH1 detector is less by 500 than the expected value of 1250, FMPL engineer suggested to maintain above spare for trouble free operations. Later analyser taken online and now the readings are as follows COD ? 9.86PPM, BOD ? 4.93PPM mA checked and found ok. Necessary training on the analyser provided by FMPL engineer. | COD/BOD Analyser calibrated | NO | Paid | Core | N |  | Analytical - CHAKAN
AFCB2409861765 | 2024-10-04 | 2024-10-05 | Cust 671 | COIMBATORE | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | OIL ANALYSER | Peristaltic sampling pump motor(PMP-SAMP-Head) | Pro active visit | Changed the new sample pump took trial run found ok also checked load amps found around 0.45A for both treated & equalization tank | Changed the new sample pump took trial run found ok also checked load amps found around 0.45A for both treated & equalization tank | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2410014766 | 2024-10-08 | 2024-10-09 | Cust 696 | COIMBATORE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Annual Calibration | Calibration of the forbes marshall instruments done and found working ok | Calibration of the forbes marshall instruments done and found working ok | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFCB2410125767 | 2024-10-10 | 2024-10-11 | Cust 697 | COIMBATORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | AQUAMAX DOT(1) | Commissioning | Commissioning of the analyser done and found working ok | Commissioning of the analyser done and found working ok | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2410708768 | 2024-10-22 | 2024-10-23 | Cust 698 | COIMBATORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Auto reset | Found PCX1 components faulty took PCX1 board with Display for service | Found PCX1 components faulty took PCX1 board with Display for service | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2410870769 | 2024-10-25 | 2024-11-11 | Cust 681 | COIMBATORE | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | Aqua4Trans | Aqua4 trans-A4TW033000(1) | LCD screen Flickering | MARK EXPRESS PRIVATE LIMITED vide LR No - 210119664 LR Dt - 07-NOV-2024 | Replacement sent through MARK EXPRESS PRIVATE LIMITED vide LR No - 210119664 LR Dt - 07-NOV-2024 | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFCB2410870769 | 2024-10-25 | 2024-11-11 | Cust 681 | COIMBATORE | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | Aqua4Trans | Aqua4 trans-A4TW033000(1) | LCD screen Flickering | MARK EXPRESS PRIVATE LIMITED vide LR No - 210119664 LR Dt - 07-NOV-2024 | Replacement sent through MARK EXPRESS PRIVATE LIMITED vide LR No - 210119664 LR Dt - 07-NOV-2024 | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFCB2411811770 | 2024-11-15 | 2024-11-18 | Cust 671 | COIMBATORE | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | OIL ANALYSER | Peristaltic sampling pump motor(PMP-SAMP-Head) | Monthly visit | Healthiness of forbes marshall analyser checked and monthly PM work done | Healthiness of forbes marshall analyser checked and monthly PM work done | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2411866771 | 2024-11-16 | 2024-11-19 | Cust 699 | COIMBATORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Electrode | Ph Electrode(FM1104(4C)) | Calibration status was failure. ph electrode issue | ph electrode not able to calibrate, need to purchase the new sensor. | ph electrode not able to calibrate, need to purchase the new sensor. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFCB2411955772 | 2024-11-18 | 2024-11-19 | Cust 664 | COIMBATORE | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | Chlorine sensor | CHLORINE SENSOR(1) | To check the analysers - 8nos. and flow meters - 10nos. | Discussed on maintaining spares, regular calibrations of analysers. Offers also submitted | Pro active visit completed | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2411977773 | 2024-11-19 | 2024-11-21 | Cust 669 | COIMBATORE | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Smartpro-con) | Proactive visit | Pro active visit done and discussion for krohne flow meter calibration done | Pro active visit done and discussion for krohne flow meter calibration done | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2411977773 | 2024-11-19 | 2024-11-21 | Cust 669 | COIMBATORE | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | Pro active visit | Pro active visit done and discussion for krohne flow meter calibration done | Pro active visit done and discussion for krohne flow meter calibration done | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2412019774 | 2024-11-19 | 2024-11-20 | Cust 681 | COIMBATORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Aqua4Trans | Display card of Aqua4Trans(1) | Dark spots in the display. | Display replaced with new one received from the division, working ok. Reconfigurations done, calibrated, mA checked and found ok. ORP - 1no. and cond. - 1no. | Issues resolved | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2412321775 | 2024-11-25 | 2024-11-30 | Cust 700 | COIMBATORE | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | DO Transmitters | VISI pro DO(VISI pro DO) | Installation guidelines for sensor | Given the installation guidelines | Given the installation guidelines. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFCB2412617776 | 2024-11-30 | 2024-11-30 | Cust 700 | COIMBATORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | Commissioning | Completed the commissioning | Completed the commissioning | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFCB2413019777 | 2024-12-06 | 2024-12-07 | Cust 489 | COIMBATORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | commissioning of the analyser | Commissioning of Forbesmarshall analyser and flowmeter done found working ok | Commissioning of Forbesmarshall analyser and flowmeter done found working ok | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2413335778 | 2024-12-12 | 2024-12-19 | Cust 693 | COIMBATORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4 trans-A4TW033000(1) | Commissioning of ORP, pH, Cl2 and Cond analysers for Raw water and clear water streams | Raw Water Analyser Panel ? 4 StreamsAqua4trans.[Model A4TW015000, S no. 24A4TQ024-W015000 and MODEL A4TW023000, S. NO. 24A4TQ027-W023000]Both transmitters powered on, necessary configurations done and all analysers found working ok..pH and ORP calibrated with buffer solutions brought by FMPL engineer.Process values after calibration being pH ? 7.78, Cond. - 131.1 Micro Siemens/cm, ORP - +557mV & Chlorine ? 1.695PPMTreated Water Analyser Panel ? 4 Streams .Aqua4trans.[Model A4TW015000, S no. 24A4TQ023-W015000 and MODEL A4TW023000, S. NO. 24A4TQ026-W023000]Both transmitters powered on, necessary configurations done and all analysers found working ok.mA checked and found matching.Process values after calibration being pH ? 7.42, Cond. - 130.9 Micro Siemens/cm, ORP - +474mV & Chlorine ? 2.74PPMBoth location analysers ? 8nos. Commissioned and it is working ok. | Analysers commissioned | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2413632779 | 2024-12-19 | 2025-02-10 | Cust 693 | COIMBATORE | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | Conductivity Electrode | Forbes Marshall Cell Model 8311, K=0.1(FM -8311) | Conductivity cable M4E-CCABLE3M TERMINATION NOT PROPER AT SENSOR END AS SCREWS WERE TIGHTENED TO THE INSULATION AND NOT TO THE CORE LEADING TO OPEN VALUE OF CONDUCTIVITY AT A LOCATION AND TEMPERATURE SENSOR OPEN AT OTHER LOCATION. PIS ATTACHED | Necessary corrective action taken. Vendor instructed to do proper lugging & termination. Change implemented in the next lot & verified on receipt at Chakan. | Necessary corrective action taken. Vendor instructed to do proper lugging &amp;amp; termination. Change implemented in the next lot &amp;amp; verified on receipt at Chakan. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFCB2413632779 | 2024-12-19 | 2025-02-10 | Cust 693 | COIMBATORE | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | Conductivity Electrode | Forbes Marshall Cell Model 8311, K=0.1(FM -8311) | Conductivity cable M4E-CCABLE3M TERMINATION NOT PROPER AT SENSOR END AS SCREWS WERE TIGHTENED TO THE INSULATION AND NOT TO THE CORE LEADING TO OPEN VALUE OF CONDUCTIVITY AT A LOCATION AND TEMPERATURE SENSOR OPEN AT OTHER LOCATION. PIS ATTACHED | Necessary corrective action taken. Vendor instructed to do proper lugging & termination. Change implemented in the next lot & verified on receipt at Chakan. | Necessary corrective action taken. Vendor instructed to do proper lugging &amp;amp; termination. Change implemented in the next lot &amp;amp; verified on receipt at Chakan. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFCB2413942780 | 2024-12-25 | 2024-12-26 | Cust 701 | COIMBATORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | AQUADO Analyzer | Visitrace DO sensor(1) | Cap failure | Found Red channel failure and also cap missing error .Took sensor for further checking | Found Red channel failure and also cap missing error .Took sensor for further checking | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2413964781 | 2024-12-26 | 2024-12-27 | Cust 702 | COIMBATORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Electrode | FM1106(1) | Not responding to buffer | Found smartpro transmitter failed due to moisture ingress replaced with spare one and checked found ok | Found smartpro transmitter failed due to moisture ingress replaced with spare one and checked found ok | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2414074782 | 2024-12-28 | 2024-12-28 | Cust 691 | COIMBATORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Electrode | FM1106(1) | Not responding to buffer | Checked the electrode with std found not responding to buffer took it for regeneration | Checked the electrode with std found not responding to buffer took it for regeneration | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2414506783 | 2025-01-01 | 2025-01-02 | Cust 677 | COIMBATORE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | ANNUAL CALIBRATION OF PH ANALYSERS | Calibration of pH analyser done and found ok | Calibration of pH analyser done and found ok | NO | Paid | Core | N |  | Analytical - CHAKAN
AFCB2414540784 | 2025-01-02 | 2025-01-03 | Cust 665 | COIMBATORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | AQUAMAX DOT(1) | Commissioning | Commissioning done and found analyser working Ok | Commissioning done and found analyser working Ok | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2415080785 | 2025-01-14 | 2025-01-17 | Cust 687 | COIMBATORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | AQUADO Analyzer | DO2 Arc Sensors(DO2 Arc Sensors) | Always shoeing zero | Found temp range out , took sensor for further analysis | Found temp range out , took sensor for further analysis | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2415207786 | 2025-01-17 | 2025-01-17 | Cust 703 | COIMBATORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | CPU with LCD of Multiparameter Display(1) | Installation of pcx after service | PCX board installed necessary configuration done found working ok | PCX board installed necessary configuration done found working ok | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2415541787 | 2025-01-24 | 2025-01-27 | Cust 671 | COIMBATORE | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | OIL ANALYSER | Flow cell(1) | monthly visit | Monthly visit done and found analyser working fine | Monthly visit done and found analyser working fine | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2415710788 | 2025-01-27 | 2025-01-29 | Cust 704 | COIMBATORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4Trans(1) | Commissioning | Commissioning of pH, TDS analyser, 3 mags, VFM, CV done. | Analysers commissioned | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2415775789 | 2025-01-28 | 2025-01-29 | Cust 670 | COIMBATORE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC visit | Cleaning of flow cell, routine maintenance and calibration of MPA carried out, analyser found working ok. | Fourth AMC visit completed | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFCB2416342790 | 2025-02-04 | 2025-02-08 | Cust 644 | COIMBATORE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Routine maintenance and calibration of MPA | Analyser checked and found having both photo detectors low, MUVX card issue, auto resetting, not responding to the standards. Suggested to upgarde to CX2000 | Analyser checked and found ok | NO | Paid | Core | N |  | Analytical - CHAKAN
AFCB2416343791 | 2025-02-04 | 2025-02-08 | Cust 671 | COIMBATORE | Analytical - CHAKAN | CLOSED | MANUFACTURING ISSUES | TOC ANALYSER | FM 101 TOC Analyser(FM 101 TOC) | Functional check | Checked the healthiness of the analyser and status are attached | Checked the healthiness of the analyser and status are attached | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFCB2416879792 | 2025-02-17 | 2025-03-22 | Cust 668 | COIMBATORE | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | Aqua4Trans | Aqua4Trans(1) | Aqua4Trans display is flickering. | The display requires repair. We will arrange for the necessary repairs to ensure the display is functioning correctly | The display requires repair. We will arrange for the necessary repairs to ensure the display is functioning correctly | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFCB2416879792 | 2025-02-17 | 2025-03-22 | Cust 668 | COIMBATORE | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | Aqua4Trans | Aqua4Trans(1) | Aqua4Trans display is flickering. | The display requires repair. We will arrange for the necessary repairs to ensure the display is functioning correctly | The display requires repair. We will arrange for the necessary repairs to ensure the display is functioning correctly | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFCB2416879792 | 2025-02-17 | 2025-03-22 | Cust 668 | COIMBATORE | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | Aqua4Trans | Aqua4Trans(1) | Aqua4Trans display is flickering. | The display requires repair. We will arrange for the necessary repairs to ensure the display is functioning correctly | The display requires repair. We will arrange for the necessary repairs to ensure the display is functioning correctly | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFCB2416924793 | 2025-02-18 | 2025-02-21 | Cust 691 | COIMBATORE | Analytical - CHAKAN | CLOSED | APPLICATION ISSUE | Ph Electrode | FM1106(1) | Pre-mature failure of pH electrodes | Electrodes found lying in dry condition with coating of dried muds, itching and regeneration tried but electrodes not responding. Suggested to procure new electrodes, maintain it frequently, to increase the measurement chamber size. | Resolutions given | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2416944794 | 2025-02-19 | 2025-02-19 | Cust 646 | COIMBATORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4Trans(1) | Commissioning | Plant not ready | Plant not ready | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2417255795 | 2025-02-26 | 2025-02-28 | Cust 704 | COIMBATORE | Krohne - MIDC | CLOSED | PRODUCT FAILURE | Accessories for KROHNE Products | FIT393M Flow Indicator Totalizer Micro-Controller based(FIT393M) | 7SEGMENT LED ISSUE | Replaced FIT fixed and found working Ok | Replaced FIT fixed and found working Ok | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2417458796 | 2025-02-28 | 2025-03-17 | Cust 679 | COIMBATORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | Sample Pump(1) | Sample Pump not running getting load. | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210134230 LR Dt - 11-MAR-2025 | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210134230 LR Dt - 11-MAR-2025 | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFCB2417571797 | 2025-02-28 | 2025-02-28 | Cust 690 | COIMBATORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | VISI pro DO(VISI pro DO) | Commissioning of DO sensor | Successfully commissioned DO sensor | Successfully commissioned DO sensor | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFCB2417719798 | 2025-03-04 | 2025-03-06 | Cust 673 | COIMBATORE | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Discussion for AMC renewal | Discussions done with inst. dept., purchase dept.PO expected shortly | AMC renewal discussions, Gas analyser installation guidelines | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2418006799 | 2025-03-10 | 2025-03-13 | Cust 665 | COIMBATORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4Trans(1) | Commissioning of Aqua4Trans TDS meter | Commissioning of TDS analyser done and found working ok | Commissioning of TDS analyser done and found working ok | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2418296800 | 2025-03-17 | 2025-03-20 | Cust 671 | COIMBATORE | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | Monthly visit | Monthly maintenance of the analysers done and found working ok | Monthly maintenance of the analysers done and found working ok | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2418506801 | 2025-03-21 | 2025-03-22 | Cust 636 | COIMBATORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Recommissioning and calibration of CX-1000-3912 | 1. Standard chemicals KHP and Kaolin powders brought by FMPL engineers handed over to M/s. Infosys.2. COD standard solutions of value 250mg/l were prepared with the help of M/s. Infosys. 3. TSS standard solutions of value 250mg/l were prepared with the help of M/s. Infosys.4. Cleaning solution [10% H2SO4] were prepared with the help of M/s. Infosys.5. MPA flow cell cleaning carried out.6. Sample pump tube found worn out, which was reversed.7. Cleaning acid pump tube contact area changed for better pumping. Acid air bubble issue resolved by changing the tube.8. Zero calibration carried out with 10%H2SO4 for COD and TSS.9. Photo Detector values PH1 ? 74 and PH2 ? 437 are very less, both detectors needs replacement.10. Key pad not found working properly in some areas, which also needs to be replaced. | MPA commissioned | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2500175802 | 2025-04-06 | 2025-04-09 | Cust 691 | COIMBATORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | VISI pro DO(VISI pro DO) | commissioning of do analyser | Sensor installed and termination done with Process Indicator, also checked the sensor in atmospheric air found working ok(7.5PPM) | Sensor installed and termination done with Process Indicator, also checked the sensor in atmospheric air found working ok(7.5PPM) | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2500273803 | 2025-04-08 | 2025-04-09 | Cust 705 | COIMBATORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Electrode | CONDUCELL 4UHF-PG-120 (4 Hastelloy C electrodes)(HT237 627) | commissioning | Conductivity sensor installed and cross verified the response of sensor with DM water & process water found working ok | Conductivity sensor installed and cross verified the response of sensor with DM water & process water found working ok | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2500566804 | 2025-04-15 | 2025-04-18 | Cust 673 | COIMBATORE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC | Calibrated the analyser against STD and found ok within accepted tolerance. | Calibrated the analyser against STD and found ok within accepted tolerance. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFCB2501637805 | 2025-05-06 | 2025-05-09 | Cust 671 | COIMBATORE | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | Monthly visit | Monthly visit done and carried out maintenance of the analyser | Monthly visit done and carried out maintenance of the analyser | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2502506806 | 2025-05-26 | 2025-05-30 | Cust 693 | COIMBATORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4 trans-A4TW033000(1) | Commissioning of pH, Chlorine and conductivity analysers at 7 locations | Analysers commissioned at 7 locations | Analysers commissioned | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2502533807 | 2025-05-26 | 2025-05-28 | Cust 665 | COIMBATORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | 12 V POWER SUPPLY CARD for Multiparameter(AWA-PS1) | Sample pump not working | Found Mother board and MUX card faulty, Suggested to send it to pune for necessary repairs | Found Mother board and MUX card faulty, Suggested to send it to pune for necessary repairs | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2503102808 | 2025-06-04 | 2025-06-07 | Cust 706 | COIMBATORE | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | commissioning | Commissioning of the MPA done and found analyser working ok | Commissioning of the MPA done and found analyser working ok | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCB2503458809 | 2025-06-13 | 2025-06-14 | Cust 671 | COIMBATORE | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | OIL ANALYSER | Flow cell(1) | monthly visit | Checked all the healthiness of analyser and found ok | Checked all the healthiness of analyser and found ok | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCD2303778810 | 2023-06-12 | 2023-06-29 | Cust 707 | CHANDIGARH | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Multiple failure observed while replacing the display unit and CUP, we observed that RS-485 cards and output cards are failed. | discussion with customer for new analyzer | discussion with customer for new analyzer | NO | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFCD2303994811 | 2023-06-14 | 2023-06-15 | Cust 708 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB connectivity needs to be done. | we took access of the Black Box and completed the CPCB connectivity and data is live on the server. | Data is live on the CPCB server hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFCD2305244812 | 2023-07-07 | 2023-07-07 | Cust 709 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software installation and CPCB + PPCB connectivity need to be done. | Software installation and CPCB + PPCB connectivity- has been completed and data is live. | Software installation and CPCB + PPCB connectivity- has been completed and data is live. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFCD2305362813 | 2023-07-10 | 2023-07-17 | Cust 710 | CHANDIGARH | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | WE HAVE GIVES THE INSTALLATION GUIDELINE FOR DO METER. | done | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFCD2305458814 | 2023-07-11 | 2023-07-12 | Cust 711 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Black Box configuration needs to be done. | The Black Box configuration has been completed and now ETP data is coming till the BB and internet are also pinging. | The Black Box configuration has been completed and now ETP data is coming till the BB and internet are also pinging. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFCD2306211815 | 2023-07-26 | 2023-07-29 | Cust 712 | CHANDIGARH | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | we have checked the DO METER and found it is not powering and the Sensor nut is also damaged. We have recommended to the customer, Please send the DO Meter to Pune for repair. | done | Closed | NO | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFCD2306411816 | 2023-07-29 | 2023-07-31 | Cust 713 | CHANDIGARH | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | MULTIPARAMETER ANALYSER | PH Input Card for Multiparameter(AWA-PH) | We checked the MPA and found the Ph value is showing -21323343893 Then We checked the PH card input -2 and found the PH input-2 is not catch the signal, Then we inserted the PH card in input-3 and reset the Multi parameters analyzer. Now MPA is working fine . | done | Closed | NO | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFCD2307147817 | 2023-08-10 | 2023-08-11 | Cust 713 | CHANDIGARH | DCS - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | MICROCON+ER SERIES | CONTROLLER MODULE (ER-201)(ER-201) | CPCB data upload issue. | We took access of the BB and made some changes so now the data is live on the server. | We took access of the BB and made some changes so now the data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFCD2307185818 | 2023-08-11 | 2023-08-11 | Cust 711 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB connectivity needs to be done. | today we took access of the Black Box and completed the CPCB connectivity and data is live on the server. | today we took access of the Black Box and completed the CPCB connectivity and data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFCD2307207819 | 2023-08-12 | 2023-08-17 | Cust 710 | CHANDIGARH | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | we have commissioned the DO Meter | now do meter is working fine | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFCD2307208820 | 2023-08-12 | 2023-08-17 | Cust 714 | CHANDIGARH | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | MULTIPARAMETER ANALYSER | Xenon Lamp Power supply for Multiparameter(AWA-XLAMPPS) | we have checked the Multi parameters analyzer and found the xenon power supply failed. Need to replace the xenon power supply. | we have submit the offer for same | Closed | NO | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFCD2307212821 | 2023-08-12 | 2023-08-17 | Cust 715 | CHANDIGARH | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | Ph Electrode | FM1104(1) | We have checked the transmitter and found that transmitter is working fine, But the PH electrode not working properly. Need to replacement the ph electrode. | we will replace the ph sensor on FOC basis. | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFCD2307278822 | 2023-08-14 | 2023-08-17 | Cust 711 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | PPCB connectivity needs to be done. | The PPCB connectivity has been completed and now the data is live on the server | The PPCB connectivity has been completed and now the data is live on the server | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFCD2307671823 | 2023-08-21 | 2023-08-21 | Cust 716 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation needs to be done. | We took the access to the system and successfully did the software reinstallation + CPCB + PPCB configuration, and now data is live on both the servers | We took the access to the system and successfully did the software reinstallation + CPCB + PPCB configuration, and now data is live on both the servers | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFCD2308467824 | 2023-09-05 | 2023-09-07 | Cust 717 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Black Box configuration needs to be done. | Black Box configuration is completed and data is live on both servers. | Black Box configuration is completed and data is live on both servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFCD2308508825 | 2023-09-06 | 2023-09-08 | Cust 717 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB and PPCB data uploaded issue, | Now the data is live on the CPCB server. please find the below screenshot for your reference. For PPCB, we are getting an Exception while uploading hence we will reopen this issue once the network issue has been resolved. | Now the data is live on the CPCB server. For PPCB, we are getting an Exception while uploading hence we will reopen this issue once the network issue has been resolved. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFCD2308637826 | 2023-09-11 | 2023-09-11 | Cust 717 | CHANDIGARH | Analytical - CHAKAN | CLOSED | COMMISSIONING | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Blackbox has been installed and its wiring has been completed as well. System was also taken online after black box configuration was completed. | Blackbox has been commissioned, MOM report attached | Complaint is closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFCD2308653827 | 2023-09-11 | 2023-10-31 | Cust 711 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data uploaded issue. | Our R&D team has made some changes in the Black Box and the data has been uploading continuously | Our R&amp;amp;D team has made some changes in the Black Box and the data has been uploading continuously | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFCD2308868828 | 2023-09-14 | 2023-09-15 | Cust 718 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation needs to be done. | We took the access to the system and successfully did the software reinstallation + local configuration, and now data is live on local server. | We took the access to the system and successfully did the software reinstallation + local configuration, and now data is live on local server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFCD2309014829 | 2023-09-18 | 2023-09-18 | Cust 717 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DIGITAL CONEECTIVITY | MULTIPARAMETER ANALYSER | LCD of Multiparameter Display(1) | THis analyser was showing incorrect PH, COD, BOD readings. We have recalibrated using online mode of the COD BOD however PH card is not working as per our observation. Same needs to be replaced. | Ph card needs to be replaced | Complaint is closed | NO | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFCD2309523830 | 2023-09-28 | 2023-09-28 | Cust 719 | CHANDIGARH | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | pH input card-FM of Multiparameter(1) | We were facing issue with PH part of the analyser, even after changing the PH card and its sensor the system was not sensing the PH correctly. IN my second visit i carried mV source with me and calibrated the card and then further calibrated the Ph sensor using buffer solution. NOw analyzer is ok | Calibration done, compliant closed. | Complaint is closed | YES | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFCD2309592831 | 2023-09-29 | 2023-09-30 | Cust 707 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CX-2000 analyser Commissioning for Nahar Spinning Mills | We have successfully configured the BB in site and now data is ping till the BB | We have successfully configured the BB in site and now data is ping till the BB | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFCD2309594832 | 2023-09-29 | 2023-09-30 | Cust 717 | CHANDIGARH | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | commissioning of CX2000 | commissioning done | commissioning done | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFCD2309665833 | 2023-09-30 | 2023-10-02 | Cust 709 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data mismatch issue with analyzer value. | We took the access to the system and checked the data on MODSCAN, and as per calculations the same data is reflecting onto the CPCB server. | We took the access to the system and checked the data on MODSCAN, and as per calculations the same data is reflecting onto the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFCD2309988834 | 2023-10-05 | 2023-10-06 | Cust 707 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | PPCB data uploaded issue. | Now the DATA is live on the PPCB server | Now the DATA is live on the PPCB server | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFCD2309989835 | 2023-10-05 | 2023-10-09 | Cust 713 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB AND PPCB DATA UPLOADED ISSUE | As of now, we are closing this complaint and will reopen this case again once we get remote access of the Black Box | As of now, we are closing this complaint and will reopen this case again once we get remote access of the Black Box | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFCD2310166836 | 2023-10-09 | 2023-11-06 | Cust 720 | CHANDIGARH | Swas - CHAKAN | CLOSED | PRODUCT FAILURE | SWAS & Analysers | Steam & Water Analysis System(SWAS & Analysers) | Aqua4Trans for Degas conductivity(SN-23A4T0461-W030100) transmitter display is flickering and keypad not working. | New transmitter dispatched | New transmitter distpacthed under goodwill. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFCD2310173837 | 2023-10-09 | 2024-06-12 | Cust 720 | CHANDIGARH | Swas - CHAKAN | CLOSED | APPLICATION ISSUE | SWAS & Analysers | Steam & Water Analysis System(SWAS & Analysers) | Degas conductivity FM8310 sensor flowthrough chamber supplied of PTFE instead of SS for 100 deg. above temprature of degas cond. sample. Kindly supplied SS flowthrough chamber for the same. | MARK EXPRESS PRIVATE LIMITED 210097268 / 27-APR-2024 D Gas heater already supplied. | MARK EXPRESS PRIVATE LIMITED 210097268 / 27-APR-2024 | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFCD2310635838 | 2023-10-18 | 2023-10-20 | Cust 709 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue. | Data is not communicating with the local system, so as of now we are closing this complaint. | Data is not communicating with the local system, so as of now we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFCD2310661839 | 2023-10-19 | 2023-10-20 | Cust 721 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB and PPCB data uploaded issue. | now the network issue has been resolved and data is live on the server. | now the network issue has been resolved and data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFCD2310696840 | 2023-10-19 | 2023-11-01 | Cust 717 | CHANDIGARH | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | pH input card-FM of Multiparameter(1) | The PH card of this multiparameter was found faulty, as we tried resetting the system as well along with online calibration and buffer solution calibration. This card was not responding to the PH of buffer solution. | RF prepared | RF prepared | YES | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFCD2311130841 | 2023-10-29 | 2023-10-31 | Cust 707 | CHANDIGARH | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | mA Output card for Multiparameter(AWA-4-20mA) | we have calibrated the multiparameter analyzer. | done | Closed | NO | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFCD2311332842 | 2023-10-31 | 2023-12-12 | Cust 711 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data uploaded issue | Data is live on the CPCB and For GPCB we are checking this issue iwth GPCB so as of now we are closing this complaint. | Data is live on the CPCB and For GPCB we are checking this issue iwth GPCB so as of now we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFCD2311827843 | 2023-11-09 | 2023-11-09 | Cust 709 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB and PPCB. | Now data is live on RSPCB server. | Now data is live on RSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFCD2311840844 | 2023-11-09 | 2023-11-10 | Cust 709 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB + PPCB data upload issue. | Now data is live on CPCB + PPCB server, hence now we are closing this complaint. | Now data is live on CPCB + PPCB server, hence now we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFCD2313003845 | 2023-12-01 | 2023-12-04 | Cust 713 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data uploaded issue. | The issue has been resolved and data is live on the server. | The issue has been resolved and data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFCD2313022846 | 2023-12-01 | 2023-12-04 | Cust 722 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB and RSPCB connectivity needs to be done. | We have taken up this issue with the AIPL Team and now the data is uploading continuously | We have taken up this issue with the AIPL Team and now the data is uploading continuously | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFCD2313530847 | 2023-12-13 | 2023-12-15 | Cust 723 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data uploded issue. | we have checked and For ETP data is not communicating with the Local PC and For Flow data is live on the CPCB sever. | we have checked and For ETP data is not communicating with the Local PC and For Flow data is live on the CPCB sever. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFCD2314398848 | 2023-12-28 | 2024-01-29 | Cust 717 | CHANDIGARH | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Need teltonica sin based router | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210078336 LR Dt - 05-JAN-2024 | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210078336 LR Dt - 05-JAN-2024 | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFCD2314583849 | 2023-12-31 | 2023-12-31 | Cust 717 | CHANDIGARH | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | MULTIPARAMETER ANALYSER | Falcon Scanner for Multiparameter(AWA-SCAN) | connection of PPCB | issue in customer internet | issue in customer internet | YES | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFCD2314829850 | 2024-01-04 | 2024-01-05 | Cust 717 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation and Local + CPCB + PPCB connectivity. | We took access to the system and re-installed the software and completed Local + CPCB + PPCB connectivity. Data is live on the same servers. | We took access to the system and re-installed the software and completed Local + CPCB + PPCB connectivity. Data is live on the same servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFCD2315397851 | 2024-01-15 | 2024-01-18 | Cust 715 | CHANDIGARH | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Electrode | Chemotrode ORP 120(HT238 740) | We have commissioned the pH electrode and calibrated it with a buffer solution. It is under observation for one week. | done | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFCD2316030852 | 2024-01-29 | 2024-01-31 | Cust 717 | CHANDIGARH | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | MULTIPARAMETER ANALYSER | FM make pH card(1) | card replacement and re calibration | replacement done and calibrated | replacement done and calibrated | NO | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFCD2316879853 | 2024-02-12 | 2024-02-12 | Cust 717 | CHANDIGARH | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | MULTIPARAMETER ANALYSER | FM make RS485 card(1) | commissioning of RS485 card along with connectivity of the system to PPCB and CPCB done | commissioning done | Closed | NO | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFCD2317964854 | 2024-02-29 | 2024-02-29 | Cust 717 | CHANDIGARH | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | MULTIPARAMETER ANALYSER | RS485 card-FM of Multiaparameter (1) | communication issue in analyzer | issue resolved and system is running smoothly | issue resolved and system is running smoothly | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFCD2318228855 | 2024-03-05 | 2024-03-08 | Cust 722 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue. | now the data is live on the server. | now the data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFCD2318229856 | 2024-03-05 | 2024-03-08 | Cust 716 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | PPCB data upload issue. | As of now we are closing this complaint and will reopen this case once we receive remote access of the system. | As of now we are closing this complaint and will reopen this case once we receive remote access of the system. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFCD2318526857 | 2024-03-13 | 2024-03-15 | Cust 717 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not updating to Local server. | We took access to your system and data is now live on the Local server. | We took access to your system and data is now live on the Local server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFCD2319538858 | 2024-03-29 | 2024-03-29 | Cust 712 | CHANDIGARH | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | MULTIPARAMETER ANALYSER | Flowcell disc of Multiparameter(1) | installation guidelines for the multiparameter analyzer and SPM analyzer | guideline are given | guideline are given | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFCD2400800859 | 2024-04-19 | 2024-04-19 | Cust 724 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | New device connectivity to CPCB | We have completed CPCB connectivity and data is live on CPCB Server. camera streaming is live on CPCB server | We have completed CPCB connectivity and data is live on CPCB Server. camera streaming is live on CPCB server | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFCD2401234860 | 2024-04-26 | 2024-04-28 | Cust 721 | CHANDIGARH | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | This visit was done to calibrate the system as per routine visit as AMC is signed with the customer. | complaint is closed. | Complaint is closed | YES | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFCD2401294861 | 2024-04-28 | 2024-04-28 | Cust 719 | CHANDIGARH | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | This visit was done to calibrate the system as per customer requirement and provide relay output for pump tripping on higher COD values. | complaint is closed | Complaint is closed | NO | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFCD2401591862 | 2024-05-01 | 2024-10-01 | Cust 725 | CHANDIGARH | Analytical - CHAKAN | CLOSED | ENGINEERING ISSUE | Conductivity Electrode | Forbes Marshall Cell Model 8312, K=1(FM -8312) | analyzer is having sensor of 0-7000 PPM capacity but actual TDS in plant is having 0-10PPM,so we have to replace the sensor having range of 0-70PPM. | There is no issue with Functionality of unit. Sales team to decide on Correct range of sensor. Hence closing the complaint. | There is no issue with Functionality of unit. Sales team to decide on Correct range of sensor. Hence closing the complaint. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFCD2401695863 | 2024-05-05 | 2024-05-09 | Cust 707 | CHANDIGARH | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | Sample Pump(1) | We have calibrated the Multi-Parameter analyzer and working fine | done | Closed | NO | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFCD2401768864 | 2024-05-06 | 2024-07-01 | Cust 726 | CHANDIGARH | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(Ph-Trans-Smartpro) | pH electrode not responding to any input buffer reading constant 6.4pH | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210102557 LR Dt - 10-JUN-2024 | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210102557 LR Dt - 10-JUN-2024 | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFCD2402318865 | 2024-05-20 | 2024-05-20 | Cust 727 | CHANDIGARH | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | As per my visit regarding the Auqamon2000 Transmitter not responding properly. We have checked the transmitter with the Simulator & looped 2wire and found the transmitter is responding and working fine. But the reading was constant so we checked the pH electrode with the buffer solution 4,7 & 10ppm and found the pH electrode ( Polilyte pro 120 ) shell was not working properly. CC Dear Amol sir, as discussed with you during the site visit, the pH electrode shell was not working, We need to replace the pH electrode within warranty. Transmitter Sr nO- 24IRPD001 | done | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFCD2402400866 | 2024-05-21 | 2024-05-23 | Cust 724 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB server data upload issue. | we have changed Slave ID and Baud Rate for dust parameter, and now data is live on CPCB server. | we have changed Slave ID and Baud Rate for dust parameter, and now data is live on CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFCD2402489867 | 2024-05-23 | 2024-05-24 | Cust 712 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DIGITAL CONEECTIVITY | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Digital connectivity needs to be done for CX2000 | Complted the cloud server connectivity and data is live on the server. | Complted the cloud server connectivity and data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFCD2402525868 | 2024-05-23 | 2024-05-24 | Cust 709 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB and PPCB data uploaded issue. | We took access of the system and it is observed that the USB serial com port got changed in the system and now we have made the necessary changes at the server end. Now the data is live on the CPCB + PPCB server. | Now the data is live on the CPCB + PPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFCD2402712869 | 2024-05-27 | 2024-05-27 | Cust 717 | CHANDIGARH | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | MULTIPARAMETER ANALYSER | CPU Display card for Multiparameter(AWA-CPU1) | system got reset, commissioning done | commissioning done | commissioning done | NO | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFCD2402713870 | 2024-05-27 | 2024-05-27 | Cust 712 | CHANDIGARH | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | Sample Pump(1) | commissioning of CX 2000, with cloud connectivity | commissioning done | commissioning done | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFCD2403616871 | 2024-06-12 | 2024-06-14 | Cust 724 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue and camera streaming was not live on the server. | Now data is live on CPCB server. Also camera streaming is live on the CPCB server. | Now data is live on CPCB server. Also camera streaming is live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFCD2404182872 | 2024-06-24 | 2024-06-26 | Cust 709 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on CPCB + PPCB server. | Now data is live on CPCB + PPCB server. | Now data is live on CPCB + PPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFCD2404390873 | 2024-06-28 | 2024-07-01 | Cust 724 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | PPCB connectivity | As of now we are closing this complaint. We will reopen this once the internet issue is resolved from customer end. | As of now we are closing this complaint. We will reopen this once the internet issue is resolved from customer end. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFCD2404559874 | 2024-06-30 | 2024-06-30 | Cust 728 | CHANDIGARH | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | pH Transmitter Aquamon | Aquamon 2000 plus(AQ2000 plus) | this visit was done to check PH sensor and its transmitter along with conductivity meters. | closed | Complaint is closed | NO | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFCD2404560875 | 2024-06-30 | 2024-06-30 | Cust 719 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | This visit was done to install and run FMBB for CX1000, system has been configured and is ready for data transmission | closed | Complaint is closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFCD2404567876 | 2024-06-30 | 2024-06-30 | Cust 719 | CHANDIGARH | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | Aqua4Trans | Communication card of Aqua4Trans(1) | Visit was done to commission the Aqua4trans TDS sensor installed along with cx1000 sensor. | closed | Complaint is closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFCD2405606877 | 2024-07-19 | 2024-07-19 | Cust 711 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data uploaded issue. | now the issue has been resolved and data is live on the CPCB server. | now the issue has been resolved and data is live on the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFCD2405715878 | 2024-07-22 | 2024-08-08 | Cust 707 | CHANDIGARH | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | MULTIPARAMETER ANALYSER | Xenon Lamp Power supply for Multiparameter(AWA-XLAMPPS) | Xenon Power Supply Of Newly Supplied CX2000 Analyser failed | material received to customer | material received to customer | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFCD2406003879 | 2024-07-26 | 2024-07-27 | Cust 724 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | PPCB configuration needs to be done. | We took access to the system ,and completed PPCB configuration and now the data is live on CPCB + PPCB servers, Also camera streaming is live on the CPCB server. | We took access to the system ,and completed PPCB configuration and now the data is live on CPCB + PPCB servers, Also camera streaming is live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFCD2406113880 | 2024-07-28 | 2024-07-28 | Cust 719 | CHANDIGARH | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | Aqua4Trans | Aqua4Trans(1) | This visit was done to install aqua4trans TDS based sensor along with the transmitter. | closed | Complaint is closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFCD2406340881 | 2024-07-30 | 2024-07-31 | Cust 719 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB configuration needs to be done. | We took access of the black box and completed CPCB connectivity, now data is live on CPCB server. | We took access of the black box and completed CPCB connectivity, now data is live on CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFCD2406405882 | 2024-07-31 | 2024-07-31 | Cust 712 | CHANDIGARH | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | MULTIPARAMETER ANALYSER | PH1 detector(1) | We have calibrated the Analyzer and reset the parameters. Now it is under observation for one week. | done | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFCD2406417883 | 2024-07-31 | 2024-07-31 | Cust 719 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | This visit was done to connect the site to CBCB server. Site is connected now | closed | Complaint is closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFCD2406430884 | 2024-07-31 | 2024-08-02 | Cust 724 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not updating to CPCB server. | Issue has been resolved and the data is now live on the CPCB and PPCB servers. | Issue has been resolved and the data is now live on the CPCB and PPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFCD2407541885 | 2024-08-23 | 2024-08-23 | Cust 719 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data upload issue. | we took access of the Black Box and observed that the internet is not coming till the BB, now the same issue is resolved from your side and now data is live on the CPCB server. | we took access of the Black Box and observed that the internet is not coming till the BB, now the same issue is resolved from your side and now data is live on the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFCD2409333886 | 2024-09-25 | 2024-09-25 | Cust 717 | CHANDIGARH | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | MULTIPARAMETER ANALYSER | Flowcell disc of Multiparameter(1) | This was a proactive visit to suggest customer that they should upgrade the existing CX1000 to 2000 model as installing spares in the existing 1000 model was not viable. | closed | Complaint is closed | NO | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFCD2409336887 | 2024-09-25 | 2024-09-25 | Cust 729 | CHANDIGARH | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | pH Transmitter Aquamon | Ph Transmitter Aquamon 2000(PH-Trans Aqu-2000) | This visit was done to convince customer and explain that Aquacon 2000 model does not have relay output, therefore either buy a DPI switch or upgrade to Aqua4trans. | closed | Complaint is closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFCD2409387888 | 2024-09-26 | 2024-09-28 | Cust 722 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on CPCB Server. | We took access to the system and now the data is live on CPCB + PPCB servers. | We took access to the system and now the data is live on CPCB + PPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFCD2410395889 | 2024-10-16 | 2024-11-28 | Cust 711 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB AND PPCB DATA UPLOADED ISSUE. | Now the network issue has been resolved from the customer end and the data is live on the CPCb and PPCb srever. | Now the network issue has been resolved from the customer end and the data is live on the CPCb and PPCb srever. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFCD2410697890 | 2024-10-22 | 2024-10-25 | Cust 707 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | PPCB data uploaded issue. | The customer will arrange access on Monday so as of now we are closing this complaint. | The customer will arrange access on Monday so as of now we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFCD2411042891 | 2024-10-29 | 2024-10-30 | Cust 730 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Re-installation of EnviroConnect software. | We have accessed the system and taken all the necessary backups. As we discussed during call, with customer they are going to provide new system access on Monday for the re-installation of the software.. so as of now we are closing this complaint and we will reopen it once we get system access. | As we discussed during call, with customer they are going to provide new system access on Monday for the re-installation of the software.. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFCD2411150892 | 2024-10-31 | 2024-10-31 | Cust 707 | CHANDIGARH | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | Mother board-PCX2 of Multiparameter Analyzer(1) | Motherboard replacement visit as new motherboard was supplied at site, PH sensor calibrated using simulator | closed | Complaint is closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFCD2411227893 | 2024-11-04 | 2024-11-08 | Cust 730 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Re-installation of EnviroConnect software. | We took access to the system and successfully completed the software re-installation and configuration. And now the data is live on CPCB + Local + PPCB servers. | We took access to the system and successfully completed the software re-installation and configuration. And now the data is live on CPCB + Local + PPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFCD2411304894 | 2024-11-05 | 2024-11-10 | Cust 707 | CHANDIGARH | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | we have calibrated the multi parameters analyzer. | resolved | Closed | YES | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFCD2411402895 | 2024-11-07 | 2024-11-07 | Cust 731 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Local host server is not accessible. | We accessed the system and found that the Apache Tomcat service was stopped on customer's local PC. After restarting it, the local server is now accessible. | We accessed the system and found that the Apache Tomcat service was stopped on customer's local PC. After restarting it, the local server is now accessible. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFCD2411793896 | 2024-11-14 | 2024-11-15 | Cust 717 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB + PPCB data upload issue. | we took access of the black box and now the issue has been resolved and data is live on CPCB + PPCB server. | we took access of the black box and now the issue has been resolved and data is live on CPCB + PPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFCD2412974897 | 2024-12-05 | 2024-12-05 | Cust 717 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Re-installation of EnviroConnect software in windows 10 | As discussed with the customer, the current PC is running Windows 7, and the data is live on the CPCB, Local, and SPCB servers. The customer plans to arrange a Windows 10 PC in the next 2 to 3 months, or when the data is no longer live. For now, they do not wish to change the PC since the data is still live. | The customer plans to arrange a Windows 10 PC in the next 2 to 3 months, or when the data is no longer live. For now, they do not wish to change the PC since the data is still live. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFCD2413046898 | 2024-12-06 | 2024-12-09 | Cust 707 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | PPCB data uploaded issue. | As of now we are closing this complaint and will reopen it once the internet side issue is resolved from your side. | As of now we are closing this complaint and will reopen it once the internet side issue is resolved from your side. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFCD2413190899 | 2024-12-09 | 2024-12-10 | Cust 732 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Re- installation of EnviroConnect software. | As per discussion with the customer during the call, we accessed the system and found that some files related to the EnviroConnect software were deleted. We have successfully re-installed the software, completed the configuration, and the data is now live on the CPCB server. | We have successfully re-installed the software, completed the configuration, and the data is now live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFCD2413306900 | 2024-12-12 | 2024-12-15 | Cust 733 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | EnviroConnect software installation in new PC | As of now, we are closing this complaint. We will reopen it once we receive the AMC or PO details to proceed further | As of now, we are closing this complaint. We will reopen it once we receive the AMC or PO details to proceed further | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFCD2413377901 | 2024-12-13 | 2024-12-16 | Cust 707 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | PPCB data uploaded issue. | We changed the new internet IP address as per the details provided by the customer and after that, the data started updating automatically on the PPCB server without any issues. | We changed the new internet IP address as per the details provided by the customer and after that, the data started updating automatically on the PPCB server without any issues. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFCD2413664902 | 2024-12-20 | 2024-12-21 | Cust 733 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | RTDAS + EnvrioConnect Software Re-installation | We have accessed the system and re-installed the EnviroConnect software, and updated to the latest version of RTDAS. And now the data is live on both CPCB , local and PPCB servers. | We have accessed the system and re-installed the EnviroConnect software, and updated to the latest version of RTDAS. And now the data is live on both CPCB , local and PPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFCD2413890903 | 2024-12-24 | 2024-12-26 | Cust 730 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on CPCB + Local + PPCB servers. | We took access to the system and now the data is live on CPCB + Local + PPCB servers. | We took access to the system and now the data is live on CPCB + Local + PPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFCD2413985904 | 2024-12-26 | 2024-12-28 | Cust 723 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to CPCB server. | Issue has been resolved and data is now live on the PPCB server. | Issue has been resolved and data is now live on the PPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFCD2414866905 | 2025-01-09 | 2025-01-10 | Cust 722 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on PPCB. | We have accessed the system, and resolved the issue . Now the data is live on the PPCB server for the Stack 1_ HRP and Ball Mill. | We have accessed the system, and resolved the issue . Now the data is live on the PPCB server for the Stack 1_ HRP and Ball Mill. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFCD2414927906 | 2025-01-10 | 2025-01-13 | Cust 711 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data update issue. | commercially not clear. | commercially not clear. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFCD2415387907 | 2025-01-21 | 2025-01-24 | Cust 730 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue . | We will reopen this compliant once the hardware issue will be resolved. | We will reopen this compliant once the hardware issue will be resolved. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFCD2415435908 | 2025-01-22 | 2025-01-24 | Cust 711 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB + PPCB data upload issue. | we took access of the black box and now the issue has been resolved and data is live on the Central CPCB server. Also we have updated the new patch in FMBB. | we took access of the black box and now the issue has been resolved and data is live on the Central CPCB server. Also we have updated the new patch in FMBB. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFCD2415880909 | 2025-01-28 | 2025-03-06 | Cust 707 | CHANDIGARH | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | FM blackbox installaed on CX2000 is not working. | whole system will be sent to chakan for repair replacement | Complaint is closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFCD2416133910 | 2025-01-31 | 2025-01-31 | Cust 730 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on CPCB + Local + PPCB servers. | We have accessed your system, and the issue has now been resolved. The data is live on the CPCB, Local, and PPCB servers. | We have accessed your system, and the issue has now been resolved. The data is live on the CPCB, Local, and PPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFCD2416409911 | 2025-02-05 | 2025-05-13 | Cust 707 | CHANDIGARH | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | This complaint pertains to a non working FMBB as we are not getting ping on both of the ports of this blackbox, digital team had taken access of it from customer's laptop and from my laptop as well but ping was not available, please supply a new BB at this site. | UNit received in Depot Reference COmplaint NO CCS25D0235 | UNit received in Depot Reference COmplaint NO CCS25D0235 | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFCD2417070912 | 2025-02-21 | 2025-02-24 | Cust 711 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data uploaded issue. | As of now, we are closing this complaint. We will reopen it once we receive remote access of the system from your end | As of now, we are closing this complaint. We will reopen it once we receive remote access of the system from your end | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFCD2417101913 | 2025-02-23 | 2025-03-12 | Cust 734 | CHANDIGARH | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | AQUADO Analyzer | Visiwater DO sensor(1) | We have recently checked the DO Meter sensor and found that it is not functioning as expected. Upon further inspection, we also tested the mA and confirmed that the sensor is not working properly. | done | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFCD2417583914 | 2025-02-28 | 2025-03-03 | Cust 735 | CHANDIGARH | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | We have commissioned the DO Meter. | done | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFCD2417876915 | 2025-03-07 | 2025-03-10 | Cust 711 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB and PPCB data uploded issue | As of now we are closing this complaint and will reopen it again once we get remote access of the system from the coustomer end. | As of now we are closing this complaint and will reopen it again once we get remote access of the system from the coustomer end. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFCD2417954916 | 2025-03-10 | 2025-03-11 | Cust 719 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data upload issue. | We took access of the black box and now the time sync issue has been resolved and data is live on the CPCB server. | We took access of the black box and now the time sync issue has been resolved and data is live on the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFCD2418207917 | 2025-03-14 | 2025-03-14 | Cust 723 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB and PPCB data uploded issue. | Our site engineer visited your site today and resolved the hardware problem. The data is now live on both systems | Our site engineer visited your site today and resolved the hardware problem. The data is now live on both systems | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFCD2500678918 | 2025-04-18 | 2025-04-21 | Cust 736 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB and PPCB data uploded issue. | As of now, we are closing this complaint and will reopen this case once we get remote access of the system from coustomer end. | As of now, we are closing this complaint and will reopen this case once we get remote access of the system from coustomer end. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFCD2500681919 | 2025-04-18 | 2025-05-13 | Cust 737 | CHANDIGARH | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Aqua4Trans | Display card of Aqua4Trans(1) | This complaint pertains to a display card of aqua4trans transmitter which is showing white display, we have tried adjusting contrast but issue still persist. Equipment is under warranty so please provide new display for this unit | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210140928 LR Dt - 07-MAY-2025 | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210140928 LR Dt - 07-MAY-2025 | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFCD2500921920 | 2025-04-24 | 2025-05-13 | Cust 712 | CHANDIGARH | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | Xenon Lamp Power supply for Multiparameter(AWA-XLAMPPS) | As per today's visit, the CX2000 Multi Parameter analyzer is not working properly. I checked the analyzer and found the Xenon Lamp not working. Then we checked the Power supply of the Xenon lamp and found the Xenon lamp card not working. We checked the DC voltage and found zero. | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210140927 LR Dt - 07-MAY-2025 | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210140927 LR Dt - 07-MAY-2025 | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFCD2500953921 | 2025-04-25 | 2025-04-25 | Cust 707 | CHANDIGARH | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | MULTIPARAMETER ANALYSER | mA Output card for Multiparameter(AWA-4-20mA) | Calibration of Multiparameter Calibration | Calibration of Multiparameter Calibration done | Calibration done | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFCD2502658922 | 2025-05-28 | 2025-05-30 | Cust 721 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB + PPCB data upload issue. | we took access of the black box and now the time sync issue has been resolved and data is live on CPCB + PPCB server. | we took access of the black box and now the time sync issue has been resolved and data is live on CPCB + PPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFCD2503378923 | 2025-06-11 | 2025-06-11 | Cust 707 | CHANDIGARH | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB + PPCB + Cloud connectivity needs to be done. | we took access of the black box and completed CPCB + PPCB + Cloud connectivity, now data is live on all the servers. | we took access of the black box and completed CPCB + PPCB + Cloud connectivity, now data is live on all the servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFCM2406978924 | 2024-08-12 | 2024-08-27 | Cust 738 | COLOMBIA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Electrode | Ph Electrode(FM1104(4C)) | After installed the ph electrode and try to calibrate, the ph electrode does not respond to different ph Buffers (7 and 10), also I have try to reactivated the electrode putting in solution with 4 pH during some time according to manual recommendation but the answer was same, the electrode does not response. | sensor shelf life is over, technical inputs provided to hernando | sensor shelf life is over, technical inputs provided to hernando | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFCN2303334925 | 2023-06-01 | 2023-06-02 | Cust 739 | CHENNAI | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | Aqua4Trans | Aqua4 trans-A4TW033000(1) | DO Analayser installation/erection | visited the site and all necessary guidelines given to technical team | visited the site on 1-6-23 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN2303399926 | 2023-06-02 | 2023-06-08 | Cust 665 | CHENNAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | THERMO SERIES | TSS SENSOR-RT2143(1) | commissioning of BOD /COD (CX1000) analyzer | Commissioning of Multiparameter Analyser Done. | Commissioning done | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFCN2303594927 | 2023-06-06 | 2023-06-10 | Cust 600 | CHENNAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | PH Sensor for Multiparameter(AWA-PH-SENS) | pH sensor commissioning | Visited the site and commissioning done | visited the site on 7-6-23 | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN2303650928 | 2023-06-07 | 2023-06-12 | Cust 739 | CHENNAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4Trans(1) | DO analyser commissioning | visited the site and 1no of DO analyser commissioning done | visited the site on 8-6-23 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN2304705929 | 2023-06-27 | 2023-07-02 | Cust 665 | CHENNAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | COD BOD TSS analyser commissioning | visited the site and commissioning completed | visited the site on 28-6-23 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN2304985930 | 2023-07-02 | 2023-07-05 | Cust 602 | CHENNAI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | RS485 CARD for Multiparameter(AWA-RS485) | RS485 reading issue | visited the site-found RS485 card burnt in condition Asked customer to procure new card for normal functioning and offer sent to customer | visited the site on 3-7-23 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN2305066931 | 2023-07-03 | 2023-07-08 | Cust 740 | CHENNAI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | TSS erratic reading | After cleaning flowcell meter working good | Troubleshooting completed meter taken on line | NO | Paid | Core | N |  | Analytical - CHAKAN
AFCN2305193932 | 2023-07-06 | 2023-07-09 | Cust 741 | CHENNAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4Trans(1) | ORP TDS meter commissioning | Visited the site and commissioning completed | visited the site on 7-7-23 | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN2305552933 | 2023-07-12 | 2023-08-04 | Cust 742 | CHENNAI | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | Aqua4Trans | Display card of Aqua4Trans(1) | Display blur issue in Aqua4trans | Replacement dispatched to client | Replacement dispatched to client | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFCN2305884934 | 2023-07-19 | 2023-07-24 | Cust 743 | CHENNAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4 trans-A4TW033000(1) | Commissioning work | Commissioning completed kept under observation | Commissioning completed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN2305907935 | 2023-07-20 | 2023-07-22 | Cust 739 | CHENNAI | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | AQUADO Analyzer | AQUAMAX DOT(1) | DO Analyzer Installation /erection | Visited the site and all necessary erection guidelines given to technical team | visited the site on 20-7-23 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN2305984936 | 2023-07-22 | 2023-07-28 | Cust 744 | CHENNAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | TOC ANALYSER | TOC 1950 Single StreamANALYSER ANATEL(1950 TOC) | TOC analyser commissioning | visited the site and analyzer having issues and spares need to buy from the client end | trouble shooting analyzer completed | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN2306147937 | 2023-07-25 | 2023-07-29 | Cust 744 | CHENNAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | Oil in water analyser commissioning | visited the site The analyzer having issues and spares need to buy from the client end | visited the site on 26-7-23 to 28-7-23 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN2306203938 | 2023-07-26 | 2023-07-28 | Cust 745 | CHENNAI | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Amc visit for the month of July 2023 | AMC visit completed and founds the MPA working fine | Closed | NO | Periodic Maintenance (AMC) | Process | N |  | Analytical - CHAKAN
AFCN2307059939 | 2023-08-08 | 2023-08-11 | Cust 746 | CHENNAI | Analytical - CHAKAN | CLOSED | AMC | Aqua4Trans | Aqua4 trans-A4TW033000(1) | pH TSS meter AMC | AMC visit done | visited the site on 9-8-23 | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFCN2307177940 | 2023-08-11 | 2023-08-15 | Cust 747 | CHENNAI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Multiparameter analyser AMC | AMC visit successfully completed | visited the site on 10-8-23 | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFCN2307277941 | 2023-08-14 | 2023-08-15 | Cust 634 | CHENNAI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | pH meter calibration | Branch level calibration done and dispatched the material to client | Branch level calibration done and dispatched the material to client | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFCN2307330942 | 2023-08-15 | 2023-08-21 | Cust 602 | CHENNAI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | CPU Display card for Multiparameter(AWA-CPU1) | Multiparameter analyser display blank | Visited the site and found RS 485 card in burnt requested client to procure new card | visited the site on 17-8-23 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN2307447943 | 2023-08-17 | 2023-08-21 | Cust 747 | CHENNAI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Multiparameter analyser AMC | AMC visit successfully completed | visited the site on 18-8-23 | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFCN2307501944 | 2023-08-18 | 2023-08-21 | Cust 739 | CHENNAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | VISI pro DO(VISI pro DO) | DO analyser commissioning | Visited the site and commissioning done | visited the site on 18-8-23 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN2307530945 | 2023-08-19 | 2023-09-11 | Cust 748 | CHENNAI | Analytical - CHAKAN | CLOSED | APPLICATION ISSUE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | pH sensor not responding | Replacement dispatched to client | Replacement dispatched to client | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFCN2307596946 | 2023-08-21 | 2023-08-24 | Cust 636 | CHENNAI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Display blank | Visited the site and Display brightness was adjusted Now analyser working satisfactory | visited the site on 17-8-23 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN2308182947 | 2023-08-30 | 2023-09-06 | Cust 749 | CHENNAI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC Multiparameter | Visited the site and AMC visit successfully done | visited the site on 6-9-23 | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFCN2308461948 | 2023-09-05 | 2023-09-08 | Cust 750 | CHENNAI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | pH TDS meter checking | Visited the site and calibration done | visited the site on 5-9-23 | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN2308465949 | 2023-09-05 | 2023-09-08 | Cust 742 | CHENNAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4Trans(1) | TDS and turbidity meter commissioning | Visited the site and commissioning done | visited the site on 7-9-23 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN2308650950 | 2023-09-11 | 2023-09-15 | Cust 751 | CHENNAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | COD BOD TSS analyser commissioning | Visited the site and commissioning done | visited the site on 12-9-23 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN2309051951 | 2023-09-18 | 2023-09-23 | Cust 740 | CHENNAI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | COD BOD analyser calibration | Visited the site and calibration done | visited the site on 19-9-23 | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN2309180952 | 2023-09-21 | 2023-09-23 | Cust 600 | CHENNAI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | COD TSS erratic reading | Visited the site, cleaned the flow cell and taken the analyser online- now readings are ok | visited the site on 22-9-23 | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN2309902953 | 2023-10-03 | 2023-10-05 | Cust 752 | CHENNAI | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | Aqua4Trans | Aqua4Trans(1) | PRO-ACTIVE VISIT | checked aqua4trans found ok | PRO-ACTIVE VISIT COMPLETED | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN2309917954 | 2023-10-03 | 2023-10-07 | Cust 753 | CHENNAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | CPU Display card for Multiparameter(AWA-CPU1) | COD BOD analyser commissioning | Visited the site - replaced the PCX1 card and done all necessary programs and taken online. analyser kept under observation for a week | visited the site on 5-10-23 | NO | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN2310105955 | 2023-10-08 | 2023-10-12 | Cust 751 | CHENNAI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | PH Sensor for Multiparameter(AWA-PH-SENS) | pH erratic reading | visited the site and pH calibration done and analyser kept under observation | visited the site on 9th and 12th sep 23 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN2310325956 | 2023-10-12 | 2023-11-18 | Cust 753 | CHENNAI | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | PCX1 programming issue after power recycle | Replacement sent to customer | Replacement sent to customer | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFCN2310327957 | 2023-10-12 | 2023-11-08 | Cust 751 | CHENNAI | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | pH reading unstable | Visited and provided additional earth and working fine so far if any issue client will revert us | Visited and provided additional earth and working fine so far if any issue client will revert us | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFCN2310530958 | 2023-10-16 | 2023-12-20 | Cust 744 | CHENNAI | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | Oil in water analyser issue | replacement sent | replacement sent | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFCN2310717959 | 2023-10-20 | 2023-10-25 | Cust 740 | CHENNAI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | COD BOD analyser erratic reading | The analyzer found to be ok and taken online | Analyzer was taken online | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN2310727960 | 2023-10-20 | 2023-10-20 | Cust 747 | CHENNAI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on Local + CPCB + APPCB | We took access to the system and data is now live on Local + CPCB + APPCB servers. | We took access to the system and data is now live on Local + CPCB + APPCB servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFCN2311034961 | 2023-10-27 | 2023-10-31 | Cust 754 | CHENNAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Electrode | Chemotrode 120(HT238 760) | ORP commissioning | Visited the site and commissioning done | visted the site on 26-10-23 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN2311223962 | 2023-10-30 | 2023-11-04 | Cust 751 | CHENNAI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Multiparameter analyser pH reading unstable | Visited the site ,provided additional grounding to panel now pH readings are ok | visited the site on 30-10-23 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN2311225963 | 2023-10-30 | 2023-11-04 | Cust 636 | CHENNAI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | TSS erratic reading | visited the site, cleaned flow cell and done zero calibration and taken online- now readings are OK | visited the site 31-10-23 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN2311226964 | 2023-10-30 | 2023-11-04 | Cust 636 | CHENNAI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | TSS erratic reading | visited the site, cleaned flow cell and done zero calibration and taken online- now readings are OK | visited the site on 31-10-23 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN2311288965 | 2023-10-31 | 2023-10-31 | Cust 755 | CHENNAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Smartpro-con) | Commissioning of Aquacon SMARTpro 8967 p | Commissioning done | Commissioning done | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFCN2311784966 | 2023-11-08 | 2023-11-11 | Cust 756 | CHENNAI | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | DO Analyser erratic reading | Visited the site - found DO sensor faulty, requests customer to procure new sensor for normal functioning | visited the site on 8-11-23 | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN2312170967 | 2023-11-19 | 2023-11-24 | Cust 757 | CHENNAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4Trans(1) | Aqua4Trans Commissioning | Commissioning completed | Commissioning completed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN2312289968 | 2023-11-22 | 2023-11-25 | Cust 753 | CHENNAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | COD BOD analyser commissioning | Visited the site and replaced PCX1 card, configuration and calibration done Taken the analyser online and found readings are OK | visited the site on 23-11-23 | NO | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN2312328969 | 2023-11-22 | 2023-11-25 | Cust 751 | CHENNAI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | COD BOD TSS erratic reading | visited the site, cleaned flow path, sampling tank and flow cell zero calibration done and taken online now readings are ok | visited the site on 25-11-23 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN2312388970 | 2023-11-23 | 2023-11-29 | Cust 755 | CHENNAI | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | Conductivity Electrode | Forbes Marshall Cond. Cell Model 8310, K=0.1(FM 8311-1.5) | Calibration | Checked at Pune, Found Ok, sending back to customer | Checked at Pune, Found Ok, sending back to customer | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFCN2312461971 | 2023-11-25 | 2023-11-29 | Cust 751 | CHENNAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Electrode | FM1106(1) | pH sensor commissioning | visited the site, connected new pH sensor and verified the readings with buffer solution - found OK | visited the site on 25-11-23 | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN2312462972 | 2023-11-25 | 2023-11-29 | Cust 681 | CHENNAI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | COD BOD TSS analyser erratic reading | Visited the site-found PCX1 touch was not working properly Asked client to send the same to pune for analysis | visited the site on 28-11-23 | NO | Paid | Core | N |  | Analytical - CHAKAN
AFCN2313328973 | 2023-12-09 | 2023-12-14 | Cust 747 | CHENNAI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC COD BOD TSS Analyser | Visited the site and AMC visit done | visited the site on 9-12-23 | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFCN2313415974 | 2023-12-11 | 2023-12-16 | Cust 758 | CHENNAI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Silica Analyser | Silica 9210(9210) | Checking silica and sodium analyser | Visited the site Silica analyzer calibration done and working satisfactory | visited the site on 12th and 13th Dec 2023 | NO | Paid | Core | N |  | Analytical - CHAKAN
AFCN2313834975 | 2023-12-19 | 2023-12-25 | Cust 747 | CHENNAI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | COD BOD TSS analyser AMC | AMC visit successfully completed | visited the site on 19-12-23 | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFCN2313878976 | 2023-12-19 | 2023-12-22 | Cust 747 | CHENNAI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB + APPCB data upload issue. | We took access to your system and now the hardware side issue is resolved and data is live on CPCB + APPCB server. | We took access to your system and now the hardware side issue is resolved and data is live on CPCB + APPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFCN2313894977 | 2023-12-19 | 2023-12-22 | Cust 644 | CHENNAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Smartpro-con) | pH TDS meter commissioning | Visited the site and pH TDS meter commissioning completed | visited the site on 19-12-23 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN2314169978 | 2023-12-25 | 2024-01-01 | Cust 759 | CHENNAI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | COD BOD TSS analyser calibration | Visited the site and calibration works has completed | visited the site on 30-12-23 | NO | Paid | Core | N |  | Analytical - CHAKAN
AFCN2314720979 | 2024-01-02 | 2024-01-06 | Cust 751 | CHENNAI | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | cod bod analyser proactive visit | Visited the site and found analyser working satisfactory | visited the site on 2-1-24 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN2314807980 | 2024-01-03 | 2024-01-06 | Cust 760 | CHENNAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4Trans(1) | Conductivity meter commissioning | Visited the site and conductivity meter commissioning done | visited the site on 4-1-24 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN2315105981 | 2024-01-09 | 2024-01-15 | Cust 761 | CHENNAI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Silica Analyser | Silica 8891(8891) | Silica analyser erratic reading | Visited the site and calibration done | visited the site on 10-1-24 | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN2315424982 | 2024-01-15 | 2024-01-20 | Cust 665 | CHENNAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | COD BOD analyser commissioning | visited the site and commissioning completed | visited the site on 19-1-24 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN2315425983 | 2024-01-15 | 2024-01-20 | Cust 712 | CHENNAI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | pH TDS meter calibration | Visited the site and calibration done | visited the site on 18-1-24 | NO | Paid | Core | N |  | Analytical - CHAKAN
AFCN2315770984 | 2024-01-22 | 2024-01-27 | Cust 564 | CHENNAI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(Ph-Trans-Smartpro) | pH TDS and TSS meter erratic reading | visited the site and found 3nos TDS sensor has faulty and 1no of pH meter faulty Balance 7pH, 2ORP and 2Turbidity meters has been calibrated with buffer and observed meters are working OK | visited the site on 24-1-24 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN2315772985 | 2024-01-22 | 2024-01-29 | Cust 665 | CHENNAI | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | MULTIPARAMETER ANALYSER | PH Sensor for Multiparameter(AWA-PH-SENS) | pH reading fluctuation | visited the site and commissioning completed | visited the site on 23-1-24 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN2315796986 | 2024-01-23 | 2024-01-25 | Cust 613 | CHENNAI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not updating on APPCB server | Data is now updating on APPCB server | Data is now updating on APPCB server | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFCN2316513987 | 2024-02-03 | 2024-02-08 | Cust 762 | CHENNAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | COD BOD TSS analyser commissioning | visited the site and commissioning completed | visited the site on 3-2-24 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN2316542988 | 2024-02-04 | 2024-02-08 | Cust 686 | CHENNAI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | DO Analyzer erratic reading | Visited the site and found DO sensor and 1no of 200mm mag meter has faulty Requested customer to procure new one | visited the site | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN2316589989 | 2024-02-05 | 2024-02-08 | Cust 613 | CHENNAI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not updating to APPCB server | We have updated the output.jar file and currently data is available on the APPCB server. | We have updated the output.jar file and currently data is available on the APPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFCN2316597990 | 2024-02-05 | 2024-02-09 | Cust 749 | CHENNAI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC _ COD BOD TSS analyser | AMC visit done | visited the site on 8-2-24 | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFCN2316657991 | 2024-02-06 | 2024-02-09 | Cust 749 | CHENNAI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Data is not updating to CPCB server. | Data is now live on the CPCB server. | Data is now live on the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | DCS - CHAKAN
AFCN2316805992 | 2024-02-09 | 2024-02-09 | Cust 613 | CHENNAI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | APPCB data upload issue | We have updated the software patch and now the data is live on the APPCB server. | We have updated the software patch and now the data is live on the APPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFCN2317007993 | 2024-02-14 | 2024-02-18 | Cust 759 | CHENNAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | COD BOD TSS analyser Spare commissioning | Commissioning done | visited the site 14-2-24 | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN2317190994 | 2024-02-18 | 2024-02-24 | Cust 763 | CHENNAI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | COD BOD TSS Analyser calibration | visited the site and calibration completed | visited the site on 20-2-24 | NO | Paid | Core | N |  | Analytical - CHAKAN
AFCN2317333995 | 2024-02-20 | 2024-02-24 | Cust 600 | CHENNAI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Multiparameter analyser showing erratic reading | Visited the site and found that PH1 and PH2 detector cards are faulty due to water ingress, asked customer to buy new cards to make the readings available | visited the site 0n 22-2-24 | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN2318133996 | 2024-03-03 | 2024-03-08 | Cust 681 | CHENNAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | LCD of Multiparameter Display(1) | PCX1 card commissioning | visited the site and found UV lamp supply card not giving output and same has send to pine for analysis and rectification | visited the site on 4-3-24 | NO | Paid | Core | N |  | Analytical - CHAKAN
AFCN2318297997 | 2024-03-08 | 2024-03-12 | Cust 747 | CHENNAI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | COD BOD Analyser AMC | AMC visit done | visited the site on 5-3-24 | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFCN2318490998 | 2024-03-12 | 2024-03-17 | Cust 686 | CHENNAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | mPA commissioning | Visited the site and commissioning completed | visited the site on 13-3-24 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN2318506999 | 2024-03-13 | 2024-03-18 | Cust 764 | CHENNAI | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | TOC ANALYSER | TOC 1950 Single StreamANALYSER ANATEL(1950 TOC) | TOC analyser proactive visit | Visited the site and checked the analyser | visited the site on 12-3-24 | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN23185161000 | 2024-03-13 | 2024-03-17 | Cust 760 | CHENNAI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Aqua4Trans | Aqua4Trans(1) | Conductivity meter erratic reading | provided telephonic support and now conductivity meter readings are satisfactory | provided telephonic support and now conductivity meter readings are satisfactory | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFCN23188201001 | 2024-03-18 | 2024-03-21 | Cust 613 | CHENNAI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | APPCB data uploading issue. | Data is live on the APPCB server. | Data is live on the APPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFCN23192701002 | 2024-03-25 | 2024-03-30 | Cust 547 | CHENNAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | OIL ANALYSER | MUVX card(1) | Oil in water analyzer commissioning | Spare fixed and commissioning completed | visited the site on 26 to 28 mar 24 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN24000391003 | 2024-04-01 | 2024-04-05 | Cust 762 | CHENNAI | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | pH reading fluctuation | Visited the site and solved pH reading issue | visited the site on 4-4-24 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN24000881004 | 2024-04-02 | 2024-04-05 | Cust 644 | CHENNAI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | pH and TDS meter erratic reading | Visited the site and verified with buffer - found no issue in meter | visited the site on 3-4-24 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN24003521005 | 2024-04-08 | 2024-04-13 | Cust 600 | CHENNAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | CPU Display card for Multiparameter(AWA-CPU1) | PCX1 commissioning | visited the site and commissioning done and found that PH1 and PH2 detector values are low, asked customer to procure detector card with filter to make the reading available | visited the site on 11/4/24 | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN24004881006 | 2024-04-10 | 2024-04-16 | Cust 686 | CHENNAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | cx 2000 commissioning | visited the site and calibration done | visited the site on 11-4-24 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN24004891007 | 2024-04-10 | 2024-04-17 | Cust 740 | CHENNAI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | cx 1000 erratic reading | Visited the site and calibration done | visited the site | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN24008321008 | 2024-04-19 | 2024-04-25 | Cust 765 | CHENNAI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | Product Failure | Spoke with customer on WhatsApp Video call and DO meters were calibrated in air Taken the meters in online- found working satisfactory | Spoke with customer on WhatsApp Video call and DO meters were calibrated in air Taken the meters in online- found working satisfactory | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFCN24011251009 | 2024-04-24 | 2024-04-27 | Cust 547 | CHENNAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | Oil in water analyser commissioning | visited the site and commissioning completed | visited the site on 26-4-24 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN24013091010 | 2024-04-29 | 2024-05-02 | Cust 665 | CHENNAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | CX 2000 Commissioning | Visited the site and commissioning completed | visited the site 30/4/24 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN24018251011 | 2024-05-07 | 2024-05-10 | Cust 749 | CHENNAI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC MPA | AMC visit successfully completed | visited the site on 10-5-24 | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFCN24028531012 | 2024-05-28 | 2024-06-01 | Cust 766 | CHENNAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | CX 2000 commissioning | Visited the site and commissioning completed | visited the site on 28-5-24 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN24029041013 | 2024-05-29 | 2024-05-31 | Cust 767 | CHENNAI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | APPCB data uploading issue | Data is now live on the APPCB server. | Data is now live on the APPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFCN24032601014 | 2024-06-04 | 2024-06-08 | Cust 497 | CHENNAI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | COD BOD analyser calibration | Visited the site and calibration done | visited the site on 5-6-24 | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN24032611015 | 2024-06-04 | 2024-06-08 | Cust 753 | CHENNAI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | COD BOD analyser calibration | visited the site and calibration done | visited the site on 6-6-24 | NO | Paid | Core | N |  | Analytical - CHAKAN
AFCN24033861016 | 2024-06-07 | 2024-06-09 | Cust 636 | CHENNAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | pH TDS and ORP meter commissioning | Visited the site and commissioning completed | visited the site on 7-6-24 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN24034551017 | 2024-06-09 | 2024-06-23 | Cust 745 | CHENNAI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC servie -- Repalcement of pH card & sensor | PH card not able to fix, calibrated & hard reset done. Under observation. | Closed | NO | Periodic Maintenance (AMC) | Process | N |  | Analytical - CHAKAN
AFCN24035131018 | 2024-06-10 | 2024-06-17 | Cust 564 | CHENNAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | Aquamax system(1) | Need service person visit at HCCB site | Visited Site, Discussed With Division, Asking to send analyser to Factory to test and upgrade the analyser cards. | Analyser need to sent to pune factory for Testing. | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFCN24035461019 | 2024-06-10 | 2024-06-14 | Cust 600 | CHENNAI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | COD BOD TSS erratic reading | Visited the site- found PH detectors are too low and requested customer to procure the spares to make the reading available | visited the site on 12-6-24 | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN24035541020 | 2024-06-11 | 2024-06-14 | Cust 647 | CHENNAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | PH, TDS, ORP, Chlorine commissioning & calibration | Visited the site and calibration done | visited the site on 13-6-24 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN24037351021 | 2024-06-14 | 2024-06-17 | Cust 665 | CHENNAI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | COD BOD TSS erratic reading found | Found sample pump not working , asked the customer to send it for analysis & rectification | Sample pump having issue, main board sensor signal getting zero ,need to do further analysis in chennai office | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN24041171022 | 2024-06-23 | 2024-06-27 | Cust 750 | CHENNAI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | pH and TDS meter checking | 9 nos flow meters are checked and 7nos of pH TDS calibration done | visited the site on 24-6-24 | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN24045011023 | 2024-06-29 | 2024-07-04 | Cust 750 | CHENNAI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | pH TDS ERRATIC READING | Visited the site and found pH sensor has faulty requested client to procure new sensor with FTC | visited the site on 1-7-24 | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN24054771024 | 2024-07-16 | 2024-07-21 | Cust 768 | CHENNAI | Krohne - MIDC | CLOSED | PRODUCT FAILURE | Vortex Meter OPTISWIRL Converter | Converter of OPTISWIRL 4200 (KMT) Compact Housing(VFC 200CCON) | VFM erratic reading | visited the site and found vibration in the vortex flow meter line , asked client to arrest the vibration | visited the site on 16/7/24 | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN24058471025 | 2024-07-24 | 2024-07-26 | Cust 497 | CHENNAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | MPHDO card commissioning | Visited the site and pH card commissioned | visited the site on 25-7-24 | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN24058481026 | 2024-07-24 | 2024-08-07 | Cust 665 | CHENNAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Cod bod commissioning after service, Visited and asked the customer to send the suspected faulty items for further analysis/ rectification. | Visited and asked the customer to send the suspected faulty items for further analysis/ rectification | Visited the site on 2-8-2024 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN24063381027 | 2024-07-30 | 2024-08-02 | Cust 766 | CHENNAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Turbidity Meter | Optex turbidity sensor(1) | Turbidity and conductivity meter commissioning | Visited the site and commissioning done | visited the site on 30-7-24 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN24063961028 | 2024-07-31 | 2024-08-08 | Cust 766 | CHENNAI | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | Conductivity Tranmitter Smartpro | Display of Smartpro(1) | Display issue | Display card sent to factory for analysis | Display card sent to factory for analysis | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFCN24067541029 | 2024-08-06 | 2024-08-12 | Cust 476 | CHENNAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Chlorine Analyzer-MAC 100 | MAC 100(1) | Chlorine analyser commissioning | Commissioning done | commissioning done | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFCN24067601030 | 2024-08-06 | 2024-08-10 | Cust 740 | CHENNAI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | COD BOD TSS erratic reading | Visited the site, display issue solved and reconfiguration done Taken online and now readings are OK | visited the site on 7-8-24 | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN24068061031 | 2024-08-07 | 2024-08-10 | Cust 764 | CHENNAI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | Oil in water analyser calibration | visited the site and carried out the zero calibration, requested to observe in online sample and revert back | visited the site 0n 8-8-24 | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN24069051032 | 2024-08-10 | 2024-08-15 | Cust 769 | CHENNAI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Smartpro-con) | Conductivity erratic reading | Visited the site, verified all programs and found there is no issue in meter, requested to check with standard solution and revert | visited the site on 9-8-24 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN24072971033 | 2024-08-19 | 2024-08-22 | Cust 613 | CHENNAI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to APPCB server. | Issue has been resolved and the data is now live on the APPCB server. | Issue has been resolved and the data is now live on the APPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFCN24075201034 | 2024-08-23 | 2024-08-23 | Cust 613 | CHENNAI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not updating to APPCB server. | the data is live on the APPCB server. | the data is live on the APPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFCN24079291035 | 2024-08-29 | 2024-08-30 | Cust 749 | CHENNAI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | RS485 card-FM of Multiaparameter (1) | AMC | AMC is successfully done,also COD,BOD OCEMS ETP online values are fetched with APPCB verified and found Ok. | visited the site on 29-8-24 | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFCN24085951036 | 2024-09-09 | 2024-09-13 | Cust 770 | CHENNAI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | pH TDS ORP meter calibration | Visited the site and calibration done | visited the site on 10th and 11th Sep 24 | NO | Paid | Core | N |  | Analytical - CHAKAN
AFCN24086681037 | 2024-09-10 | 2024-09-14 | Cust 761 | CHENNAI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | SODIUM ANALYSER | SODIUM 9073 single channel ANLAYSER(9073 Sodium) | Sodium analyser calibration | visited the site and calibrated the analyzer | visited the site on 11-9-24 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN24090911038 | 2024-09-19 | 2024-09-23 | Cust 764 | CHENNAI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | OIL ANALYSER | PM Board(PM-BD) | ANALYSER TRIP | PM CARD CHANGED NOW THE ANALYSER WORKING FINE. | visited the site on 9-8-24,19-9-24 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN24093021039 | 2024-09-24 | 2024-09-26 | Cust 665 | CHENNAI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | SAMPLING PUMP NOT WORKING | AFTER SERVICE NOW WORKING FINE AND ZERO CALIBRATION DONE WITH DM WATER, CONNECTED THROUGH TNPCB | visited the site on 25-9-24 | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN24095661040 | 2024-09-29 | 2024-09-30 | Cust 771 | CHENNAI | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | Ph Electrode | Ph Electrode(FM1104(4C)) | Showing Incorrect Values | Checked Transmitter, found ok, checked electrode, found failed, asked to replace. | Guidelines given. | NO | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFCN24099371041 | 2024-10-07 | 2025-01-03 | Cust 564 | CHENNAI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Electrode | Chemotrode 120(HT238 760) | Sensor values not changing - showing 6.8 pH always | Replacement given to customer | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFCN24099371041 | 2024-10-07 | 2025-01-03 | Cust 564 | CHENNAI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Electrode | Chemotrode ORP 120(HT238 740) | Sensor values not changing - showing zero value | Replacement given to customer | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFCN24099371041 | 2024-10-07 | 2025-01-03 | Cust 564 | CHENNAI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Conductivity Electrode | Forbes Marshall Cell Model 8311, K=0.1(FM -8311) | Showing C sensor shot. | Replacement given to customer | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFCN24099581042 | 2024-10-07 | 2024-10-12 | Cust 768 | CHENNAI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Aqua4Trans | Aqua4Trans(1) | TDS checking | FMPL RECOMMENDED TO CHNAGE TO HMI AND TDS SENSOR | visited the site on 08-10-24 | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN24100471043 | 2024-10-08 | 2024-10-12 | Cust 751 | CHENNAI | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | CX2000 HEALTHINESS | CX 2000 WORKING FINE | visited the site on 08-10-24 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN24103101044 | 2024-10-14 | 2024-10-17 | Cust 772 | CHENNAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | pH TDS meter commissioning | PH TDS METER COMMISSIONING DONE | visited the site on 15-10-2024 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN24107821045 | 2024-10-23 | 2024-10-28 | Cust 773 | CHENNAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4Trans(1) | pH TDS ORP commissioning | VISITED THE SITE AND COMMISSIONED THE PH TDS ORP AND DO METER | visited the site on 23-10-24 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN24107851046 | 2024-10-23 | 2024-10-27 | Cust 681 | CHENNAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | Xenon Lamp Power supply for Multiparameter(AWA-XLAMPPS) | Xenon lamp power supply card commissioning | REPLACED DONE BUT XENON POWER SUPPLY CARD FAILURE NEED TO CHECK OUR FACTORY | visited the site on 23-10-24 | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN24109881047 | 2024-10-28 | 2024-10-31 | Cust 774 | CHENNAI | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | COD , BOD values showing zero all the time | Calibration done, working now | Working fine now | NO | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFCN24115451048 | 2024-11-11 | 2024-11-18 | Cust 712 | CHENNAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Communication card of Aqua4Trans(1) | Commissioning of TDS and pH sensor | Commissioning done. | Commissioning done | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFCN24116231049 | 2024-11-12 | 2024-11-15 | Cust 613 | CHENNAI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Local data upload issue. | data is live on the server. | Data is live on the server | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFCN24120311050 | 2024-11-19 | 2024-11-22 | Cust 764 | CHENNAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | Oil in water analyser Spare commissioning | Visited the site, spare card replaced and done zero and span calibration, online sample not available, requested client to check in online and revert back | visited the site on 22-11-24 | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN24123241051 | 2024-11-25 | 2024-11-28 | Cust 764 | CHENNAI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | TOC ANALYSER | TOC 1950 Single StreamANALYSER ANATEL(1950 TOC) | TOC analyser checking | visited the site and checked the analyser healthiness | visited the site on 22-11-24 | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN24133001052 | 2024-12-12 | 2024-12-16 | Cust 775 | CHENNAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | THERMO SERIES | TDS SENSOR(1) | TDS COMMISSIONING | Commissioning done | visited the site on 10-12-24 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN24133851053 | 2024-12-14 | 2024-12-21 | Cust 564 | CHENNAI | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | Aqua4Trans | PH input card(1) | Checks on pH transmitter | About 10 instruments was calibrated in the water treatment systems. | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFCN24140521054 | 2024-12-27 | 2024-12-31 | Cust 776 | CHENNAI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | ORP Meter display issue | Visited the site, noticed that power supply issue, separate 24v dc supply provided and powered up the meter, now working ok | visited the site on 27-12-24 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN24141151055 | 2024-12-28 | 2024-12-31 | Cust 760 | CHENNAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4Trans(1) | pH and ORP Meter commissioning | visited the site and commissioning done | visited the site on 30-12-24 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN24141171056 | 2024-12-28 | 2025-01-02 | Cust 777 | CHENNAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | pH TDS Turbidity analyser commissioning | Ph TDS Turbidity analyser commissioning dione | Visited the site on 30-12-2024&31-12-24 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN24143591057 | 2024-12-31 | 2025-01-11 | Cust 778 | CHENNAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Chlorine Analyzer-MAC 100 | MAC 100(1) | Engineer Required for Supervision of Erection & Commissioning of Residual Chlorine Analyser for Nagercoil Project. Exte | CHLORINE ANALYZER COMMISSIONED DONE NOW READINGS ARE SATISFACTORY | CHLORINE ANALYZER COMMISSIONED DONE | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFCN24160291058 | 2025-01-30 | 2025-01-31 | Cust 779 | CHENNAI | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | No display of COD and BOD readings, also not to transmit the data to TNPCB | Checked and given the repair estimate to customer | Closed | NO | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFCN24163041059 | 2025-02-03 | 2025-02-07 | Cust 749 | CHENNAI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | MPA AMC | AMC visit done | visited the site on 7-2-25 | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFCN24164431060 | 2025-02-07 | 2025-02-11 | Cust 760 | CHENNAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Relay card of Aqua4Trans(1) | Relay output commissioning | relay output configured and checked in office - found working OK | relay output configured and checked in office - found working OK | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFCN24165641061 | 2025-02-10 | 2025-02-13 | Cust 747 | CHENNAI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software re-installation | we re-installed the software with an updated version and completed Local and CPCB configuration. Data is now live on the Local, CPCB and APPCB servers. | we re-installed the software with an updated version and completed Local and CPCB configuration. Data is now live on the Local, CPCB and APPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFCN24165811062 | 2025-02-10 | 2025-02-13 | Cust 747 | CHENNAI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | MPA AMC | AMC Visit completed | visited the site on 10-2-25 | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFCN24166081063 | 2025-02-11 | 2025-03-08 | Cust 780 | CHENNAI | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | pH Transmitter Aquamon | Ph Transmitter Aquamon 2000(PH2000) | the PH Transmitter display is not showing. (Video Attached for your reference) you are requested to your depute Engineer to resolve the issue at the earliest. Kindly note the site address along with contact details as mentioned below: M/s Calcimag India Pvt Ltd Sipcot industrial Park phase -2 Plot No-A-15/4cheyyar Tiruvannamalai Chennai Tamilnadu-606601 Contact Person ? Mr.Madhav Musale MOB. No. - 7722060681 | replacement dispatched to the customer | replacement dispatched to the customer | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFCN24166191064 | 2025-02-11 | 2025-02-16 | Cust 781 | CHENNAI | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | Aqua4Trans | Aqua4Trans(1) | pH transmitter with Dtron Pre commissioning | Visited the site and procommissioning done | visited the site on 11-2-25 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN24167481065 | 2025-02-13 | 2025-02-18 | Cust 764 | CHENNAI | Analytical - CHAKAN | CLOSED | AMC | TOC ANALYSER | FM 101 TOC Analyser(FM 101 TOC) | TOC Analyzer AMC | AMC visit done | visited the site on 14th and 15th FEB 25 | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFCN24168041066 | 2025-02-16 | 2025-02-20 | Cust 764 | CHENNAI | Analytical - CHAKAN | CLOSED | AMC | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | oil in water analyser calibration | AMC visit done | visited the site on 14th and 15th Feb 25 | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFCN24169901067 | 2025-02-20 | 2025-02-22 | Cust 740 | CHENNAI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | COD BOD TSS calibration | visited the site and calibration done | visited the site on 20-2-25 | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN24171731068 | 2025-02-24 | 2025-02-28 | Cust 665 | CHENNAI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | MPA Calibration | Visited the site and found UV lamp power supply card has not powering up and same sent to factory for further analysis. | visited the site on 24-2-25 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN24172431069 | 2025-02-25 | 2025-02-28 | Cust 751 | CHENNAI | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | MPA General visit | Visited the site, checked all necessary programs, verified with DM water- found analyser working satisfactory. | visited the site on 25-2-25 | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN24178351070 | 2025-03-06 | 2025-03-08 | Cust 759 | CHENNAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | COD BOD TSS analyser spare commissioning | Visited the site, optic filters commissioned and analyser taken online | visited the site on 7-3-25 | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN24178951071 | 2025-03-08 | 2025-03-13 | Cust 782 | CHENNAI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Turbidity Meter | Optex turbidity sensor(1) | Turbidity meter calibration | Visited the site and calibration done | visited the site on 10-3-25 | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN24182181072 | 2025-03-14 | 2025-03-17 | Cust 781 | CHENNAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Silica Analyser | Silica 8891(8891) | Silica analyser commissioning | Visited the site and commissioning completed | visited the site on 14-3-25 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN24183631073 | 2025-03-18 | 2025-03-22 | Cust 766 | CHENNAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | Resistivity meter commissioning | Visited the site and commissioning completed | visited the site on 17-3-25 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN24183661074 | 2025-03-18 | 2025-03-22 | Cust 749 | CHENNAI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | COD BOD TSS Analyser AMC | AMC visit completed | visited the site on 20-3-25 | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFCN24184791075 | 2025-03-21 | 2025-03-25 | Cust 665 | CHENNAI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | COD BOD TSS analyser erratic reading | Visited the site, done zero calibration and observed the readings Analyzer working satisfactory. | visited the site on 21-3-25 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN24185211076 | 2025-03-22 | 2025-03-25 | Cust 763 | CHENNAI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | MPA Checking | Visited the site, found most of components are faulty and requested to procure the same to take the analyser online | visited the site on 24-3-25 | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN24185451077 | 2025-03-23 | 2025-03-25 | Cust 749 | CHENNAI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | MPA AMC | AMC visit done | visited the site on 24-3-25 | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFCN24187061078 | 2025-03-26 | 2025-03-29 | Cust 665 | CHENNAI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | Xenon Lamp Power supply for Multiparameter(AWA-XLAMPPS) | MPA spare fixing | Visited the site, replaced the spare and taken the analyser online. Now readings are satisfactory. | visited the site on 27-3-25 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN25000561079 | 2025-04-02 | 2025-04-06 | Cust 636 | CHENNAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | pH TDS ORP meter commissioning | Visited the site and commissioning completed | visited the site on 3-4-25 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN25000731080 | 2025-04-03 | 2025-04-06 | Cust 665 | CHENNAI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | TSS erratic reading | Visited and zero calibration done, verified with online sample reading are satisfactory. | visited the site on 3-4-25 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN25001601081 | 2025-04-05 | 2025-04-09 | Cust 665 | CHENNAI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | Xenon Lamp Power supply for Multiparameter(AWA-XLAMPPS) | TSS erratic reading | Visited the site and done zero calibration, now readings are OK | visited the site on 8-4-25 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN25002691082 | 2025-04-08 | 2025-04-11 | Cust 613 | CHENNAI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to APPCB server. | Currently data is live on the APPCB server. We have checked on the modscan and observed that out of 275 poll requests, only 145 responses are being received on the modscan. This is resulting in reduced data availability on the APPCB server. | We have checked and observed that out of 275 poll requests,only 145 responses are being received on the modscan. This is resulting in reduced data availability on the APPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFCN25005431083 | 2025-04-15 | 2025-04-30 | Cust 692 | CHENNAI | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | Chlorine Analyzer-MAC 100 | MAC 100(1) | Sample value not matching with actual and founds the readings struck | Founds parameter changed - needs to reconfigrure factory setting | Closed | NO | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFCN25006561084 | 2025-04-17 | 2025-04-21 | Cust 636 | CHENNAI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | CX1000 Calibration | Visited the site and calibration done | visited the site on 18-4-25 | NO | Paid | Core | N |  | Analytical - CHAKAN
AFCN25017171085 | 2025-05-08 | 2025-05-11 | Cust 783 | CHENNAI | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | pH TDS ORP and mag meter checking | visited the site and checked the healthiness of field devices | visited the site on 8-5-25 | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN25023481086 | 2025-05-22 | 2025-05-24 | Cust 665 | CHENNAI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | COD BOD TSS Erratic reading | Visited the site, done zero calibration and taken the analyser online, provided training to the technical team | visited the site on 23-5-25 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN25024441087 | 2025-05-25 | 2025-05-29 | Cust 497 | CHENNAI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | MPA Calibration | visited the site and calibration done | visited the site on 27-5-25 | NO | Paid | Core | N |  | Analytical - CHAKAN
AFCN25025401088 | 2025-05-27 | 2025-06-01 | Cust 784 | CHENNAI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | PH CALIBRATION | Found pH sensor faulty non warranty | Visited the site 27-5-2025 | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN25029591089 | 2025-06-01 | 2025-06-05 | Cust 760 | CHENNAI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Aqua4Trans | Aqua4Trans(1) | pH ORP Meter calibration | Visited the site and noticed that Aqua4Trans display hanging issue happening continuously and the transmitter sent to pune for further analysis. | visited the site 0n 3-6-25 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFCN25032091090 | 2025-06-06 | 2025-06-16 | Cust 785 | CHENNAI | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | Conductivity transmitter not responding | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210145279 LR Dt - 14-JUN-2025 | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210145279 LR Dt - 14-JUN-2025 | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFCN25036041091 | 2025-06-17 |  | Cust 781 | CHENNAI | Analytical - CHAKAN | OPEN | CALIBRATION JOB | Aqua4Trans | Aqua4Trans(1) | pH Silica analyser calibration work |  |  | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23033531092 | 2023-06-01 | 2023-06-05 | Cust 786 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | The flow meter totalizer reading is not matching with the CPCB server. | Connectivity of flow meter done through separate USB converter. Now flow meter readings were matching on CPCB server. Details are in attached MOM. | Connectivity of flow meter done through saperate USB converter. Now flow meter readings were matched on CPCB server. Details are in attached MOM. | EXTENDED | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23033571093 | 2023-06-01 | 2023-06-02 | Cust 787 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | DATCON | DATCON(DATCON) | CPCB connectivity needs to be done. | CPCB connectivity has been completed and data is live on the server. | CPCB connectivity has been completed and data is live on the server. Hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23033721094 | 2023-06-01 | 2023-06-04 | Cust 788 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 2nd AMC visit | 2nd AMC visit completed | Working fine | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23033861095 | 2023-06-02 | 2023-06-06 | Cust 789 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | commissioning | MPA has been commissioned and found working fine | MPA has been commissioned and found working fine | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23034611096 | 2023-06-05 | 2023-06-08 | Cust 790 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera streaming issue on CPCB server. | Here the issue is from the customer network end and now the same has been resolved and camera streaming is live. | The issue has been resolved Hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23034631097 | 2023-06-05 | 2023-06-10 | Cust 791 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4Trans(1) | COMMISSIOING | Commissioning has been done. | Commissioning has been done. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23034691098 | 2023-06-05 | 2023-06-05 | Cust 717 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Lens broken, UV Power supply failed, Calibration job | Flowcell disc, UV power supply & MUVX replaced. Calibration done working fine. | Flowcell disc, UV power supply & MUVX replaced. Calibration done working fine. | YES | Paid | Core | N |  | Analytical - CHAKAN
AFDL23034851099 | 2023-06-05 | 2023-06-30 | Cust 792 | DELHI | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Installation guidelines of MPA | Installation guidelines of MPA given. | Installation guidelines given | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFDL23034891100 | 2023-06-05 | 2023-07-07 | Cust 793 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB and RSPCB data uploaded issue. | Our development team made some upgrades from the software end and now the data availability on the RSPCB server is almost 99 %. for ETP station. | The issue has been resolved and data is live hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23035891101 | 2023-06-06 | 2023-06-10 | Cust 794 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Transmitter Aquacon | Aquacon 2000(AQCN 2000 p) | Commissioning and mA issue. | Commissioning has been done and analyzer under observation. | Commissioning has been done and analyzer under observation. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23035891101 | 2023-06-06 | 2023-06-10 | Cust 794 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Transmitter Aquacon | Aquacon 2000(AQCN 2000 p) | Commissioning and mA issue. | Commissioning has been done and analyzer under observation. | Commissioning has been done and analyzer under observation. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23036291102 | 2023-06-07 | 2023-06-10 | Cust 795 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | DO meter showing error - recovering copy | try to resolve the issue on line but could not resolved hence same has been dispatched to pune works for further checking /repairing on chargeable basis. | try to resolve the issue on line but could not resolved hence same has been dispatched to pune works for further checking /repairing on chargeable basis. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFDL23036711103 | 2023-06-08 | 2023-06-09 | Cust 796 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | LocalHost data uploaded issue. | Not found any issue data is live. | Not found any issue data is live. Hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23036941104 | 2023-06-08 | 2023-06-12 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | TOC ANALYSER | FM 101 TOC Analyser(FM 101 TOC) | Visit organised for NOC certificate formalities with CISF & Fire & Safety department against the Invoice no. 553501282. | The NOC formalities for the final invoice have been completed and submitted to the CISF office, after confirmation of the certificate being completed, it will be collected from the CISF office. | The NOC formalities for the final invoice have been completed and submitted to the CISF office, after confirmation of the certificate being completed, it will be collected from the CISF office. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23037381105 | 2023-06-10 | 2023-06-14 | Cust 711 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | MPA troubleshooting | MPA troubleshooting done as per attached mail communication.now MPA is working fine. | MPA troubleshooting done as per attached mail communication.now MPA is working fine. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFDL23038301106 | 2023-06-12 | 2023-06-19 | Cust 798 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Display of Smartpro(1) | NEELKANTH FOOD AND BEVERAGES - AGRA KHASRA NO.325,VILLAGE BHAINYA, GONDA-KHAIR ROAD,TEHSIL IGLAS, DISTRICT ALIGARH-202124 Mr.Mridul Singhal - 96341 57767 | The pH reading problem has been resolved through telephonic support by Suneet Kumar with help of your site person Mr. Raman (Mobile No. 7082982858). We appreciate to Mr. Raman for supporting us. Now the same complaint is closing now. | pH reading problem has been resolved through telephonic support by Suneet Kumar with help of your site person Mr. Raman (Mobile No. 7082982858). We appreciate to Mr. Raman for supporting us. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFDL23038321107 | 2023-06-12 | 2023-06-15 | Cust 799 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Scheduled visit for online Analyzer against AMC. | Visit for Multi Parameter Analyser against AMC has completed, Analyser has been working fine. Details are in attached MOM. | Visit for Multi Parameter Analyser against AMC has completed, Analyser has been working fine. Details are in attached MOM. | EXTENDED | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23039051108 | 2023-06-13 | 2023-06-15 | Cust 800 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | UPPCB domain URL needs to be change. | We took the access to your system and changed the UPPCB domain URL and now data is live on UPPCB server. | Data is live on the UPPCB server hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23040101109 | 2023-06-15 | 2023-06-16 | Cust 801 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 1st AMC visit | 1st AMC visit completed | Working fine | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23040501110 | 2023-06-15 | 2023-06-17 | Cust 608 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Scheduled visit for ETP online Analyzer against AMC. | Visit for Multi Parameter Analyser against AMC has completed, Analyser has been working fine. Details are in attached MOM. | Visit for Multi Parameter Analyser against AMC has completed, Analyser has been working fine. Details are in attached MOM. | EXTENDED | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23040561111 | 2023-06-15 | 2023-06-15 | Cust 802 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading on CPCB and HSPCB server. | The issue has been resolved, data is uploading on CPCB and HSPCB server. | The issue has been resolved, data is uploading on CPCB and HSPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23040781112 | 2023-06-15 | 2023-06-16 | Cust 790 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 4th AMC visit | 4th AMC visit completed | Working fine | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23040991113 | 2023-06-16 | 2023-06-16 | Cust 803 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera Configuration on CPCB + RSPCB server. | As you have changed the camera IP so we have changed the same Camera IP on our configuration and now the camera streaming is live on the CPCB server. | Camera streaming is live on both servers hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23041171114 | 2023-06-16 | 2023-06-17 | Cust 786 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 1st CAMC visit | 1st CAMC visit completed | Working fine | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23041821115 | 2023-06-19 | 2023-06-21 | Cust 804 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Reconfiguration of software and online connectivity at Uklana STP, PHED NO.3 HISAR | The Software installation and CPCB + HSPCB connectivity for the Uklana site has been completed and data is live on the server. | Connectivity has been completed and data is live hence we are closing this complaint. | YES | No Trip | Core | Y | Product / equipment failure impacting safety of equipment, process or people. | Analytical - CHAKAN
AFDL23042421116 | 2023-06-19 | 2023-06-21 | Cust 786 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Emergency visit to attend TSS and pH reading issue | Resolved the issue after replacing RS485 card and 2nos EEPROM IC | Working fine | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23042641117 | 2023-06-20 | 2023-07-11 | Cust 805 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | Dissolved Oxygen Analyzer Panel mounted(AQUADO4000P) | received mail nagesh to arrange visit for commissioning of the DO transmitter | Issue of reading mismatch resolved after site calibration. Commissioning Done. | Issue of reading mismatch resolved after site calibration. Commissioning Done. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFDL23042651118 | 2023-06-20 | 2023-06-21 | Cust 786 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Required Changes in the configuration as RS485 - FM Card | For ETP Outlet, we have made changes in the configuration as per the FM-Make RS485 card and now data is uploading on the CPCB server. | we have made changes in the configuration as per the FM-Make RS485 card and now data is uploading on the CPCB server. Hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23043131119 | 2023-06-21 | 2023-06-23 | Cust 806 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data uploded issue. | Here issue form hardware side and same has been resolved and the data is live. | Here issue form the hardware side the same has been resolved and the data is live. Hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23043141120 | 2023-06-21 | 2023-06-24 | Cust 807 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data uploaded issue. | We have observed your site and now the data is uploaded continuously on the server. | We have observed your site and now the data is uploaded continuously on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23043261121 | 2023-06-21 | 2023-06-27 | Cust 808 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Electrode | CS42 SENSOR(1) | Installation of Conductivity sensor and Programming FMCS42 | Done the same , the Analyzer is working fine. | Provided the telephonic support for installation of Conductivity sensor and programming. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFDL23044361122 | 2023-06-23 | 2023-06-24 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | SALES PROMOTION | TOC ANALYSER | FM 101 TOC Analyser(FM 101 TOC) | Meeting to discuss FMTOC Analyzer visit, which are installed at DHDT Unit (BS-VI Project). | Visit for TOC Analyzer, meeting done with the customer. FMPL will send offer for CAMC / AMC for the same. | Visit for TOC Analyzer, meeting done with the customer. FMPL will send offer for CAMC / AMC for the same. | NO | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23044481123 | 2023-06-23 | 2023-06-23 | Cust 796 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Local Data uploaded issue. | Not found any issue hence we are closing this complaint. | Not found any issue hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23045701124 | 2023-06-25 | 2023-06-28 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Schedule visit for ETP#3 Analyzers as COD, Phenol & Sulphide Aanalyzers against AMC. | Visit for COD, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | Visit for COD, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | EXTENDED | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23045731125 | 2023-06-25 | 2023-07-02 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | visit for checking of 22 nos water quality analyzers at ETP | visit done from 26/06 to 30/06/2023 under ARC.details as per attached MOM. | visit done from 26/06 to 30/06/2023 under ARC.details as per attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23046301126 | 2023-06-26 | 2023-06-29 | Cust 807 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | UPPCB connectivity needs to be done. | The UPPCB connectivity has been completed and data is live. | Data is live on the UPPCB server hence we are closing this compliant. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23046901127 | 2023-06-27 | 2023-07-03 | Cust 802 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Amc visit | Done the AMC visit | AMC visit Done. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23047101128 | 2023-06-27 | 2023-07-05 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | AMC | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | Schedule visit for ETP#1 Analyzers as COD, TOC, Oil, Phenol & Sulphide Aanalyzers against AMC. | Visit for COD,TOC, Oil in Water, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | Visit for COD,TOC, Oil in Water, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | EXTENDED | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23048021129 | 2023-06-29 | 2023-07-17 | Cust 804 | DELHI | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | AQUADO Analyzer | AQUAMAX DOT(1) | urchased the DO Analysers against Inv. No. 553506097 dt. 12.10.2023. Inv. Copy is also attached for your reference.Please note 1 no. of DO Controller (Sr. No. U0482223605) is not working. | Material Received in Depot & Repaired Under Depot complaint CCS23D0499 | Material Received in Depot &amp;amp; Repaired Under Depot complaint CCS23D0499 | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFDL23048061130 | 2023-06-29 | 2023-06-30 | Cust 809 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | need to send log files to rspcb | We have observed that from yesterday , data is continuously updating on the RSPCB server.But we have made some changes , and will keep the system under observation. | We have observed that from yesterday , data is continuously updating on the RSPCB server.But we have made some changes , and will keep the system under observation. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23048871131 | 2023-06-30 | 2023-06-30 | Cust 804 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation needs to be done. | We took the access to the system and successfully did the software reinstallation + local + CPCB + HSPCB configuration and now data is live on all the servers. | We took the access to the system and successfully did the software reinstallation + local + CPCB + HSPCB configuration and now data is live on all the servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23048901132 | 2023-06-30 | 2023-07-03 | Cust 800 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | need to do software reinstallation and configuration. | we are closing the complaint. We will reopen the same once the hardware issue is resolved. | we are closing the complaint. We will reopen the same once the hardware issue is resolved. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23050001133 | 2023-07-02 | 2023-07-05 | Cust 810 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | visit for calibration of ETP analyzer | visit done .mom attached. | visit done .mom attached. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL23050261134 | 2023-07-03 | 2023-07-06 | Cust 559 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | RSPCB data upload issue. | Data is live on HSPCB server, hence we are closing this complaint. | Data is live on HSPCB server, hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23050441135 | 2023-07-03 | 2023-07-10 | Cust 709 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | MPA display was not on, Customer arranged spare and visit will be make. | mA output card, found faulty and replaced. Analyzer working ok. | mA output card, found faulty and replaced. Analyzer working ok. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23050751136 | 2023-07-03 | 2023-07-06 | Cust 811 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Material received against invoice no: FC/RF552410379, Needs visit for re-commissioning | Modbus repeater and Power supply module (24 to 5 V DC) Installation has been done, same has been working fine. | Modbus repeater and Power supply module (24 to 5 V DC) Installation has been done, same has been working fine. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23051001137 | 2023-07-04 | 2023-07-10 | Cust 800 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | DATCON | DATCON(DATCON) | Falcon Mini supplied to M/s Creamy Food Modbus output not working. | Faulty material sent for repair. | Attended | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFDL23051221138 | 2023-07-04 | 2023-07-07 | Cust 812 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data uploaded issue. | We have asked them to open the NTP port from their network end and currently the data is live. | We have asked them to open the NTP port from their network end and currently the data is live. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23051281139 | 2023-07-04 | 2023-07-06 | Cust 813 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | As per our telephonic discussion please arrange the installation of the received ETP Online monitoring system OA NO-5523111121, supplied by Forbes Marshall & EMF meter. OA NO-702316953 supplied by Krohne Marshall. | Installation and Commissioning have been done, and found working fine. | Installation and Commissioning have been done, and found working fine. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23051501140 | 2023-07-05 | 2023-07-06 | Cust 804 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | HSPCB data upload issue. | we took the access to your system and now the issue has been resolved and data is live on the HSPCB server. | we took the access to your system and now the issue has been resolved and data is live on the HSPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23051551141 | 2023-07-05 | 2023-07-09 | Cust 814 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | DO Transmitters | VISI pro DO(VISI pro DO) | visit for DO analyzer | visit done details in mom. | visit done details in MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23051581142 | 2023-07-05 | 2023-07-10 | Cust 815 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | VISI pro DO(VISI pro DO) | Calibration of DO Analyser at Haldiram Noida | Dome the commissioning of DO Analyzers | commssioned the Analyzers | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23051581142 | 2023-07-05 | 2023-07-10 | Cust 815 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | VISI pro DO(VISI pro DO) | Calibration of DO Analyser at Haldiram Noida | Done the commissioning of Analyzers | commssioned the Analyzers | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23052021143 | 2023-07-06 | 2023-07-07 | Cust 532 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | RSPCB data uploaded issue. | We took the access to the system and successfully updated the enviroconnect 6.4 patch 14 to enviroconnect 6.4 patch 49. And now data is live on all the servers. | We took the access to the system and successfully updated the enviroconnect 6.4 patch 14 to enviroconnect 6.4 patch 49. And now data is live on all the servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23052411144 | 2023-07-07 | 2023-07-31 | Cust 814 | DELHI | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | AQUADO Analyzer | AQUAMAX DOT(1) | reading not showing less than 1 PPM( decimal reading not appearing less than 1 PPM) | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210055428 LR Dt - 29-JUL-2023 | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210055428 LR Dt - 29-JUL-2023 | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFDL23052591145 | 2023-07-07 | 2023-07-08 | Cust 816 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data uploaded issue. | We took access and made some changes in the configuration from server end and now the data is live. | Data is live hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23052611146 | 2023-07-07 | 2023-07-10 | Cust 719 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | PH1 detector showing zero and OFFSET error. | Analyzer working healthy. MOM attached. | Analyzer working healthy. MOM attached. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL23053581147 | 2023-07-10 | 2023-07-15 | Cust 817 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Troubleshooting | attended all instruments. Details mentioned in MOM | attended all instruments. Details mentioned in MOM | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL23053861148 | 2023-07-10 | 2023-07-11 | Cust 818 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | UPPCB server url change | we took the access to the system and successfully updated the UPPCB server url, and now data is live on UPPCB server. | we took the access to the system and successfully updated the UPPCB server url, and now data is live on UPPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23055191149 | 2023-07-12 | 2023-07-14 | Cust 819 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | UPPCB domain url change. | we took the access to the BB, and successfully updated the UPPCB domain url. now data is live on UPPCB server. | we took the access to the BB, and successfully updated the UPPCB domain url. now data is live on UPPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23055681150 | 2023-07-12 | 2023-07-14 | Cust 820 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | visit for MPA | visit done .xenon lamp power supply card replaced .now analyzer is working fine. | visit done .xenon lamp power supply card replaced .now analyzer is working fine. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23056501151 | 2023-07-14 | 2023-07-16 | Cust 730 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC | AMC done. | AMC done. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23057891152 | 2023-07-17 | 2023-07-24 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | visit under AMC for 22 nos water Quality analyzers | visit done for 5 days as per attached MOM. | visit done for 5 days as per attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23058121153 | 2023-07-18 | 2023-07-21 | Cust 812 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data uploaded issue. | We have made some changes in the BB configuration and now the data is live on the server. | We have made some changes in the BB configuration and now the data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23058791154 | 2023-07-19 | 2023-07-25 | Cust 821 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MIL card of Multiparameter(1) | AMC-1 MPA | Checked its performance and carry out calibration, found working ok. | Checked healthiness of the analyzer and carried out the calibration, found working ok. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23059111155 | 2023-07-20 | 2023-07-21 | Cust 802 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB + HSPCB data upload issue. | We took the access to the system and after restarted enviroconnect services data is live on CPCB + HSPCB server. | We took the access to the system and after restarted enviroconnect services data is live on CPCB + HSPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23059711156 | 2023-07-21 | 2023-07-24 | Cust 821 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | UPPCB server url need to be change | Now data is live on UPPCB server, hence we are closing this complaint. | Now data is live on UPPCB server, hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23059741157 | 2023-07-21 | 2023-08-10 | Cust 822 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB connectivity needs to be done. | cpcb connec tovty is completed .for uppcb , ftp ports and url is blocked. | cpcb connec tovty is completed .for uppcb , ftp ports and url is blocked. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23060111158 | 2023-07-23 | 2023-07-28 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Schedule visit for ETP#3 Analyzers as COD, Phenol & Sulphide Aanalyzers against AMC. | Visit for COD, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | Visit for COD, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23060951159 | 2023-07-24 | 2023-07-27 | Cust 823 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB and RSPCB data uploaded issue. | We took access of the system and updated the jar file so now the data is uploading continuously on the RSPCB server. | The issue has been resolved hence we are closing this complaint,. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23061171160 | 2023-07-24 | 2023-07-29 | Cust 824 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | The calibration schedule of the EQMS analyser was due on 09.07.2023. You are requested to depute your service engineer immediately at site to calibrate and submit necessary calibration and healthiness report for the same. | Calibration done. MOM attached | Calibration done. MOM attached | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL23061431161 | 2023-07-25 | 2023-07-28 | Cust 812 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data uploaded issue. | we took access of the BB and updated it with the latest version so now the data is live on the server. | we took access of the BB and updated it with the latest version so now the data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23061661162 | 2023-07-25 | 2023-08-07 | Cust 807 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data uploaded issue. | Here we found a network issue from the customer end and now the same has been resolved and data is live on the server. | Here we found a network issue from the customer end and now the same has been resolved and data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23062191163 | 2023-07-26 | 2023-08-01 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | AMC | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | Schedule visit for ETP#1 Analyzers as COD, TOC, Oil, Phenol & Sulphide Aanalyzers against AMC. | Visit for COD, TOC, Oil, Phenol & Sulphide Aanalyzers against AMC has completed, Details are in attached MOM. | Visit for COD, TOC, Oil, Phenol & Sulphide Aanalyzers against AMC has completed, Details are in attached MOM. | EXTENDED | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23062301164 | 2023-07-26 | 2023-08-01 | Cust 814 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | calibration job for MPA through telephonic support | MPA Calibration done through telephonic support .mail communication attached. | MPA Calibration done through telephonic support .mail communication attached. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFDL23062491165 | 2023-07-26 | 2023-07-29 | Cust 811 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Troubleshooting | Replaced the faulty parts with new spares & found working fine. | Replaced the faulty parts with new spares & found working fine. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23063341166 | 2023-07-28 | 2023-08-01 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Visit to attend COD Analyzers of PTA ETP#3 Water Quality Analysers against as an emergency AMC visit. | Visit for COD Analysers against AMC Emergency visit has completed, Details are in attached MOM. | Visit for COD Analysers against AMC Emergency visit has completed, Details are in attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23064271167 | 2023-07-29 | 2023-08-01 | Cust 788 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | PH Sensor for Multiparameter(AWA-PH-SENS) | pH reading is not matching with the lab reading. | Resolve pH issue through telephonic support by our Suneet Kumar. Now it's working satisfactorily. Details are in attached mail correspondence. | Resolve pH issue through telephonic support by our Suneet Kumar. Now it's working satisfactorily. Details are in attached mail correspondence. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFDL23065411168 | 2023-07-31 | 2023-07-31 | Cust 600 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | AQUAMAX DOT(1) | DO Value was stuck at zero | DO meter calibration is done, it is under observation | Calibiration of DO meter is done, now it is under observation | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFDL23065431169 | 2023-07-31 | 2023-07-31 | Cust 825 | DELHI | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Installation guidelines | Installation guidelines given and pH sensor offer send. | Installation guidelines given, offer of pH sensor send | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFDL23065481170 | 2023-07-31 | 2023-07-31 | Cust 826 | DELHI | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC discussion and troubleshooting | pH is maintained and offer of AMC is send | pH is maintained , AMC offer send | YES | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFDL23065551171 | 2023-07-31 | 2023-08-03 | Cust 827 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Device add on CPCB server. | We took the access to the system and successfully completed the addition of 2 mass flow meters in the CPCB server, now data is live on CPCB server. | We took the access to the system and successfully completed the addition of 2 mass flow meters in the CPCB server, now data is live on CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23065621172 | 2023-07-31 | 2023-07-31 | Cust 828 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | pH reading high and Data transfer to CPCB | DATCON files was found corrupted | Datcon files was corrupted , send it pune | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFDL23065671173 | 2023-07-31 | 2023-07-31 | Cust 829 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | pH 1 and pH 2 issue. | MPA is to be send pune | MPA condition was very critical , need to send pune | YES | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFDL23065801174 | 2023-07-31 | 2023-08-02 | Cust 830 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | cpcb and hspcb data uploaded issue. | We made some changes in the configuration and now the data is live on the server. | We made some changes in the configuration and now the data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23066731175 | 2023-08-01 | 2023-08-18 | Cust 831 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MUVX Card of Multiparameter(1) | Multiparameter MUVX Card failed, need to supply on FOC Basis | New MUVX Card and xenon lamp power supply card supplied on FOC Basis and replaced the same at site, Now Multiparameter Analyser working fine. | New MUVX Card and xenon lamp power supply card supplied on FOC Basis and replaced the same at site, Now Multiparameter Analyser working fine. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFDL23066821176 | 2023-08-01 | 2023-08-07 | Cust 832 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Polymetron Phosphate Analyzer | 0 - 50 ppm PO4 measurement module 09210=A=1520(09210=A=1520) | Checking and calibration of Phosphate Analyser CX1000-201 | Checked the Phosphate Analyser and done the span calibration. | Checked the Phosphate Analyser and done the span calibration. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23066821176 | 2023-08-01 | 2023-08-07 | Cust 832 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Polymetron Phosphate Analyzer | 0 - 50 ppm PO4 measurement module 09210=A=1520(09210=A=1520) | Checking and calibration of Phosphate Analyser CX1000-201 | Checked the Phosphate Analyser and done the span calibration. | Checked the Phosphate Analyser and done the span calibration. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23066911177 | 2023-08-01 | 2023-08-07 | Cust 814 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | AQUADO Analyzer | AQUAMAX DOT(1) | visit for installation of new AQUAMAXDO | visit done & installed DPI 201 & Taken in to service & found working ok. | visit done & installed DPI 201 & Taken in to service & found working ok. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23066951178 | 2023-08-01 | 2023-08-04 | Cust 833 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue. | we took access of the router and updated the module of the router so now the data is live on the CPCB server. | we took access of the router and updated the module of the router so now the data is live on the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23066971179 | 2023-08-01 | 2023-08-03 | Cust 692 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue. | this is RSPCB server issue hence the customer taken up this issue with RSPCB, as of now we are closing this complaint. | this is RSPCB server issue hence the customer taken up this issue with RSPCB, as of now we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23067061180 | 2023-08-01 | 2023-08-04 | Cust 809 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 2nd AMC visit | 2nd AMC visit completed | Working fine | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23067061180 | 2023-08-01 | 2023-08-04 | Cust 809 | DELHI | Analytical - CHAKAN | CLOSED | AMC | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | 2nd AMC visit | 2nd AMC visit completed | Working fine | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23067071181 | 2023-08-01 | 2023-08-04 | Cust 834 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Troubleshooting/calibration/ replacement of spare | Replaced faulty items. During checking we datcon and xenon lamp power supply is not working. details mentioned in mom. | Replaced faulty items. During checking we datcon and xenon lamp power supply is not working. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL23067341182 | 2023-08-02 | 2023-08-03 | Cust 802 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | DATA upload issue on hspcb server | issue is resolved from customer end. and data was updating on hspcb server. | issue is resolved from customer end. and data was updating on hspcb server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23067491183 | 2023-08-02 | 2023-08-04 | Cust 835 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software installation and CPCB connectivity need to be done | We took the access to the system and successfully did the software installation + CPCB configuration, now data is live on CPCB server. | We took the access to the system and successfully did the software installation + CPCB configuration, now data is live on CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23068111184 | 2023-08-04 | 2023-08-07 | Cust 800 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | hardware issue is resolved and need to configure data on server. | We have done the configuration as per requirement and data is now live on the CPCB server. Values cross checked with you. | We have done the configuration as per requirement and data is now live on the CPCB server. Values cross checked with you. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23070001185 | 2023-08-07 | 2023-08-13 | Cust 786 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Schedule visit for ETP online Analyzer against CAMC. | Visit for Multi Parameter Analyser against CAMC has completed, Analyser has been working fine. Details are in attached MOM. | Visit for Multi Parameter Analyser against CAMC has completed, Analyser has been working fine. Details are in attached MOM. | EXTENDED | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23070251186 | 2023-08-08 | 2023-08-11 | Cust 804 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB connectivity needs to be done for 7.5 MLD STP Sirsa | We took access of the system and completed the CPCB connectivity and now the data is live on the server. | We took access of the system and completed the CPCB connectivity and now the data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23072741187 | 2023-08-14 | 2023-09-22 | Cust 836 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Chlorine sensor | CHLORINE SENSOR(1) | Chlorine range showing a maximum of 1.4 ppm but we required 4 to 6 ppm. Kindly arrange your technical person at our site For Your Reference Tax Invoice No :- 5545003256 Dated :- 13-07-2023 Below site address M/s.CRD Foods Address :- Village -Koshi Khurd NR.NK School,Bharatpur,Road, Mathura Uttar Pradesh - 281005 Kind Attn. : Mr.Dinesh Sharma Mob No :- 84710 30873 | Chlorine Analyzer sensor is 0-5PPM, According to customer process water 0-10PPM range sensor is required. | Chlorine Analyzer sensor is of range 0-5PPM but customers process water pH is above 10PPM | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFDL23072791188 | 2023-08-14 | 2023-08-17 | Cust 790 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera Connectivity needs to be done. | As of now we are closing this complaint and we will reopen it after getting system access from customers end. | As of now we are closing this complaint and we will reopen it after getting system access from customers end. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23073441189 | 2023-08-15 | 2023-08-18 | Cust 837 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Schedule visit for online Multi Parameter Analyzer in ETP against AMC | Visit for Multi Parameter Analyser against AMC has completed, Analyser has been working fine. Details are in attached MOM. | Visit for Multi Parameter Analyser against AMC has completed, Analyser has been working fine. Details are in attached MOM. | EXTENDED | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23073481190 | 2023-08-16 | 2023-08-17 | Cust 546 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 2nd AMC visit | 2nd AMC visit completed | Working fine | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23073821191 | 2023-08-16 | 2023-08-20 | Cust 824 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Kindly plan your visit to resolve the following problems in the EQMS Analyser & ETP Flow meter: TSS & PH need to be checked properly. Flowmeter of ETP to be calibrated. | Checked MPA, pH and Magnetic flow meter, found working fine, detailed mentioned in attached MOM. | Checked MPA, pH and Magnetic flow meter, found working fine, detailed mentioned in attached MOM. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL23073871192 | 2023-08-16 | 2023-08-19 | Cust 838 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning to be done of Multi Parameter Analyzer. | Commissioned the MPA, DO Analyzers, LT & Mag Flow Meters in 25 MLD STP, Panipat , observed all the instruments are working fine. Details are in attached MOM. | Commissioned the MPA, DO Analyzers, LT & Mag Flow Meters in 25 MLD STP, Panipat , observed all the instruments are working fine. Details are in attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23074101193 | 2023-08-17 | 2023-08-18 | Cust 532 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | RSPCB data upload issue. | Data is not coming till the local system checked on MODSCAN. as of now we are closing this complaint. | Data is not coming till the local system checked on MODSCAN. as of now we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23074121194 | 2023-08-17 | 2023-09-12 | Cust 839 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation needs to be done. | Data is live on all the servers, hence we are closing this complaint. | Data is live on RSPCB server, hence we are closing this compliant. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23074861195 | 2023-08-18 | 2023-09-11 | Cust 836 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Chlorine sensor | CHLORINE SENSOR(1) | Chlorine range showing a maximum of 1.4 ppm but we required 4 to 6 ppm. Kindly arrange your technical person at our site For Your Reference Tax Invoice No :- 5545003256 Dated :- 13-07-2023 Below site address M/s.CRD Foods Address :- Village -Koshi Khurd NR.NK School,Bharatpur,Road, Mathura Uttar Pradesh - 281005 Kind Attn. : Mr.Dinesh Sharma Mob No :- 84710 30873 | FM Engineer visited site and resolved the issue. | FM Engineer visited site and resolved the issue. Attached MOM | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFDL23074881196 | 2023-08-18 | 2023-08-21 | Cust 831 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | visit for checking/card installation & calibration of MPA | visit done MPA replaced card /resent done /calibration done & found working ok.MOM attached. | visit done MPA replaced card /resent done /calibration done & found working ok.MOM attached. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23075191197 | 2023-08-19 | 2023-08-21 | Cust 840 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data mismatch issue. | We took the access to the system and did some changes from configuration end and now data is matching to the actual value. | We took the access to the system and did some changes from configuration end and now data is matching to the actual value. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23075371198 | 2023-08-20 | 2023-08-22 | Cust 786 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Schedule visit for online Multi Parameter Analyzer in ETP against CAMC | Visit for Multi Parameter Analyser against CAMC has completed, Analyser has been working fine. Details are in attached MOM. | Visit for Multi Parameter Analyser against CAMC has completed, Analyser has been working fine. Details are in attached MOM. | EXTENDED | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23075431199 | 2023-08-20 | 2023-08-20 | Cust 841 | DELHI | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Discussion on AMC for MPA and burner | Discussion on AMC for MPA and burner done. | Discussion regarding to MPA AMC and burner | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFDL23075671200 | 2023-08-20 | 2023-08-20 | Cust 831 | DELHI | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | MULTIPARAMETER ANALYSER | MUVX Card of Multiparameter(1) | FM Engineer visit require for replacement of MUVX Card, Xenon Power supply card & PH sensor at site. | FM Engineer visited site and replaced MUVX Card, Xenon Power supply card & PH sensor with new and now MPA working fine. Attached MOM | FM Engineer visited site and replaced MUVX Card, Xenon Power supply card & PH sensor with new and now MPA working fine. Attached MOM | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFDL23076111201 | 2023-08-21 | 2023-08-29 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | visit for calibration of 22 nos water Quality analyzers at ETP. | visit done details in mom. | visit done details in mom. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL23076701202 | 2023-08-21 | 2023-08-23 | Cust 722 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data uploaded issue. | We took access of the system and updated the DAS locally and data is live on the server. | We took access of the system and updated the DAS locally and data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23076751203 | 2023-08-21 | 2023-08-23 | Cust 532 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | CPCB and RSPCB Data transmission and calibration of analyzer | CPCB and RSPCB data transmission restored & calibration completed | Working fine | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL23076781204 | 2023-08-21 | 2023-08-28 | Cust 716 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | ENVIROCONNECT Software PC was crashed. CPCB data not live. | MOM attached. | MOM attached | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL23077001205 | 2023-08-22 | 2023-08-28 | Cust 599 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | COMMISSIONING OF DO METER | MOM attached. In SMARTPro-23SDG047/Oxysense-1735 scratch found on sensor membrane. Presently, both analyzer working satisfactory. | MOM attached. In SMARTPro-23SDG047/Oxysense-1735 scratch found on sensor membrane. Presently, both analyzer working satisfactory. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23077221206 | 2023-08-22 | 2023-08-25 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Schedule visit for ETP#3 Analyzers as COD, Phenol & Sulphide Aanalyzers against AMC. | Visit for COD, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | Visit for COD, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23077511207 | 2023-08-23 | 2023-08-23 | Cust 793 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC visit required for Quater-3 | Engineer visited site and done the calibration. Report attached. | Resolved | NO | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFDL23078581208 | 2023-08-24 | 2023-08-31 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | AMC | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | Schedule visit for ETP#1 Analyzers as COD, TOC, Oil, Phenol & Sulphide Aanalyzers against AMC. | Visit for COD, TOC, Oil in Water, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | Visit for COD, TOC, Oil in Water, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | EXTENDED | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23078611209 | 2023-08-25 | 2023-08-27 | Cust 842 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | To rectify the analyzer with recently supplied spares | Rectified and calibrated the analyzer | Working fine | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL23079231210 | 2023-08-26 | 2023-08-28 | Cust 526 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | Checking and Calibration of pH and Conductivity Analyzer. | Checked all Ph and Conductivity Analyser and calibrate the same | Checked and calibrated the pH Analysers and Conductivity Analyzer | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23079311211 | 2023-08-26 | 2023-08-27 | Cust 843 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Display keypad issue and calibration | Resolved display keypad issue and calibrated the analyzer | Working fine | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23081241212 | 2023-08-29 | 2023-08-29 | Cust 844 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | DATCON | DATCON(DATCON) | CPCB data uploaded issue. | We took access of the DATCon and set the DNS in the datcon configuration so now the data is live on the server. | We took access of the DATCon and set the DNS in the datcon configuration so now the data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23082171213 | 2023-08-31 | 2023-09-01 | Cust 802 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation needs to be done. | In customers network the APPCB server url is blocked, we have convey the same to the customer and they are working on that, as of now we are closing this complaint. | In customers network the APPCB server url is blocked, we have convey the same to the customer and they are working on that, as of now we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23082361214 | 2023-08-31 | 2023-09-01 | Cust 806 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB and HSPCB data uploaded issue. | We took access of the system and restart the local services so now the data is live on the server. | We took access of the system and restart the local services so now the data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23082831215 | 2023-09-01 | 2023-09-03 | Cust 845 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | DO Transmitters | optex series(optex series) | Error -1 & 4 are displaying on DO Analyzers. | Visit for troubleshooting of DO Analyzer has complete, details are in attached MOM. | Visit for troubleshooting of DO Analyzer has complete, details are in attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23084011216 | 2023-09-04 | 2023-09-09 | Cust 788 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Schedule visit for online Multi Parameter Analyzer in ETP against AMC. | Visit for Multi Parameter Analyser against AMC has completed, Analyser has been working fine. Details are in attached MOM. | Visit for Multi Parameter Analyser against AMC has completed, Analyser has been working fine. Details are in attached MOM. | EXTENDED | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23084581217 | 2023-09-05 | 2023-09-07 | Cust 846 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation | Software installation has been completed and data is live on the server. | Software installation has been completed and data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23084751218 | 2023-09-05 | 2023-09-09 | Cust 788 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Schedule visit for online Multi Parameter Analyzer in ETP against AMC. | Visit for Multi Parameter Analyser against AMC has completed, Analyser has been working fine. Details are in attached MOM. | Visit for Multi Parameter Analyser against AMC has completed, Analyser has been working fine. Details are in attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23085591219 | 2023-09-08 | 2023-09-08 | Cust 839 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | RSPCB data upload issue. | after restarted the RTDAS services, now data is live on RSPCB server. | after restarted the RTDAS services, now data is live on RSPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23085631220 | 2023-09-08 | 2023-09-13 | Cust 795 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | recovery page error appear on screen .Emergency visit under AMC | visit done details in mom. | visit done details in MOM. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23085801221 | 2023-09-08 | 2023-09-08 | Cust 600 | DELHI | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | COD value was high | Installation giudelines given and AMC offer send | Installation guidelines given and AMC offer given | NO | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFDL23086061222 | 2023-09-09 | 2023-09-13 | Cust 810 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | visit for checking of MPA | visit done for MPA as per attached MOM. | visit done for MPA as per attached MOM. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL23086761223 | 2023-09-11 | 2023-09-18 | Cust 610 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | please provide support to install hardness analizer. Po & analizer photo attached. Site name:- haldiram snacks pvt LTD Sector - 63 , noida, uttar Pradesh Contact person: BIPIN DWIVEDI Mobile no :- 8115724613 | UPS supply was not at site for Hardness Analyzer due to which we were unable to power on the system. Now customer has confirmed about UPS connectivity. We will reopen the complaint and will visit at this on Wed. | Now customer has confirmed about UPS connectivity. We will reopen the complaint and will visit at this on Wed. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23087561224 | 2023-09-12 | 2023-10-31 | Cust 807 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data uploaded issue. | We have made the necessary changes and now the data is uploading continuously on the CPCB server | We have made the necessary changes and now the data is uploading continuously on the CPCB server | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23087641225 | 2023-09-12 | 2023-09-18 | Cust 847 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Chlorine Analyzer-MAC 100 | MAC 100(1) | End user--Moon Beverages, Hapur, Chlorine analyzer-- not working | Chlorine analyzer sensor range is 0-5PPM but according to process it should be 0-20PPM. when sensor of range 0-20PPM will be available at site i will visit again. pH issue has been resolved | I have visited and found the range of chlorine is 0-5PPM but it should be 0-20 PPM, as soon as sensor will be available at site we will visit again | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFDL23088481226 | 2023-09-14 | 2023-09-14 | Cust 800 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC visit pending for MPA. | Engineer visited site and done the MPA calibration. | Resolved | NO | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFDL23089081227 | 2023-09-15 | 2023-09-25 | Cust 848 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | 24 VDC SMPS (1) | Checking of MPA and Flow meter | MPA is working fine the Flow tube need to replace. | Checked the MPA and Flowmeter. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL23089181228 | 2023-09-15 | 2023-09-17 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | SALES PROMOTION | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | Visit for ETP Analyzers with our sales engineer Mr. Sunil Mor. | Visit completed for sales promotion with our sales engineer Mr Sunil Mor. We will share the offer of discussed proposal shortly. | Visit completed for sales promotion with our sales engineer Mr Sunil Mor. We will share the offer of discussed proposal shortly. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23089731229 | 2023-09-17 | 2023-09-21 | Cust 799 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Schedule visit for online Multi Parameter Analyzer in ETP against AMC. | Visit for Multi Parameter Analyser & Camera against AMC has completed, Analyser has been working fine. Details are in attached MOM. | Visit for Multi Parameter Analyser & Camera against AMC has completed, Analyser has been working fine. Details are in attached MOM. | EXTENDED | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23090691230 | 2023-09-19 | 2023-10-01 | Cust 849 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Electrode | Chemotrode 120(HT238 760) | installation guidelines & commissioning | Given the installation guidelines to the customer and commissioned the Inlet MPA. | Given the installation guidelines to the customer. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23090901231 | 2023-09-19 | 2023-09-24 | Cust 850 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Visit against receiving service PO for Multi Parameter Analyzer. | Attended Multi Parameter Analyser, Now its working satisfactorily. Details are in attached MOM. | Attended Multi Parameter Analyser, Now its working satisfactorily. Details are in attached MOM. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL23091291232 | 2023-09-20 | 2023-09-22 | Cust 799 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera streaming issue on CPCB | Took access of the system and changed IP address of Camera in the configuration. | Took access of the system and changed IP address of Camera in the configuration. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23091441233 | 2023-09-21 | 2023-09-24 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Schedule visit for ETP#3 Analyzers as COD, Phenol & Sulphide Aanalyzers against AMC. | Visit for COD, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | Visit for COD, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23091631234 | 2023-09-21 | 2023-09-23 | Cust 610 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | THERMO SERIES | AquaPro Analyzer(1) | visit for commissioning of Calcium Hardness Analyzer Make- Thermo scientific 2100 series SN:- 00017264 with Hisam (Process Engineer) | visit done details in attached mail communication. | visit done at Haldiram site noida for hardness analyzer.details as per attached mail communication. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23091771235 | 2023-09-21 | 2023-09-22 | Cust 851 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data uploaded issue. | The issue has been resolved now the data is live on the server. | The issue has been resolved now the data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23092041236 | 2023-09-22 | 2023-09-25 | Cust 796 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Aqua4Trans | Aqua4Trans(1) | Calibration of pH/TSS analyzer | Calibration of TSS and pH could not be done as AQUA4TRANS, TSS sensor and pH sensor cable were not working. Customer is sending the pH/TSS transmitter, TSS sensor and pH sensor cable to Pune | pH/TSS transmitter, TSS sensor and pH sensor cable are being sent to Pune by the customer | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL23092421237 | 2023-09-23 | 2023-10-01 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | visit for checking & calibration of ETP analyzers under AMC | visit done for 22 nos ETP Analyzers .details in mom. | visit done for 22 nos ETP Analyzers .details in mom. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23092611238 | 2023-09-24 | 2023-09-25 | Cust 816 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Aqua4Trans | Aqua4 trans-A4TW033000(1) | Calibration of pH/TSS Analyzer | Calibration of pH/TSS analyzer completed | Working fine | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL23092701239 | 2023-09-24 | 2023-09-28 | Cust 802 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Schedule visit for online Multi Parameter Analyzer in ETP against AMC. | Visit for Multi Parameter Analyser against AMC has completed, Analyser has been working fine. Details are in attached MOM. | Visit for Multi Parameter Analyser against AMC has completed, Analyser has been working fine. Details are in attached MOM. | EXTENDED | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23092721240 | 2023-09-25 | 2023-09-28 | Cust 852 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning of ETP Analyzer | Commissioning completed | Working fine | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23092761241 | 2023-09-25 | 2023-09-27 | Cust 802 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB + HSPCB data upload issue. | Now hardware issue is resolved from customer end and now data is live on CPCB + HSPCB server. | Now hardware issue is resolved from customer end and now data is live on CPCB + HSPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23093451242 | 2023-09-25 | 2023-09-29 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | AMC | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | Schedule visit for ETP#1 Analyzers as COD, T OC, Oil, Phenol & Sulphide Aanalyzers against AMC. | Visit for COD, TOC, Oil in Water, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | Visit for COD, TOC, Oil in Water, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | EXTENDED | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23093511243 | 2023-09-25 | 2023-09-30 | Cust 853 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 1 AMC of multiparameter | MOM attached | MOM attached | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23093721244 | 2023-09-26 | 2023-09-26 | Cust 841 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC Calibration Pending | Engineer Visited site and done the calibration of MPA | Resolved | NO | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFDL23093901245 | 2023-09-26 | 2023-10-01 | Cust 854 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Need to visit to commission black box for digital connectivity of MPA. | commissioned black box for digital connectivity of MPA. Data is successfully transferring to CPCB. | BB has been commissioned and MPA data is successfully transfering to CPCB | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFDL23093911246 | 2023-09-26 | 2023-10-01 | Cust 855 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | THERMO SERIES | AquaPro Analyzer(1) | Commissioning of Hardness Analyzer is to be done at site | Awaiting for material at site. | Awaiting for material at site. | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFDL23094551247 | 2023-09-27 | 2023-09-30 | Cust 827 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Reinstallation of software. | Successfully did the software reinstallation and all the configuration now data is live on CPCB server, but for UPPCB there is UPPCB server url and ftp port is block, hence as of now we are closing this complaint. | Successfully did the software reinstallation and all the configuration now data is live on CPCB server, but for UPPCB there is UPPCB server url and ftp port is block, | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23094821248 | 2023-09-28 | 2023-09-29 | Cust 786 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 2nd CAMC visit | 2nd CAMC visit completed | Working fine | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23096121249 | 2023-09-29 | 2023-10-02 | Cust 813 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB connectivity needs to be done. | we completed the CPCB connectivity but we are getting Exceptions in logs and we have asked the customer to check this issue with their network team hence as of now we are closing this complaint. | we completed the CPCB connectivity but we are getting Exceptions in logs and we have asked the customer to check this issue with their network team hence as of now we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23096331250 | 2023-09-29 | 2023-10-01 | Cust 814 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | PH Sensor for Multiparameter(AWA-PH-SENS) | visit for checking of PH reading response | visit done .details in mom. pH sensor was not working. | visit done .details in mom. pH sensor was not working. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23096491251 | 2023-09-29 | 2023-10-01 | Cust 801 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 2nd AMC visit | 2nd AMC visit completed | Working fine | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23098141252 | 2023-10-01 | 2023-12-12 | Cust 814 | DELHI | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | MULTIPARAMETER ANALYSER | CPU with LCD of Multiparameter Display(1) | display character missing with in 3 months after commissioning.Model-CX1000-3922,Analyzer s/n- 204150 | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210077845 LR Dt - 24-NOV-2023 | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210077845 LR Dt - 24-NOV-2023 | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFDL23098141252 | 2023-10-01 | 2023-12-12 | Cust 814 | DELHI | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | MULTIPARAMETER ANALYSER | PH Sensor for Multiparameter(AWA-PH-SENS) | PH sensor not working & fail with in 3 months after commissioning .model- Polilyte pro 120,S/n- 38142 | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210077845 LR Dt - 24-NOV-2023 | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210077845 LR Dt - 24-NOV-2023 | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFDL23098901253 | 2023-10-02 | 2023-11-03 | Cust 807 | DELHI | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | FM Black Box was found failure under warranty hence need to replace FMBB | FMBB dispatched through RF OA NO.: 5524800905, LR No - 210070890 LR Dt - 09-OCT-2023 | FMBB dispatched through RF OA NO.: 5524800905, LR No - 210070890 LR Dt - 09-OCT-2023 | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23099181254 | 2023-10-03 | 2023-10-08 | Cust 810 | DELHI | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Installation Guideline | CHECKED MATERIAL & INSTALLATION GUIDELINE WAS GIVEN TO THE CUSTOMER. | CHECKED MATERIAL & INSTALLATION GUIDELINE WAS GIVEN TO THE CUSTOMER. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23099191255 | 2023-10-03 | 2023-10-09 | Cust 608 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Checking and Calibration of Analyzer | Done the calibration of MPA. | Done the calibration of Multiparameter. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL23099291256 | 2023-10-03 | 2023-10-08 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | AMC | pH Transmitter Monec | Monec 8965 Ph Transmitter(8965) | Visit for truboolshooting in SWAS Analyzers as an emergency visit. | Calibration done through simulator of pH Transmitter, but there pH electrode was not responding. | Emergency visit for Oil in water Analyser against AMC has completed, Analyser has been working fine. Details are in attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23099291256 | 2023-10-03 | 2023-10-08 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | AMC | Silica Analyser | Silica 9210(9210) | Silica Analyzer was not in functioning. | Replaced LC Board, Display, OFC and Reagents in Silca Analyzer. Now Analyzer is functionly working satisfactorilly. Customer will prepaired R1 reagent and to be connected with the Analyzer for perfect perform. | Emergency visit for Oil in water Analyser against AMC has completed, Analyser has been working fine. Details are in attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23099291256 | 2023-10-03 | 2023-10-08 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | AMC | Conductivity Tranmitter Polymetron | Conductivity Tranmitter Polymetron 9125(9125) | Visit for truboolshooting in SWAS Analyzer as an emergency visit. | Calibration done through simulator, Analyzer is working fine. | Emergency visit for Oil in water Analyser against AMC has completed, Analyser has been working fine. Details are in attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23099361257 | 2023-10-04 | 2023-10-07 | Cust 856 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | SODIUM ANALYSER | SODIUM 9073 single channel ANLAYSER(9073 Sodium) | Service of SWAN make sodium and silica analyzer | visit done .details in attached mail communication. | visit done .details in attached mail communication. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23099371258 | 2023-10-04 | 2023-10-10 | Cust 709 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC-3rd | MOM attached | MOM attached. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23099711259 | 2023-10-04 | 2023-10-05 | Cust 497 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue. | Data is not coming till MODSCAN, after reconnecting of cable which is comes from hardware side, then data is live on CPCB server. | Data is not coming till MODSCAN, after reconnecting of cable which is comes from hardware side, then data is live on CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23101081260 | 2023-10-08 | 2023-10-08 | Cust 841 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | TSS value was showing low. | TSS in process water is under 35mg/lt and disc is B5 so it can detect. B10 disc needs to be use. | TSS in process water is under 35mg/lt and disc is B5 so it can detect. | YES | Periodic Maintenance (AMC) | Process | N |  | Analytical - CHAKAN
AFDL23101121261 | 2023-10-08 | 2023-10-14 | Cust 786 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Through Email against Emergency AMC visit for online Multi Parameter Analyzer due to abnormal spike in TSS reading. | Emergency visit for Multi Parameter Analyser against CAMC has completed, Analyser has been working fine. Details are in attached MOM. | Emergency visit for Multi Parameter Analyser against CAMC has completed, Analyser has been working fine. Details are in attached MOM. | EXTENDED | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23101141262 | 2023-10-08 | 2024-02-28 | Cust 810 | DELHI | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Fan damage for replacement. | Transit damage of Panel Cooling fan. Replacement sent to site | Transit damage of Panel Cooling fan. Replacement sent to site | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFDL23101941263 | 2023-10-09 | 2023-10-11 | Cust 857 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation needs to be done. | Successfully did the software reinstallation + Local + CPCB configuration, but currently data is not communicating with the local system, checked on MODSCAN, (Scanner faulty). | Successfully did the software reinstallation + Local + CPCB configuration, but currently data is not communicating with the local system, checked on MODSCAN, (Scanner faulty). | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23101971264 | 2023-10-09 | 2023-10-16 | Cust 858 | DELHI | Analytical - CHAKAN | CLOSED | SALES PROMOTION | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Visit to inspection of the online Multi-Parameter Analyzer. | Visit for Multi Parameter Analyser to check healthiness. Spares offer to be submitted shortly. | Visit for Multi Parameter Analyser to check healthiness. Spares offer to be submitted shortly.s completed, Analyser has been working fine. Details are in attached MOM. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23102221265 | 2023-10-10 | 2023-10-11 | Cust 790 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 1st AMC visit | 1st AMC visit completed | Working fine | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23103121266 | 2023-10-11 | 2023-10-13 | Cust 859 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Transmitter Aquacon | Aquacon 2000(AQCN 2000 p) | Calibration of Cond, pH | Calibration done & MOM attached | Calibration done and MOM attached | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL23103581267 | 2023-10-12 | 2023-10-19 | Cust 860 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Visit for MFM and Multi Parameter Analyzers calibration against PO. | Visit for Multi Parameter Analyser & Magnetic Flow Meter calibration against ARC has been completed, Analysers & Flow Meters have been working fine. Details are in attached MOM. | Visit for Multi Parameter Analyser & Magnetic Flow Meter calibration against ARC has been completed, Analysers & Flow Meters have been working fine. Details are in attached MOM. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL23103661268 | 2023-10-13 | 2023-12-19 | Cust 861 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | COMMISSIONING OF pH Transmitter | FM Engg. visited the site and found pH sensor was faulty, offer has been send and order is received awaiting for the material. | FM Engg. visited the site and pH sensor was found faulty. | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFDL23105391269 | 2023-10-16 | 2023-10-18 | Cust 730 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | MPA AMC 4 | MOM attached | MOM attached | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23105761270 | 2023-10-17 | 2023-10-20 | Cust 824 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | The calibration is due on date 25/10/23, required service engineer visit for the same. | Checked the healthiness and carried out the calibration of the analyzer found working Ok. | Checked the healthiness and carried out the calibration of the analyzer found working Ok. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL23106141271 | 2023-10-18 | 2023-10-20 | Cust 804 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning of MPA at 5 MLD STP Tohna | Done the commissioning, | Done the commissioning of MPA | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23106151272 | 2023-10-18 | 2023-10-20 | Cust 862 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Service visit against WO: 1011704329 | Checked the healthiness and carried out the calibration of the analyzer found working ok. | Checked the healthiness and carried out the calibration of the analyzer found working ok. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL23106191273 | 2023-10-18 | 2023-10-18 | Cust 863 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Display received at site after repairing, Recommissioning of Analyzer | Re-Commissioned the MPA | Commissioned the MPA Analyzer | NO | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23106231274 | 2023-10-18 | 2023-10-20 | Cust 863 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Online connectivity | Connectivity completed for HSPCB/CPCB | Connectivity completed | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23107401275 | 2023-10-20 | 2023-10-20 | Cust 802 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | HSPCB data upload issue. | Here is the apache service is stopped in local PC after restarted services data is live on HSPCB server. | Here is the apache service is stopped in local PC after restarted services data is live on HSPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23107841276 | 2023-10-23 | 2024-03-04 | Cust 864 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Chlorine Analyzer-MAC 100 | MAC 100(1) | poor performance of Forbes Marshall make Chlorine analyzer installed at our CAN line | Sensor Replaced. pH Compensation offered | Sensor Replaced. pH Compensation offered | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFDL23108141277 | 2023-10-23 | 2023-10-29 | Cust 820 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | visit for checking of MPA & it is getting rest | visit done for MPA checking as per attached MOM. | visit done for MPA checking as per attached MOM. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23109051278 | 2023-10-25 | 2023-10-28 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Schedule visit for ETP#3 Analyzers as COD, Phenol & Sulphide Aanalyzers against AMC. | Visit for COD, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | Visit for COD, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | EXTENDED | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23109181279 | 2023-10-26 | 2023-10-31 | Cust 803 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Addition of New Station on CPCB server | Inlet and outlet ETP flow meters are connected and data is continuously reflecting on the Local and CPCB servers for the same | Inlet and outlet ETP flow meters are connected and data is continuously reflecting on the Local and CPCB servers for the same | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23109391280 | 2023-10-26 | 2023-10-28 | Cust 812 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data uploaded issue. | Now the Data is live on the CPCB server for STP Outlet. | Now the Data is live on the CPCB server for STP Outlet. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23109461281 | 2023-10-26 | 2023-10-29 | Cust 821 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | EUS visit Against WO: 6000001233 | Checked the performance and carried out the calibration, found working ok. | Checked the performance and carried out the calibration, found working ok. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23109651282 | 2023-10-26 | 2023-11-02 | Cust 730 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | General healthiness checking of MPA | MOM attached. General healthiness checked | MOM attached. General healthiness checked | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23110151283 | 2023-10-27 | 2023-10-30 | Cust 865 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation needs to be done. | we successfully did the software reinstallation + local + CPCB configuration but data is not communicating with the local system (checked on MODSCAN). as of now we are closing this complaint. | we successfully did the software reinstallation + local + CPCB configuration but data is not communicating with the local system (checked on MODSCAN). as of now we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23110481284 | 2023-10-28 | 2023-11-01 | Cust 816 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Aqua4Trans | Aqua4Trans(1) | pH reading issue | Necessary support provided over the phone and video call, email attached. | Necessary support provided over the phone and video call, email attached. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFDL23110601285 | 2023-10-28 | 2023-11-03 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | AMC | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | Scheduled AMC visit for COD, TOC, Oil in Water, Sulphide Analyzers, installed at ETP # 1, 2 & TTPRO. | Visit for COD,TOC, Oil in Water & Sulphide Analysers against AMC has completed, Details are in attached MOM. | Visit for COD,TOC, Oil in Water & Sulphide Analysers against AMC has completed, Details are in attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23111171286 | 2023-10-29 | 2023-11-10 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | visit for checking & calibration of 22 nos ETP Analyzers. | visit done for all 22 nos analyzers under AMC .MOM attached. | visit done .MOM attached . | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23113141287 | 2023-10-31 | 2023-10-31 | Cust 800 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC Visit pending for MPA. | Engineer visited site and done the calibration. | Resolved | NO | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFDL23113151288 | 2023-10-31 | 2023-10-31 | Cust 841 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | MPA Breakdown needs an engineer visit. | Engineer visited the site and xeneon power supply card was faulty and replace this card with available spares. | Resolved | NO | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFDL23113181289 | 2023-10-31 | 2023-11-05 | Cust 866 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Service visit against WO-PSB/WC/05/BA | Checked the healthiness of the analyzer, found ok. | Checked the healthiness of the analyzer, found ok. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL23113361290 | 2023-10-31 | 2023-11-02 | Cust 866 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Local Connectivity issue. | The software installation and Local connectivity of your industry has been completed and data is live on the server. | The software installation and Local connectivity of your industry has been completed and data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23114301291 | 2023-11-01 | 2023-11-06 | Cust 799 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | TSS readings are fluctuating. | Issue resolved of STP Analyzer by our Suneet Kumar through telephonic support. Now the Multiparameter Analyzer is working properly. Therefore this complaint is now being closed. | Issue resolved of STP Analyzer by our Suneet Kumar through telephonic support. Now the Multiparameter Analyzer is working properly. Therefore this complaint is now being closed. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFDL23114921292 | 2023-11-02 | 2023-11-04 | Cust 849 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Installation guidelines and commissioning of analyzers | Provided installation guidelines and commissioned the analyzers as per the attached MOM | Working fine | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23115131293 | 2023-11-02 | 2023-11-02 | Cust 804 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation and Local + CPCB + HSPCB connectivity. | We took access to the system and reinstalled the software and completed the Local + CPCB + HSPCB connectivity. Now data is live on the servers. | We took access to the system and completed the Local + CPCB + HSPCB connectivity. Now data is live on the servers. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFDL23115651294 | 2023-11-04 | 2023-11-08 | Cust 809 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 3rd AMC visit | 3rd AMC visit completed | Working fine | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23115651294 | 2023-11-04 | 2023-11-08 | Cust 809 | DELHI | Analytical - CHAKAN | CLOSED | AMC | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | 3rd AMC visit | 3rd AMC visit completed | Working fine | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23116061295 | 2023-11-05 | 2023-11-08 | Cust 786 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Schedule visit for ETP online Analyzer against CAMC. | Visit for Multi Parameter Analyser against CAMC has completed, Analyser has been working fine. Details are in attached MOM. | Visit for Multi Parameter Analyser against CAMC has completed, Analyser has been working fine. Details are in attached MOM. | EXTENDED | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23116661296 | 2023-11-06 | 2023-11-08 | Cust 848 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Schedule visit for ETP online Analyzer against AMC at Mankind Research Center, Manesar. | Visit for Multi Parameter Analyser against AMC has completed, Analyser has been working fine. Details are in attached MOM. | Visit for Multi Parameter Analyser against AMC has completed, Analyser has been working fine. Details are in attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23116671297 | 2023-11-06 | 2023-11-08 | Cust 848 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Schedule visit for ETP online Analyzer against AMC at Mankind Research Center, Manesar. | Visit for Multi Parameter Analyser against AMC has completed, Analyser has been working fine. Details are in attached MOM. | Visit for Multi Parameter Analyser against AMC has completed, Analyser has been working fine. Details are in attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23117031298 | 2023-11-07 | 2023-11-08 | Cust 806 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not uploading on CPCB and HSPCB servers. | Data is now live on the server. | Data is now live on the server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL23117071299 | 2023-11-07 | 2023-11-07 | Cust 804 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation needs to be done. | We took the access to your system and did the software reinstallation + CPCB + HSPCB configuration. Now data is live on both the server | We took the access to your system and did the software reinstallation + CPCB + HSPCB configuration. Now data is live on both the server | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23117611300 | 2023-11-08 | 2023-11-19 | Cust 867 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Silica Analyser | Hach Silica 5500 series(Hach Silica 5500 series) | Please arrange to depute PLC engineer & Silica Analyzer engineer to Jindal Thekand site urgently. Commissioning & Plant hand over is held up. Contact person of Tekhand Site ? Mr. Radhe Shaym - 9711272411 ADDRESS :-TEHKHAND WASTE TO ELETRICITY PROJECT LIMITED GOVERNMENT LAND , NEAREXISTING SLF / TEHKHAND BUS DEPOT , TEHKHAND MA | Commissioned Silica Analyzer successfully. Observed Analyzer is working satisfactorily. Details are in attached MOM. | Commissioned Silica Analyzer successfully. Observed Analyzer is working satisfactorily. Details are in attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23118231301 | 2023-11-09 | 2023-11-27 | Cust 866 | DELHI | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | MULTIPARAMETER ANALYSER | pH input card-FM of Multiparameter(1) | Need EPROM with version 9.42 for newly received MPHDO card. | Replacement material has been booked through TCI EXPRESS vide LR No - 110007584385 LR Dt - 23-NOV-2023 | Replacement material has been booked through TCI EXPRESS vide LR No - 110007584385 LR Dt - 23-NOV-2023 | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFDL23118241302 | 2023-11-09 | 2023-11-10 | Cust 559 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating to RSPCB Server. | Now data is live on RSPCB server, hence we are closing this complaint. | Now data is live on RSPCB server, hence we are closing this complaint. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL23118301303 | 2023-11-09 | 2023-11-13 | Cust 868 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | visit for checking & calibration | visit done for calibration & checking of MPA & found working ok. | visit done for calibration & checking of MPA & found working ok. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL23118521304 | 2023-11-10 | 2023-11-11 | Cust 496 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | RSPCB data upload issue. | Now data is liv eon RSPCB server. | Now data is liv eon RSPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23119731305 | 2023-11-15 | 2023-11-22 | Cust 869 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | visit for checking & calibration of MPA | visit done .mom attached | visit done .MOM attached . | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL23120411306 | 2023-11-16 | 2023-11-18 | Cust 478 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | CPCB + RSPCB data upload issue. | here is internet issue, after the internet issue is resolved from customers end then data is now live on CPCB server. | here is internet issue, after the internet issue is resolved from customers end then data is now live on CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23120441307 | 2023-11-16 | 2023-11-18 | Cust 559 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | RSPCB data upload issue. | Apache service is stopped in customers local PC, after restart all the services now data is live on RSPCB server. | Apache service is stopped in customers local PC, after restart all the services now data is live on RSPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23120501308 | 2023-11-16 | 2023-11-27 | Cust 711 | DELHI | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | ETP-2 Blackbox frequently hanging so pl provide replacement. | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210077835 LR Dt - 23-NOV-2023 | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210077835 LR Dt - 23-NOV-2023 | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFDL23120561309 | 2023-11-16 | 2023-11-18 | Cust 807 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data uploaded issue. | We have checked and currently, the Data is live on the Central portal. | We have checked and currently, the Data is live on the Central portal. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23120871310 | 2023-11-17 | 2023-11-27 | Cust 870 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | visit for calibration of MPA | visit done .mom attached. | visit done .MOM attached . | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL23120881311 | 2023-11-17 | 2023-11-27 | Cust 871 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | visit for calibration of MPA | visit done .details in mom | visit done .details in mom | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL23122301312 | 2023-11-20 | 2023-11-23 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Schedule visit for ETP#3 Analyzers as COD, Phenol & Sulphide Aanalyzers against AMC. | Visit for COD, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | Visit for COD, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23122411313 | 2023-11-21 | 2023-11-23 | Cust 841 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | HSPCB data uploaded issue. | Now the issue has been resolved and data is live on the server. | Now the issue has been resolved and data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23123351314 | 2023-11-22 | 2023-11-27 | Cust 812 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Checking and calibration of Analyzer | Done the Calibration of MPA. | Done the calibration of Multiparameter. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23123351314 | 2023-11-22 | 2023-11-27 | Cust 812 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Checking and calibration of Analyzer | Done the Calibration of MPA | Done the calibration of Multiparameter. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23123361315 | 2023-11-22 | 2023-11-27 | Cust 837 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Checking and Calibration | Done the calibration of MPA | Done the calibration of Multiparameter. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23123411316 | 2023-11-22 | 2023-11-25 | Cust 786 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Schedule visit for ETP online Analyzer against CAMC. | Visit for Multi Parameter Analyser against CAMC has completed, Analyser has been working fine. Details are in attached MOM. | Visit for Multi Parameter Analyser against CAMC has completed, Analyser has been working fine. Details are in attached MOM. | EXTENDED | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23124191317 | 2023-11-24 | 2023-11-30 | Cust 872 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 2 AMC | MOM attached | MOM attached | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23124641318 | 2023-11-25 | 2023-11-28 | Cust 807 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data uploaded issue. | Data is live on the server. | Data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23124931319 | 2023-11-26 | 2023-12-19 | Cust 855 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | THERMO SERIES | AquaPro Analyzer(1) | Hardness Analyzer is not performing even after taken the help of Mr. Nitin Jain Thermo engineer, i have done 4 visits. pls do needful. | Mr. Nitin Jain has visited the site and done the calibration. system is under observation. | Thermo Engg. Mr. Nitin Jain has visited the site and done the calibration. system is under observation. | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFDL23124981320 | 2023-11-26 | 2023-11-26 | Cust 873 | DELHI | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | Conductivity Transmitter Aquacon | Aquacon 2000(AQCN 2000 p) | Install base checkup and discussion | Install base check up and discussion regarding offer | Install base checkup and discussion done | YES | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFDL23125641321 | 2023-11-27 | 2023-12-01 | Cust 532 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Local host page is not opening. | Took access to the system and found server was crashed so did the reconfiguration and now the data is live on CPCB+LOCAL+RSPCB. | Took access to the system and found server was crashed so did the reconfiguration and now the data is live on CPCB+LOCAL+RSPCB. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL23126421322 | 2023-11-27 | 2023-11-30 | Cust 869 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data uploaded issue. | the data is live on the CPCB server | the data is live on the CPCB server | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23126471323 | 2023-11-27 | 2023-12-02 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | AMC | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | Scheduled AMC visit for COD, TOC, Oil in Water, Sulphide Analyzers, installed at ETP # 1, 2 & TTPRO. | Visit for COD, TOC, Oil in Water, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | Visit for COD, TOC, Oil in Water, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23127131324 | 2023-11-28 | 2023-12-01 | Cust 841 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | HSPCB data uploded issue. | we have made some changes in the configuration and now the data is live on the server. | we have made some changes in the configuration and now the data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23127671325 | 2023-11-29 | 2023-12-05 | Cust 799 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software installation | As of now we are closing this complaint. We will reopen this once access is provided by customer. | As of now we are closing this complaint. We will reopen this once access is provided by customer. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23129151326 | 2023-11-30 | 2023-12-06 | Cust 874 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Service visit against WO no :4100074634 | Checked analyzer performance and found working fine. | Checked healthiness of the analyzers and carried out the calibration, found working ok. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL23129881327 | 2023-12-01 | 2023-12-02 | Cust 850 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | To check the analyzer | Cleaning, corrections, calibration carried out & analyzer working alright | Working fine | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL23130211328 | 2023-12-01 | 2023-12-04 | Cust 793 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB connectivity needs to be done. | we have completed the CPCB and RSPCB connectivity for both the Flow meters and data is live on the server. | we have completed the CPCB and RSPCB connectivity for both the Flow meters and data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23130461329 | 2023-12-03 | 2023-12-08 | Cust 802 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Ph Sensor calibration | Installation of new sensor and re-calibration. | Replacement of pH sensor and re calibration | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFDL23130571330 | 2023-12-03 | 2023-12-10 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | visit for checking & calibration of 22 nos ETP analyzers under AMC | visit was made by our Shubham Tinkhere from 24/11 to 02/12/2023 under AMC of month November . | visit was made by our Shubham Tinkhere from 24/11 to 02/12/2023 under AMC of month November . | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFDL23131221331 | 2023-12-04 | 2023-12-06 | Cust 804 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | 20 MLD STP MPA analyzer reading is not updating | Did some configuration changes and data is live on CPCB + Local + HSPCB servers. | Did some configuration changes and data is live on CPCB + Local + HSPCB servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23131911332 | 2023-12-05 | 2023-12-08 | Cust 807 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Multi-Parameter Analyzer is not working. | Attended the Multi-Parameter Analyzer, Now online MPA is working satisfactorily, after replaced spares. Details are in attached MOM. | Attended the Multi-Parameter Analyzer, Now online MPA is working satisfactorily, after replaced spares. Details are in attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23131951333 | 2023-12-06 | 2023-12-12 | Cust 801 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Emergency visit during RSPCB team visit | RSPCB team was shown the satisfactory functioning of the analyzer | Working fine | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23132031334 | 2023-12-06 | 2023-12-08 | Cust 874 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | Peristaltic sampling pump motor for Multiparameter(AWA-PSPM) | Require service visit against work order no 4100074374 | Checked the analyzer healthiness and carried out the calibration found ok. | Checked the analyzer healthiness and carried out the calibration found ok. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL23132151335 | 2023-12-06 | 2023-12-08 | Cust 804 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue | Data is not continuous on modscan. need to check the RS 485 PCB. As of now we are closing this complaint. We will reopen this once the issue from hardware side is resolved. | Data is not continuous on modscan. need to check the RS 485 PCB. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23132251336 | 2023-12-06 | 2023-12-10 | Cust 814 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | PH Sensor for Multiparameter(AWA-PH-SENS) | visit for replacement of PH sensor & PCX1 with Display Board under warranty | visit done .PH sensor replaced .New PCX1 board was getting reset issue hence could not replaced .MOM attached for refernce. | visit done .PH sensor replaced .New PCX1 board was getting reset issue hence could not replaced .MOM attached for refernce. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23132891337 | 2023-12-08 | 2023-12-12 | Cust 866 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | MPHDO card received need service engineer visit for the installation. | Replaced the MPHDO card and carried out the calibration, found working ok. | Replaced the MPHDO card and carried out the calibration, found working ok. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23133211338 | 2023-12-09 | 2023-12-14 | Cust 599 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Turbidity Input card of Aqua4Trans(1) | commissioning visit USER:GMADA, MOHALI | Commissioning done and MOM attached | MOM attached | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23133501339 | 2023-12-10 | 2023-12-13 | Cust 810 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | visit for calibration of MPA | visit done .MOM attached. | visit done .MOM attached . | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL23134071340 | 2023-12-11 | 2023-12-13 | Cust 875 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Electrode | FM1104(1) | Supplied pH Electrode is not working properly. | Discussed with the customer in detail and drop a mail. | Discussed with the customer in detail and suggest the solutions. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFDL23134331341 | 2023-12-11 | 2023-12-14 | Cust 876 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | New connectivity CPCB + UPPCB server connectivity. | We took the access to your system and did UPPCB configuration, and now we observed that the UPPCB url is blocked in your network, hence as of now we are closing this complaint. | We took the access to your system and did UPPCB configuration, and now we observed that the UPPCB url is blocked in your network, hence as of now we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23134451342 | 2023-12-11 | 2023-12-14 | Cust 717 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation | As of now we are closing this complaint. We will reopen this once the H/W issue is resolved. (there is an issue with the scanner) | As of now we are closing this complaint. We will reopen this once the H/W issue is resolved. (there is an issue with the scanner) | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23134531343 | 2023-12-11 | 2023-12-15 | Cust 877 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Service visit against OA no 5524106598 | Checked the healthiness of the analyzer and found working ok. | Checked healthiness of the analyzers and carried out the calibration, found working ok. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL23135181344 | 2023-12-13 | 2023-12-15 | Cust 878 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not Playaing on CPCB and RSPCB Server. | As of now we are closing this complaint as there is issue from customer end related to camera. | As of now we are closing this complaint as there is issue from customer end related to camera. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL23135361345 | 2023-12-13 | 2023-12-22 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | visit for calibration of 22 nos water quality analyzers | visit done for all 22 nos ETP analyzers. | visit done for all 22 nos ETP analyzers. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23135841346 | 2023-12-14 | 2023-12-17 | Cust 532 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Troubleshooting | Modbus Communication has been established. MPA is working fine | Modbus Communication has been established. MPA is working fine | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL23136041347 | 2023-12-14 | 2023-12-15 | Cust 532 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB+LOCAL+RSPCB | We took access of the system and data is live now and CPCB+MPCB+LOCAL | We took access of the system and data is live now and CPCB+MPCB+LOCAL | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL23136151348 | 2023-12-14 | 2023-12-15 | Cust 793 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | RSPCB data availability issue. | Now the Data Availability on the RSPCB server is almost 99 %. | Now the Data Availability on the RSPCB server is almost 99 %. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23136201349 | 2023-12-14 | 2023-12-15 | Cust 804 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation | We have completed the Local + CPCB + HSPCB connectivity and data is now live on the same servers. | We have completed the Local + CPCB + HSPCB connectivity and data is now live on the same servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23136711350 | 2023-12-16 | 2023-12-17 | Cust 879 | DELHI | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | MULTIPARAMETER ANALYSER | CPU Display card for Multiparameter(AWA-CPU1) | CPU Display became dull. | Display needs be replace. | Display needs to be replace and offer is send. | YES | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFDL23138001351 | 2023-12-18 | 2023-12-18 | Cust 831 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | MULTIPARAMETER ANALYSER | PH Sensor for Multiparameter(AWA-PH-SENS) | Data to RSPCB is going only 50% where as our OCEMS is working 24X7 | System given online to HO team changes have been done and system working fine and under observation. | System working. | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFDL23139021352 | 2023-12-19 | 2023-12-24 | Cust 716 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | BILL BOARD COMMISSIONING.---MAKE-VERTAXIS | cOMMISSIONED.. 3RD ROW DATA UPDATING ISSUE WILL BEE= ATTENDED BY OUR ho TEAM. | cOMMISSIONED.. 3RD ROW DATA UPDATING ISSUE WILL BEE= ATTENDED BY OUR ho TEAM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23141071353 | 2023-12-24 | 2024-01-01 | Cust 880 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | RS485 card-FM of Multiaparameter (1) | New Output card not showing data of actual remotely hence visit(CHargeable) | MOM attached | MOM attached | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL23141911354 | 2023-12-25 | 2023-12-31 | Cust 881 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4Trans(1) | Commissioning of pH and Chlorine Analyser | Commissioned the Analysers | Done the commissioning of Analyzers | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23141961355 | 2023-12-26 | 2023-12-29 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Schedule visit for ETP#3 Analyzers as COD, Phenol & Sulphide Analyzers against AMC. | Visit for COD, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | Visit for COD, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23142131356 | 2023-12-26 | 2024-01-02 | Cust 882 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | received mail from customer to attend the problem of pH and conductivity @ tata motors Lucknow. | Visited site and suggested to perform 2 point calibration & temperature to auto from manual. | Visited site and suggested to perform 2 point calibration & temperature to auto from manual. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23142351357 | 2023-12-26 | 2024-01-02 | Cust 875 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Electrode | FM1104(1) | Supplied pH sensor is not working properly. | Spoken to customer and guided him on the problem. | Spoken to customer and guided him on the problem. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFDL23142371358 | 2023-12-26 | 2024-01-02 | Cust 874 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | depute service engineer for calibration and certification of analyser as early possible. | Analyzer calibration done. | Analyzer calibration done. | YES | Paid | Core | N |  | Analytical - CHAKAN
AFDL23142571359 | 2023-12-26 | 2024-01-01 | Cust 821 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Visit for multiparameter AMC | Analyzer is working ok & calibration done. | System is working fine & calibration done | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23143321360 | 2023-12-27 | 2024-01-03 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | AMC | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | Scheduled AMC visit for COD, TOC, Oil in Water, Sulphide Analyzers, installed at ETP # 1, 2 & TTPRO. | Visit for COD, TOC, Oil in Water, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | Visit for COD, TOC, Oil in Water, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23143991361 | 2023-12-28 | 2023-12-30 | Cust 533 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB and RSPCB data uploaded issue. | We have checked and ETP data is not communicating with the Local PC. Hence need to check this issue from the hardware site so as of now we are closing this complaint. | We have checked and ETP data is not communicating with the Local PC. Hence need to check this issue from the hardware site so as of now we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23145281362 | 2023-12-30 | 2024-01-01 | Cust 608 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | To check multiparameter analyzer | Cleaning and calibration carried out for COD/BOD and pH and spares are required to calibrate TSS, offer of required spares to be submitted | Working fine except TSS, required spares offer to be submitted | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23145661363 | 2023-12-31 | 2024-01-04 | Cust 802 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration of MPA. | Done the calibration of MPA. | Done the calibration of Multiparameter. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23146381364 | 2023-12-31 | 2024-01-01 | Cust 870 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | on line support for Multiparameter ananlyzer as analyzer COD value is showing 0.0 PPM | COD calibration done through online support & found working fine as per attached mail communication . | COD calibration done through online support &amp;amp; found working fine as per attached mail communication . | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFDL23147051365 | 2024-01-02 | 2024-01-05 | Cust 716 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB CENTRAL ISSUE | We have taken access to check this issue, it is observed that your system timing was running ahead of actual time hence data upload to the CPCB central server stopped. Now we have corrected the time settings in your system manually and the data is live on the server. | Now we have corrected the time settings in your system manually and the data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23147071366 | 2024-01-02 | 2024-01-29 | Cust 883 | DELHI | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | Ph Electrode | FM1106(1) | The probe is malfunctioning and showing non linear readings. | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210078346 LR Dt - 08-JAN-2024 | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210078346 LR Dt - 08-JAN-2024 | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFDL23147921367 | 2024-01-03 | 2024-01-05 | Cust 842 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB central portal | We took access to the system and data is live now on CPCB Portal. | We took access to the system and data is live now on CPCB Portal. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL23148001368 | 2024-01-03 | 2024-01-05 | Cust 839 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB and RSPCB | Took access to the system and data is not updating on CPCB and RSPCB. | Took access to the system and data is not updating on CPCB and RSPCB. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL23148041369 | 2024-01-03 | 2024-01-08 | Cust 533 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | RS485 card-FM of Multiaparameter (1) | To check the analyzer and connectivity with CPCB/RSPCB | Checked the analyzer and restored the CPCB/RSPCB connectivity | Working fine | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL23148261370 | 2024-01-04 | 2024-01-05 | Cust 533 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | RSPCB data uploaded issue. | Now the issue has been resolved and data is live on the RSPCB server. | Now the issue has been resolved and data is live on the RSPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23148531371 | 2024-01-04 | 2024-01-07 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | AMC | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | Three months visitor Gate pass formalities against Analyzer AMC. | Three months visitor Gate pass formalities against Analyzer AMC has been done. | Three months visitor Gate pass formalities against Analyzer AMC has been done. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23148561372 | 2024-01-04 | 2024-01-08 | Cust 848 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Schedule visit for ETP online Analyzer against AMC at Mankind Research Center, Manesar. | Visit for online Multi-Parameter Analyzer has been completed, observed Analyzer is working satisfactorily. Details are in attached MOM. | Visit for online Multi-Parameter Analyzer has been completed, observed Analyzer is working satisfactorily. Details are in attached MOM. | EXTENDED | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23148711373 | 2024-01-04 | 2024-01-13 | Cust 848 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Checking and calibration of MPA. | Done the calibration, | Done the calibration of Multiparameter. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL23148971374 | 2024-01-05 | 2024-01-05 | Cust 839 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on RSPCB Server. | Took access to the system and data is not updating on CPCB and RSPCB. | Took access to the system and data is not updating on CPCB and RSPCB. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL23149141375 | 2024-01-05 | 2024-01-15 | Cust 692 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | RSPCB data upload issue. | we took the access to the system and updated RTDAS the version, and Now the data is continuously uploading on CPCB + RSPCB server. | we took the access to the system and updated RTDAS the version, and Now the data is continuously uploading on CPCB + RSPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23149251376 | 2024-01-05 | 2024-01-15 | Cust 884 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | FM make pH card(1) | PH card failure | PH card sent by client for testing & repairing to our pune works other wise need to procure new PH card . | PH card sent by client for testing &amp;amp; repairing to our pune works other wise need to procure new PH card .communication attached for reference. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFDL23149591377 | 2024-01-06 | 2024-01-08 | Cust 849 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning of Multiprameter and other analyzers | Commissioning work done | Working fine | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23149621378 | 2024-01-06 | 2024-01-11 | Cust 876 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Customer reported for commissioning of CX2000 Multiparameter analyser | Commissioning done. Analyser working ok. | Commissioning done analyser found working ok | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23149631379 | 2024-01-07 | 2024-01-08 | Cust 717 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration | MOM attached | MOM attached | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL23149651380 | 2024-01-07 | 2024-01-21 | Cust 856 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT UPGRADE | SODIUM ANALYSER | Transmitter(08941=A=0000) | Spares upgradation and calibration | visit made on 4, 8, 10 jan 24 alone conductivty high and no sample issue not resolved, so visit made with Mr. Vinay chaudhary on 18 & 19 jan 24. Both MOM attached. Work completed | MOM attached | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL23149981381 | 2024-01-08 | 2024-01-12 | Cust 843 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | CAMC visit | CAMC visit completed | Working fine | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23150421382 | 2024-01-08 | 2024-01-11 | Cust 839 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB+RSPCB | We took access to the system and data is now updating on RSPCB and CPCB. | We took access to the system and data is now updating on RSPCB and CPCB. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL23150621383 | 2024-01-08 | 2024-01-13 | Cust 799 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Schedule visit for online Multi Parameter Analyzer in ETP against AMC. | Visit for online Multi-Parameter Analyzer has been completed, observed Analyzer is working satisfactorily. Details are in attached MOM. | Visit for online Multi-Parameter Analyzer has been completed, observed Analyzer is working satisfactorily. Details are in attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23150771384 | 2024-01-09 | 2024-01-12 | Cust 469 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB + RSPCB data upload issue. | We did software reinstallation + Local + CPCB + RSPCB configuration, and now data is live on all the server. | We did software reinstallation + Local + CPCB + RSPCB configuration, and now data is live on all the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23150991385 | 2024-01-09 | 2024-01-17 | Cust 730 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC 5th | MOM attached | MOM attached | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23151221386 | 2024-01-09 | 2024-01-12 | Cust 799 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation. | We have re-installed the software and completed the Local connectivity. Data is live on local server. | We have re-installed the software and completed the Local connectivity. Data is live on local server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23151531387 | 2024-01-10 | 2024-06-28 | Cust 798 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Site Address: Granary Beverages LLP Address : KHASRA NO. 699-704, VILLAGE ASAUDHA, SONEPAT ROAD, BAHADURGARH, Sankhol, Jhajjar, Haryana,124505 | Sanjay visited the site and has calibrated the pH sensor and there are some difference in the actual pH and our meter readings. MOM made. | Sanjay visited the site and has calibrated the pH sensor and there are some difference in the actual pH and our meter readings. MOM made. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFDL23153361388 | 2024-01-14 | 2024-01-17 | Cust 863 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Data not transfer online/ | Visit done and sent the quotation to customer. | Visit done and sent the quotation to customer. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23154301389 | 2024-01-15 | 2024-01-20 | Cust 885 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Visit for spares installation and recalibration of online Multi-Parameter Analyzer. | Visit for online Multi-Parameter Analyzer has been completed, observed Analyzer is working satisfactorily. Details are in attached MOM. | Visit for online Multi-Parameter Analyzer has been completed, observed Analyzer is working satisfactorily. Details are in attached MOM. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL23154731390 | 2024-01-16 | 2024-01-19 | Cust 839 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on RSPCB | We took access to the system and data is live now on RSPCB Server. | We took access to the system and data is live now on RSPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL23155021391 | 2024-01-16 | 2024-01-22 | Cust 856 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | SODIUM ANALYSER | SODIUM 9245 Multi Channel ANLAYSER(9245 Sodium) | visit for troubleshooting of SWAN based sodium analyzer soditrace. | visit done .MOM attached for reference. | visit done .MOM attached for reference. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL23156091392 | 2024-01-19 | 2024-01-20 | Cust 815 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | pH Transmitter Aquamon | Ph Transmitter Aquamon 2000(PH2000) | To check, pH - 5 nos. Cond/TDS - 4 nos. ORP - 2 nos. and DO - 2 nos. for ETP and STP at Lenskart Solutions Pvt. Ltd, Bhiwadi site | Attended as per attached MOM - 1 no. pH Tx and 1 no. DO sensor to be sent to Pune by customer for necessary checks and repairs. | Working fine except one 1 DO and 1 pH Tx which customer will send to Pune for checks and repairs | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23156121393 | 2024-01-19 | 2024-01-23 | Cust 849 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning of Multiparameter and other analyzers | Work done as per attached MOM | Working fine | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23156301394 | 2024-01-19 | 2024-01-19 | Cust 869 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB Server. | Took access to the system and data is live on CPCB Server | Took access to the system and data is live on CPCB Server | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL23157971395 | 2024-01-23 | 2024-01-25 | Cust 886 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not updating on HSPCB. | Data is live on the HSPCB server. | Data is live on the HSPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23158291396 | 2024-01-23 | 2024-01-26 | Cust 790 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 2nd AMC visit | 2nd AMC visit completed | Working fine | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23158361397 | 2024-01-23 | 2024-01-25 | Cust 887 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software re- installation and Local + CPCB + HSPCB connectivity | We have re-installed the software with the updated version and completed Local + CPCB + HSPCB configuration and the data is live on the same servers. | We have re-installed the software with the updated version and completed Local + CPCB + HSPCB configuration and the data is live on the same servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL23158441398 | 2024-01-23 | 2024-01-30 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Schedule visit for ETP#3 Analyzers as COD, Phenol & Sulphide Analyzers against AMC. | Visit for COD, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | Visit for COD, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23158491399 | 2024-01-24 | 2024-01-31 | Cust 709 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 1AMC | MOM attached | MOM attached | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23158501400 | 2024-01-24 | 2024-01-31 | Cust 853 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC-2nd | AMC done | MOM attached | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23158631401 | 2024-01-24 | 2024-01-31 | Cust 888 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | AQUAMAX DOT(1) | COMMISSIOINING guideline | installation & pre-commissioning guideline given | MOM attached | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23159011402 | 2024-01-24 | 2024-02-05 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | visit for calibration & checking of 22 nos WQ analyzers | visit done under AMC | visit done .MOM attached . | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL23159341403 | 2024-01-25 | 2024-01-31 | Cust 882 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Smartpro-con) | Customer reported for checking of 6 nos of Conductivity/PH transmitters installed at TATA motors Lucknow. | Calibration done. | Calibration done | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23159341403 | 2024-01-25 | 2024-01-31 | Cust 882 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Smartpro-con) | Customer reported for checking of 6 nos of Conductivity/PH transmitters installed at TATA motors Lucknow. | Calibration done. | Calibration done | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23159341403 | 2024-01-25 | 2024-01-31 | Cust 882 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Smartpro-con) | Customer reported for checking of 6 nos of Conductivity/PH transmitters installed at TATA motors Lucknow. | Calibration done. | Calibration done | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23159341403 | 2024-01-25 | 2024-01-31 | Cust 882 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Smartpro-con) | Customer reported for checking of 6 nos of Conductivity/PH transmitters installed at TATA motors Lucknow. | Calibration done. | Calibration done | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23159341403 | 2024-01-25 | 2024-01-31 | Cust 882 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Smartpro-con) | Customer reported for checking of 6 nos of Conductivity/PH transmitters installed at TATA motors Lucknow. | Calibration done. | Calibration done | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23159341403 | 2024-01-25 | 2024-01-31 | Cust 882 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Smartpro-con) | Customer reported for checking of 6 nos of Conductivity/PH transmitters installed at TATA motors Lucknow. | Calibration done. | Calibration done | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23160381404 | 2024-01-29 | 2024-02-01 | Cust 839 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on RSPCB Server. | We took access to your system and now data is live on RSCB Server.Please find the attached screenshot for your reference. | We took access to your system and now data is live on RSCB Server.Please find the attached screenshot for your reference. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL23160461405 | 2024-01-29 | 2024-01-31 | Cust 839 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | RSPCB data upload issue. | After restart the enviroconnectDAS service, now data is live on RSPCB server. | After restart the enviroconnectDAS service, now data is live on RSPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23160711406 | 2024-01-29 | 2024-02-04 | Cust 834 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration | Datcon configuration has been done. ph2 detector is not working suggested to the customer to send the same to our ho for repair if possible. | Datcon configuration has been done. ph2 detector is not working suggested to the customer to send the same to our ho for repair if possible. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL23161221407 | 2024-01-30 | 2024-02-02 | Cust 546 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 3rd AMC visit | 3rd AMC visit completed | Working fine | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23161761408 | 2024-01-30 | 2024-02-04 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | AMC | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | Scheduled AMC visit for COD, TOC, Oil in Water, Sulphide Analyzers, installed at ETP # 1, 2 & TTPRO. | Visit for COD, TOC, Oil in Water, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | Visit for COD, TOC, Oil in Water, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23162191409 | 2024-01-31 | 2024-02-02 | Cust 889 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Chlorine sensor | CHLORINE SENSOR(1) | Installation & Commissioning of FRC Analyzer and Electromagnetic Flow Meter | 2 nos. FRC Analyzers and Electromagnetic Flow Meters have been commissioned. 1no. FRC sensor Sr. no. 1223 was found faulty and carried for necessary checking / replacement | Working fine | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23162741410 | 2024-01-31 | 2024-02-02 | Cust 834 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | DATCON | DATCON(DATCON) | Datcon configuration needs to be done. | we have completed the DATCon configuration and data is live on the CPCB server. | we have completed the DATCon configuration and data is live on the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23162821411 | 2024-01-31 | 2024-02-02 | Cust 816 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | camera is not live on RSPCB | As of now we are closing this complaint, we will reopen this once the issue is resolved from customer end. | As of now we are closing this complaint, we will reopen this once the issue is resolved from customer end. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23164631412 | 2024-02-01 | 2024-02-11 | Cust 872 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Emergency visit after 2 AMC for spare replacement | Emergency visit mom attached | Emergency visit MOM attached. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23164651413 | 2024-02-01 | 2024-02-04 | Cust 890 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | commissioning at goraya stp | Did MPA commissioning, Connectivity pending due to docs & mA card not supplied with it. | MOM attached | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23164711414 | 2024-02-02 | 2024-02-05 | Cust 532 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | RSPCB data availability issue. | As of now we are closing this complaint, and will reopen it after the hardware side issue is resolved. | As of now we are closing this complaint, and will reopen it after the hardware side issue is resolved. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23164721415 | 2024-02-02 | 2024-02-05 | Cust 891 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | HSPCB data upload issue (Need to reinstall software) | We took the access to your system and successfully did software reinstallation + CPCB + HSPCB configuration. And now data is live on CPCB + HSPCB server. | We took the access to your system and successfully did software reinstallation + CPCB + HSPCB configuration. And now data is live on CPCB + HSPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23164961416 | 2024-02-02 | 2024-02-02 | Cust 804 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Addition of new device | We have completed the CPCB and Local connectivity and the data is now live on the same servers. | We have completed the CPCB and Local connectivity and the data is now live on the same servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23164971417 | 2024-02-02 | 2024-07-01 | Cust 892 | DELHI | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | DO METER IS NOT WORKING | Previously visit planned but customer engineer not available so guided to install the DO sensor. Customer started to take sensor in process. | Previously visit planned but customer engineer not available so guided to install the DO sensor. Customer started to take sensor in process. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFDL23164981418 | 2024-02-02 | 2024-07-01 | Cust 892 | DELHI | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | DO TRANSMITTER FAILURE | Previously visit planned but customer engineer not available so guided to install the DO sensor. Customer started to take sensor in process. | Previously visit planned but customer engineer not available so guided to install the DO sensor. Customer started to take sensor in process. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFDL23164991419 | 2024-02-02 | 2024-07-01 | Cust 892 | DELHI | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | DO TRANSMITTER FAILURE | Previously visit planned but customer engineer not available so guided to install the DO sensor. Customer started to take sensor in process. | Previously visit planned but customer engineer not available so guided to install the DO sensor. Customer started to take sensor in process. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFDL23165011420 | 2024-02-02 | 2024-02-11 | Cust 599 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Turbidity Meter | Thermo series(Thermo series) | TURBIDITY METER NOT WORKING | MOM attached | MOM attached | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23165411421 | 2024-02-04 | 2024-02-08 | Cust 786 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Schedule visit for ETP online Analyzer against CAMC. | Visit for online Multi-Parameter Analyzer has been completed, against CAMC, observed Analyzer is working satisfactorily. Details are in attached MOM. | Visit for online Multi-Parameter Analyzer has been completed, against CAMC, observed Analyzer is working satisfactorily. Details are in attached MOM. | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23165711422 | 2024-02-05 | 2024-04-29 | Cust 893 | DELHI | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | Ph Electrode | Chemotrode 120(HT238 760) | FAULTY ELECTRONICS OF PH METER | Replacement material has been booked through TCI EXPRESS vide LR No - 110007585203 LR Dt - 13-APR-2024 | Replacement material has been booked through TCI EXPRESS vide LR No - 110007585203 LR Dt - 13-APR-2024 | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFDL23165931423 | 2024-02-05 | 2024-02-08 | Cust 804 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not updating to CPCB server. | Data is now live on the CPCB server. | Data is now live on the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23166431424 | 2024-02-06 | 2024-02-09 | Cust 804 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | DATCON | DATCON(DATCON) | CPCB and HSPCB data uploaded issue. | The issue has been resolved and now the data is live on both servers. | The issue has been resolved and now the data is live on both servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23166721425 | 2024-02-07 | 2024-02-08 | Cust 849 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | Xenon Lamp Power supply for Multiparameter(AWA-XLAMPPS) | Commissioning of water quality analyzers | Work done as per attached MOM | Working fine | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23169691426 | 2024-02-13 | 2024-02-16 | Cust 891 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation + CPCB + HSPCB configuration needs to be done. | we took access to your system and did software reinstallation + local + CPCB + HSPCB configuration And now data is live on all the servers. | we took access to your system and did software reinstallation + local + CPCB + HSPCB configuration And now data is live on all the servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23170191427 | 2024-02-14 | 2024-02-21 | Cust 862 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Customer reported for calibration of CX1000 Multiparameter analyzer & one mag flow meter at there site. | Calibration done. | Calibration done. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL23170211428 | 2024-02-14 | 2024-02-16 | Cust 894 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Reinstallation of Software PC formatted by the customer. | We took access to the system and successfully completed the installation and configuration of software now data is live on LOCAL+CPCB+HSPCB | We took access to the system and successfully completed the installation and configuration of software now data is live on LOCAL+CPCB+HSPCB | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL23170471429 | 2024-02-15 | 2024-02-17 | Cust 895 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Chlorine sensor | CHLORINE SENSOR(1) | To check FRC analyzer with FM5296H Sr. no. 2203H29808 and chlorine sensor FCL505 | Checked the chlorine analyzer at group tech industries Bahadurgarh and found that membrane cap was not available at site, also chlorine analyzer demo setup was not ready at site and customer was advised to arrange the same to check the analyzer reading. Customer wants to check the analyzer with on/off control in presence of Toshniwal engineer who has supplied the on/off valve. Customer will confirm once site is ready. Membrane cap is being sent to customer by analytical div. | Chlorine analyzer to be checked once site is ready | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23170481430 | 2024-02-15 | 2024-02-17 | Cust 786 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 3rd CAMC visit | 3rd CAMC visit completed | Working fine | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23170491431 | 2024-02-15 | 2024-02-17 | Cust 809 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 4th AMC visit for Multiparameter and Oil in Water analyzers | 4th AMC visit completed | Working fine | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23170571432 | 2024-02-15 | 2024-03-22 | Cust 807 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Data Breaking issue on CPCB server. | we have updated OS of the router so now the data is updating continuously | we have updated OS of the router so now the data is updating continuously | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23170711433 | 2024-02-15 | 2024-02-17 | Cust 801 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MUVX Card of Multiparameter(1) | 3rd AMC visit | 3rd AMC visit completed | Working fine | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23171221434 | 2024-02-16 | 2024-02-16 | Cust 801 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB + RSPCB data upload issue. | we took the access to your system and now the issue has been resolved and data is live on CPCB + RSPCB server. | we took the access to your system and now the issue has been resolved and data is live on CPCB + RSPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23172421435 | 2024-02-19 | 2024-02-26 | Cust 896 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | AQUAMAX DOT(1) | Issue :- Material is dispatched & need CSD engineer for commissioning support . End User :- 21.5 MLD STP, Secotr-2, Near UP Jal Nigam End User Site address :- 21.5 MLD STP, Secotr-2, Near UP Jal Nigam , Ghaziabad , 201009 Uttar Pradesh , India End User Name & Contact Details :- Mr. Anil Pandit- 9130983038 | CHECKED ANALYZER IN AIR AND FOUND WORKING FINE. COMMISSIONING IS PENDING DUE TO DIP TYPE SYSTEM AND FLOW THROUGH CHAMBER WAS NOT AVAILABLE AT SITE. | CHECKED ANALYZER IN AIR AND FOUND WORKING FINE. COMMISSIONING IS PENDING DUE TO DIP TYPE SYSTEM AND FLOW THROUGH CHAMBER WAS NOT AVAILABLE AT SITE. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23172591436 | 2024-02-19 | 2024-02-22 | Cust 897 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | RSPCB data upload issue. | Now the issue has been resolved and data is live on the server. | Now the issue has been resolved and data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23172931437 | 2024-02-20 | 2024-02-25 | Cust 795 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | Troubleshooting | Checked do calibration and found working fine. | Checked do calibration and found working fine. | NO | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23173061438 | 2024-02-20 | 2024-02-26 | Cust 532 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | To replace RS485 and calibrate the analyzer | Replaced RS485 card and calibrated the analyzer | Working fine | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23173421439 | 2024-02-20 | 2024-02-25 | Cust 898 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | pH Transmitter Aquamon | Aquamon 2000 plus(AQ2000 plus) | Commissioning | pH and conductivity analyser have been commissioned through remote support. Now system is working fine | pH and conductivity analyser have been commissioned through remote support. Now system is working fine. Hence this complaint has been closed. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFDL23174371440 | 2024-02-22 | 2024-02-28 | Cust 888 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | VISI pro DO(VISI pro DO) | commissioining of DO METER- 2 No. | MOM attached | MOM attached | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23174381441 | 2024-02-22 | 2024-02-28 | Cust 723 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | FM make RS485 card(1) | MODBUS output not getting | MOM attached | MOM attached | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23174681442 | 2024-02-22 | 2024-02-28 | Cust 786 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 1st CAMC visit | 1st CAMC visit completed | Working fine | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23174951443 | 2024-02-22 | 2024-02-29 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Schedule visit for ETP#3 Analyzers as COD, Phenol & Sulphide Analyzers against AMC. | Visit for COD, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | Visit for COD, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23175561444 | 2024-02-23 | 2024-03-03 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | visit for checking & calibration of 22 nos WQA ETP plant. | visit done for all 22 nos WQ analyzers as per attached mom.. | visit done for all 22 nos WQ analyzers as per attached mom. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23176211445 | 2024-02-25 | 2024-02-29 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | AMC | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | Scheduled AMC visit for COD, TOC, Oi l in Water, Sulphide Analyzers, installed at ETP # 1, 2 & TTPRO. | Visit for COD, TOC, Oil in Water, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | Visit for COD, TOC, Oil in Water, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23176471446 | 2024-02-26 | 2024-02-29 | Cust 804 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software re-installation | We took access to the system and re-installed the software. completed the HSPCB + CPCB + Local configuration and data is now live on the same servers | We took access to the system and re-installed the software. completed the HSPCB + CPCB + Local configuration and data is now live on the same servers | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23179181447 | 2024-02-28 | 2024-03-05 | Cust 892 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | DO showing 0 ppm in process and not performing calibration | MOM attached | MOM attached | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23180471448 | 2024-02-29 | 2024-03-01 | Cust 608 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Scheduled visit for ETP online Analyzer against AMC. | Visit for online Multi-Parameter Analyzer has been completed, observed Analyzer is working satisfactorily. Details are in attached MOM. | Visit for online Multi-Parameter Analyzer has been completed, observed Analyzer is working satisfactorily. Details are in attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23180861449 | 2024-03-01 | 2024-03-01 | Cust 804 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | DATCON | DATCON(DATCON) | CPCB and HSPCB problem issued. | I took access of the DATCon and found that the ETP data is not coming till the DATCon Hence as of now we are closing this complaint. | I took access of the DATCon and found that the ETP data is not coming till the DATCon Hence as of now we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23180931450 | 2024-03-01 | 2024-03-12 | Cust 899 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | COMMISSIONING | Commissioning done and mail attached | commissioning done &amp;amp; Mail attached. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFDL23181081451 | 2024-03-01 | 2024-03-02 | Cust 533 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Maintenance and Calibration | Maintenance and calibration work completed | Working fine | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL23181641452 | 2024-03-04 | 2024-03-07 | Cust 900 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera streaming. | Issue has been resolved and Camera streaming is live on the CPCB server. | Issue has been resolved and Camera streaming is live on the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23181951453 | 2024-03-05 | 2024-03-06 | Cust 901 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Electrode Cable(con-cab) | Commissioning at jwalamukhi, Dahera LWSS turbidity, conductivity & pH analyzer- Total 6 No. | MOM attached | MOM attached | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23182591454 | 2024-03-06 | 2024-03-15 | Cust 815 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | AQUADO Analyzer | AQUAMAX DOT(1) | The pH meter and DO meter are not working an dwe require Forbes service engineer for rectification. | Telephonically supported. Mail attached | Telephonically supported. Mail attached | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFDL23182821455 | 2024-03-07 | 2024-03-08 | Cust 902 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | CPCB + UPPCB configuration needs to be done in UBUNTU system. | As of now we are closing this complaint, and will reopen it after hardware side connection is ok. | As of now we are closing this complaint, and will reopen it after hardware side connection is ok. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23183431456 | 2024-03-09 | 2024-03-11 | Cust 849 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Support for analyzers at CETP Bhiwadi | UF feed MPA recommissioned after replacing the PCX1, dual channel DO commissioned, replaced pH sensor, calibrated neuralization tank pH meter, flow cell of MPS MPA cleaned & calibrated with standard solution | Working fine | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23184951457 | 2024-03-12 | 2024-03-21 | Cust 849 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Site visit for MPA. | Successfully commission of multiple Analyzers, observed all are working satisfactorily. Details are in attached MOM. | Successfully commission of multiple Analyzers, observed all are working satisfactorily. Details are in attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23185111458 | 2024-03-13 | 2024-03-16 | Cust 843 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Replacement of tubing and calibration under CAMC | Tubing replacement and calibration work completed | Working fine | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23185331459 | 2024-03-13 | 2024-03-15 | Cust 863 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Black Box configuration and CPCB + HSPCB connectivity needs to be done. | the Black Box configuration and CPCB + HSPCB connectivity for STP 5 MLD HUDA site has been completed and now data is live on both sites. | the Black Box configuration and CPCB + HSPCB connectivity for STP 5 MLD HUDA site has been completed and now data is live on both sites. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23185631460 | 2024-03-14 | 2024-03-14 | Cust 903 | DELHI | Steam System Division | CLOSED | PRODUCT FAILURE | PRV - MEASURETROL | FM MEASURETROL(SIZE 050NB) | Cable of Measure trol faulty | New cable changed and working | New cable changed and working | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFDL23185811461 | 2024-03-14 | 2024-03-20 | Cust 837 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Checking & Calibration of MPA | Done the calibration | Done the calibration of Multiparameter. | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23185821462 | 2024-03-14 | 2024-03-20 | Cust 863 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | Falcon Scanner for Multiparameter(AWA-SCAN) | Commissioning of Black box and reconfiguration. | Done the commissioning of FM black box and connectivity. | Done the commissioning of FM black box and connectivity. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23185831463 | 2024-03-14 | 2024-03-16 | Cust 849 | DELHI | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | MULTIPARAMETER ANALYSER | pH input card-FM of Multiparameter(1) | Meeting with customer | We visited Triveni Noida office on 12th March to discuss and take necessary action to resolve the issues at CETP Bhiwadi. Suneet moved to CETP Bhiwadi site from the next day. | We visited Triveni Noida office on 12th March to discuss and take necessary action to resolve the issues at CETP Bhiwadi. Suneet moved to CETP Bhiwadi site from the next day. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFDL23185921464 | 2024-03-14 | 2024-03-16 | Cust 788 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 3rd AMC visit | 3rd AMC visit completed | Working fine | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23186011465 | 2024-03-14 | 2024-03-25 | Cust 874 | DELHI | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | MULTIPARAMETER ANALYSER | RS485 card-FM of Multiaparameter (1) | Customer reported for communication failure & by telephonic support found RS 485 card faulty. | Replacement material has been booked through TCI EXPRESS vide LR No - 110007585063 LR Dt - 21-MAR-2024 | Replacement material has been booked through TCI EXPRESS vide LR No - 110007585063 LR Dt - 21-MAR-2024 | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFDL23186391466 | 2024-03-15 | 2024-03-16 | Cust 788 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 4th AMC visit | 4th AMC visit completed | Working fine | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23187851467 | 2024-03-18 | 2024-03-18 | Cust 904 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Engineer visit require to calibrate MPA | Engineer visited and calibrate the analyzer. Report attached | Resolved | EXTENDED | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFDL23188131468 | 2024-03-18 | 2024-03-21 | Cust 806 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to the CPCB and HSPCB server. | Data is uploading to the HSPCB, CPCB and Local server. | Data is uploading to the HSPCB, CPCB and Local server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL23188601469 | 2024-03-19 | 2024-03-21 | Cust 533 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Issue in Enviro Connect Data Portal | Data is live on the server | Data is live on the server | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23188651470 | 2024-03-19 | 2024-03-20 | Cust 802 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | HSPCB data upload issue. | JAVA file is installed in your local PC After uninstalling the same now local host server is accessible and data is live on the local + CPCB + HSPCB server. | JAVA file is installed in your local PC After uninstalling the same now local host server is accessible and data is live on the local + CPCB + HSPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23188781471 | 2024-03-19 | 2024-03-31 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | visit for checking & calibration of 22 nos water quality analyzers under ETP AMC | visit done as per attached mom. | visit done .MOM attached . | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23190491472 | 2024-03-21 | 2024-03-23 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Schedule visit for ETP#3 Analyzers as COD, Phenol & Sulphide Analyzers against AMC. | Visit for COD, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | Visit for COD, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23191201473 | 2024-03-22 | 2024-03-25 | Cust 905 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software re-installation | As of now we are closing this complaint, we will reopen this once we get access to the system. | As of now we are closing this complaint, we will reopen this once we get access to the system. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL23192781474 | 2024-03-25 | 2024-03-30 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | AMC | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | Scheduled AMC visit for COD, TOC, Oil in Water , Sulphide Analyzers, installed at ETP # 1, 2 & TTPRO. | Visit for COD, TOC, Oil in Water, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | Visit for COD, TOC, Oil in Water, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23192971475 | 2024-03-26 | 2024-03-26 | Cust 874 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | PH Input Card for Multiparameter(AWA-PH) | Customer reported for EQMS PH card not responding | EQMS reset done now working ok. | EQMS reset done now working ok. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23193051476 | 2024-03-26 | 2024-04-03 | Cust 802 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Checking and Calibration Of Multiparameter | Done the Calibration | Done the calibration of Multiparameter. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23193801477 | 2024-03-27 | 2024-03-29 | Cust 905 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | HSPCB data uploading issue. | we took the access to your system, and now the issue has been resolved and data is live on the RSPCB server. | we took the access to your system, and now the issue has been resolved and data is live on the RSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL23194021478 | 2024-03-27 | 2024-04-03 | Cust 837 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | MULTIPARAMETER ANALYSER | Falcon Scanner for Multiparameter(AWA-SCAN) | Online data transfer issue | During checking found that Internet cable issue replaced the cable by customer . | During checking found that Internet cable issue. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL23194791479 | 2024-03-28 | 2024-03-29 | Cust 839 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | RSPCB data upload issue. | we took the access to your system, and now the issue has been resolved and data is live on the RSPCB server. | we took the access to your system, and now the issue has been resolved and data is live on the RSPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL23195781480 | 2024-03-29 | 2024-03-29 | Cust 841 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | PH Sensor for Multiparameter(AWA-PH-SENS) | Engineer visit required to check PH reading issue. | Engineer visited site and found PH sensor faulty and replaced from stock. | Resolved | YES | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFDL23195931481 | 2024-03-29 | 2024-03-30 | Cust 906 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Display of Smartpro(1) | commission of auto-cleaning mounting pH meter OA 5524107771 | Now material is in transit. Hence I closed this complaint. | Now material is in transit. Hence I closed this complaint. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFDL23197721482 | 2024-03-31 | 2024-04-01 | Cust 796 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Aqua4Trans | Aqua4Trans(1) | pH/TSS analyzer checking and calibration | pH/TSS checked & calibrated successfully | Working fine | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL23C03621483 | 2023-06-23 | 2023-07-13 | Cust 907 | DELHI | Swas - CHAKAN | CLOSED | SHORT SUPPLY | SWAS & Analysers | Steam & Water Analysis System(SWAS & Analysers) | conductivity meter cable short supply. | Material is sent as per requirement , hence there is no short supply from FM. | Material supplied as per requirement | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFDL24000121484 | 2024-04-01 | 2024-04-04 | Cust 908 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | pH input card-FM of Multiparameter(1) | installation guideline visit at safdarjung hospital STP site delhi | visit done | visit done .mail communication attached. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24000541485 | 2024-04-01 | 2024-04-04 | Cust 810 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | visit for checking & calibration of MPA | visit done & mom attached. | visit done & mom attached. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL24000561486 | 2024-04-01 | 2024-04-04 | Cust 848 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Schedule visit for ETP online Analyzer against AMC at Mankind Research Center, Manesar. | Visit for online Multi-Parameter Analyzer has been completed, observed Analyzer is working satisfactorily. Details are in attached MOM. | Visit for online Multi-Parameter Analyzer has been completed, observed Analyzer is working satisfactorily. Details are in attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24000571487 | 2024-04-01 | 2024-04-07 | Cust 890 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning of Multiparameter analyzer at bhulana 2 MLD Kapoorthala | MOM attached, connectivity pending. Doing after receiving of rel. docs. | MOM attached, Connectivity pending | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24000911488 | 2024-04-02 | 2024-04-09 | Cust 822 | DELHI | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | MULTIPARAMETER ANALYSER | FM make pH card(1) | Customer reported for problem with PH sensor | RF no138243274 done card will be dispatched soon | PH card RF138243274 done it will be dispatched soon | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24000921489 | 2024-04-02 | 2024-04-26 | Cust 909 | DELHI | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | Ph Tranmitter Smartpro | Display of Smartpro(1) | Installation guide lines to be provided to customer for Chlorine and Turbidity Analyzer. | The installation guidelines for the Chlorine and Turbidity Analyzer have been provided to the customer representative. | The installation guidelines for the Chlorine and Turbidity Analyzer have been provided to the customer representative. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24001431490 | 2024-04-03 | 2024-04-05 | Cust 807 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data uploaded issue. | RS485 issue has been resolved and now the data is live on the server. | RS485 issue has been resolved and now the data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24001531491 | 2024-04-03 | 2024-04-08 | Cust 858 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | 12 V POWER SUPPLY CARD for Multiparameter(AWA-PS1) | AMC visit | AMC visit done. | Done the calibration of Multiparameter. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24001681492 | 2024-04-04 | 2024-04-13 | Cust 730 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration of MPA | mom attached | MOM attached | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL24001741493 | 2024-04-04 | 2024-04-05 | Cust 724 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | New device addition on CPCB. | We have completed CPCB connectivity and data is live now on CPCB Server. Camera IP is not playing in customers network, we will reopen this complaint once the issue is resolve from customer end. | We have completed CPCB connectivity and data is live now on CPCB Server. Camera IP is not playing in customers network, we will reopen this complaint once the issue is resolve from customer end. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24001991494 | 2024-04-04 | 2024-04-29 | Cust 807 | DELHI | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | MULTIPARAMETER ANALYSER | CPU with LCD of Multiparameter Display(1) | PCX1 with display card having issue of brightness hence need to replace it under warranty (24 months warranty) | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210097142 LR Dt - 20-APR-2024 | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210097142 LR Dt - 20-APR-2024 | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFDL24002181495 | 2024-04-04 | 2024-04-07 | Cust 910 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | Visit to commissioning of Oil in water Analyzer. | Commissioned Oil-In-Water Analyzer, Observed working satisfactorily. Details are in attached MOM. | Commissioned Oil-In-Water Analyzer, Observed working satisfactorily. Details are in attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24002191496 | 2024-04-04 | 2024-06-08 | Cust 849 | DELHI | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | Aqua4Trans | Display card of Aqua4Trans(1) | Observed that the Analog output was not being received from Analog Output (Ch1 & 2) terminals. Therefore, the output was configured on Ch 3 & 4 terminals. Please send one Display cum CPU for this transmitter. | Replaced all the Display card of Aqua4Trans at site. | Replaced all the Display card of Aqua4Trans at site. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFDL24002191496 | 2024-04-04 | 2024-06-08 | Cust 849 | DELHI | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | Aqua4Trans | Display card of Aqua4Trans(1) | Observed that the display contrast and brightness were high, resulting in text that was not readable (Tx Sr No. 23A4TO016-W040000). Please send the display cum CPU for this transmitter. | Replaced all the Display card of Aqua4Trans at site. | Replaced all the Display card of Aqua4Trans at site. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFDL24002551497 | 2024-04-05 | 2024-04-09 | Cust 906 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Smartpro Electronics Assembly(Smartpro-Elc-cbl) | Commissioning of pH measurement system with auto cleaning | Commissioned & found working fine. | Commissioned & found working fine. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24002631498 | 2024-04-05 | 2024-04-26 | Cust 786 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | RSPCB data uploaded issue., | for this the NTP server port 123 is blocked in the customer network hence BB is not able to sync with internet time hence currently the data is live on the server. | for this the NTP server port 123 is blocked in the customer network hence BB is not able to sync with internet time hence currently the data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24003231499 | 2024-04-08 | 2024-04-10 | Cust 806 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not getting uploaded to CPCB and HSPCB | Data is live on the CPCB and HSPCB servers. | Data is live on the CPCB and HSPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24003261500 | 2024-04-08 | 2024-04-12 | Cust 824 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Customer reported for calibration of EQMS at there site | Calibration done | Calibration done. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL24003391501 | 2024-04-08 | 2024-04-08 | Cust 800 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Engineer visit required for calibration of MPA as per AMC schedule. | Engineer visited site and done the calibration. MOM attached. | Resolved | EXTENDED | Periodic Maintenance (AMC) | Process | N |  | Analytical - CHAKAN
AFDL24003711502 | 2024-04-08 | 2024-04-26 | Cust 799 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Checking and calibration of MPA | Done the calibration | Done the calibration of Multiparameter. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24004111503 | 2024-04-09 | 2024-04-10 | Cust 886 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue. | The hardware issue has been resolved as the data is now live on the server. | The hardware issue has been resolved as the data is now live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24004231504 | 2024-04-09 | 2024-04-29 | Cust 876 | DELHI | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Customer reported for PH readings fluctuating and hence checked with old sensor PH found new sensor faulty. | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210097141 LR Dt - 20-APR-2024 | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210097141 LR Dt - 20-APR-2024 | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFDL24005001505 | 2024-04-10 | 2024-04-20 | Cust 849 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | As per customer , the online Multiparameter Analyzer reading is not matching with the Lab readings. | Attended multiple Analyzers and resolved issues & concludes solutions for site, Details are in attached MOM. | Attended multiple Analyzers and resolved issues & concludes solutions for site, Details are in attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24005211506 | 2024-04-11 | 2024-04-14 | Cust 709 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration | Calibrated & found MPA is working fine. | Calibrated & found MPA is working fine. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL24005411507 | 2024-04-12 | 2024-04-18 | Cust 821 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Customer reported for calibration of EQMS system. | Calibration done | Calibration done. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL24005451508 | 2024-04-12 | 2024-04-12 | Cust 518 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | RSPCB data upload issue. | we have checked the RSPCB data availability and found that it is above 95% from the last 2 days. | we have checked the RSPCB data availability and found that it is above 95% from the last 2 days. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24007561509 | 2024-04-18 | 2024-04-22 | Cust 911 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Transmitter Aquacon | Aquacon 2000(AQCN 2000 p) | Calibration and health mapping TDS sensors. | Calibration done | Calibration done of TDS meter and pH electrode not working. | NO | Paid | Process | N |  | Analytical - CHAKAN
AFDL24007561509 | 2024-04-18 | 2024-04-22 | Cust 911 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Calibration and health mapping of pH sensors. | pH Electrode not working | Calibration done of TDS meter and pH electrode not working. | NO | Paid | Process | N |  | Analytical - CHAKAN
AFDL24007561509 | 2024-04-18 | 2024-04-22 | Cust 911 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Calibration and health mapping of pH sensors. | pH electrode not working. | Calibration done of TDS meter and pH electrode not working. | NO | Paid | Process | N |  | Analytical - CHAKAN
AFDL24007981510 | 2024-04-19 | 2024-04-26 | Cust 912 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | commissioning at 14MLD stp plant at faridkot | MOM attached | MOM attached | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24008551511 | 2024-04-20 | 2024-04-25 | Cust 820 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | visit for replacement of newly supplied PCX1 with display board of MPA. | visit done for PCX1 Display card replacement at site . | visit done for PCX1 Display card replacement at site .MOM attached. | NO | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24008951512 | 2024-04-21 | 2024-04-26 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Schedule visit for ETP#3 Analyzers as COD, Phenol & Sulphide Analyzers against AMC. | Visit for COD, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | Visit for COD, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24009791513 | 2024-04-22 | 2024-04-24 | Cust 807 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | calibration | Checked Calibration & found working fine. | Checked Calibration & found working fine. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL24009801514 | 2024-04-22 | 2024-05-04 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | Iocl mathura ETP Analyzers AMC | visit done under April ARC for WQ ETP analyzers .MOM attached . | visit done under April ARC for WQ ETP analyzers .MOM attached . | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24010711515 | 2024-04-24 | 2024-05-31 | Cust 564 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(Ph-Trans-Smartpro) | received mail from Amol Bhamare to address the pH transmitter issue. Site address ,Devyani Food Industries Limited Khasra No 31 to 34,40 to 44, 366 and 367 Village Nanan, SIDCO Industrial Complex, Phase-II, Ghatti, Kathua, Pin:-184143 GST No: 01AABCD7572N2ZH Devyani PO No: 5400111979 Thermax Contact Person at Site: Mr. Sauraj Gil & his Mobile No: +91-9719991127 & +91-9634078214 Rupneet Singh M:7009699218 | Called to end user and concluded issue in pH smartpro buffer amplifier circuit. Mail to end user to send the transmitter ro pune for testing and repairing. Mail attached | Called to end user and concluded issue in pH smartpro buffer amplifier circuit. Mail to end user to send the transmitter ro pune for testing and repairing. Mail attached | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFDL24010791516 | 2024-04-24 | 2024-04-30 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | AMC | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | Scheduled AMC visit for COD, TOC, Oi l in Water, Sulphide Analyzers, installed at ETP # 1, 2 & TTPRO. | Visit for COD, TOC, Oil in Water, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | Visit for COD, TOC, Oil in Water, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24010811517 | 2024-04-24 | 2024-04-27 | Cust 814 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | visit for checking if PCX2 Board as RS485 card was burnt inside Panel. | visit done & Found PCX1 card Input slot 4 is working fine.mail communication attached . | visit done & Found PCX1 card Input slot 4 is working fine.mail communication attached . | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24011001518 | 2024-04-24 | 2024-06-10 | Cust 913 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | commissioning | Commissioning of MPA done | MOM attached | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24011041519 | 2024-04-24 | 2024-04-26 | Cust 790 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 3rd AMC visit | 3rd AMC visit completed | Working fine | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24011731520 | 2024-04-25 | 2024-04-27 | Cust 801 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 4th AMC visit | 4th AMC visit completed | Working fine | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24013531521 | 2024-04-29 | 2024-05-02 | Cust 532 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data availability is low | We took access to the system and data is live on the Local +CPCB servers. For RSPCB, we are checking this issue with the RSPCB team. As of now we are closing this complaint. | We took access to the system and data is live on the Local +CPCB servers. For RSPCB, we are checking this issue with the RSPCB team. As of now we are closing this complaint. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24014261522 | 2024-04-30 | 2024-06-14 | Cust 839 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | RSPCB data upload issue. | We took access to the BB and updated some files in the configuration of the BB and Now the Data Availability on the RSPCB server is 99 % for Stack 7. | We took access to the BB and updated some files in the configuration of the BB and Now the Data Availability on the RSPCB server is 99 % for Stack 7. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24016051523 | 2024-05-01 | 2024-05-05 | Cust 910 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | Visit for training and demonstration of operation, maintenance, and calibration of OIL IN WATER ANALYZER for the IOCL representatives. | Demonstration / Training has been provided to IOCL representative & commissioned successfully Oil-In-Water, observed the analyzer is working fine. Details are in attached MOM. | Demonstration / Training has been provided to IOCL representative & commissioned successfully Oil-In-Water, observed the analyzer is working fine. Details are in attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24016061524 | 2024-05-01 | 2024-05-05 | Cust 786 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Schedule visit for ETP online Analyzer against CAMC. | Visit for online Multi-Parameter Analyzer has been completed against CAMC, observed Analyzer is working satisfactorily. Details are in attached MOM. | Visit for online Multi-Parameter Analyzer has been completed against CAMC, observed Analyzer is working satisfactorily. Details are in attached MOM. | EXTENDED | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24017921525 | 2024-05-06 | 2024-05-09 | Cust 849 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Aqua4Trans | Display card of Aqua4Trans(1) | Display to be replaced of Aqua4Trans Transmitters. | Attended Multi-Parameter Analyzer, Turbidity, DO & Chlorine Analyzers and replaced display of Aqua4Trans Transmitters, observed all are working satisfactorily. Details are in attached MOM. | Attended Multi-Parameter Analyzer, Turbidity, DO & Chlorine Analyzers and replaced display of Aqua4Trans Transmitters, observed all are working satisfactorily. Details are in attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24018171526 | 2024-05-07 | 2024-05-09 | Cust 546 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 4th AMC visit | AMC visit completed | Working fine | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24018351527 | 2024-05-07 | 2024-05-13 | Cust 914 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Smartpro-con) | Telephonic support for commissioning | Commissioned through telephonic support | Provided telephonic support for commissioning | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFDL24018611528 | 2024-05-08 | 2024-05-10 | Cust 915 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB + UPPCB data upload issue. | We took access of the system and made some changes in the BB configuration. Currently data is live on CPCB + UPPCB server. | We took access of the system and made some changes in the BB configuration. Currently data is live on CPCB + UPPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24018721529 | 2024-05-08 | 2024-05-10 | Cust 724 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | camera streaming failed in online connectivity . | now the camera streaming is live on the server. | now the camera streaming is live on the server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24019271530 | 2024-05-10 | 2024-05-10 | Cust 803 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | need Software reinstallation | We have re-installed the software with updated version and completed Local + CPCB + RSPCB connectivity and the data is now live on the same servers. | We have re-installed the software with updated version and completed Local + CPCB + RSPCB connectivity and the data is now live on the same servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24019291531 | 2024-05-10 | 2024-05-10 | Cust 793 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | As per AMC calibration is due, need engineer visit. | Engineer visited the site and done the calibration. | Resolved | EXTENDED | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFDL24019641532 | 2024-05-11 | 2024-05-12 | Cust 916 | DELHI | Analytical - CHAKAN | CLOSED | SALES PROMOTION | Aqua4Trans | Conductivity Input card of Aqua4trans(1) | Proactive visit. | Aqua4trans offer has been send. | Aqua4trans offer has been send | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFDL24020321533 | 2024-05-13 | 2024-05-20 | Cust 802 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | CPU Display card for Multiparameter(AWA-CPU1) | Display Blank | Replaced the MPA display and reconfigured. | Done the calibration of Multiparameter. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24020501534 | 2024-05-14 | 2024-05-18 | Cust 823 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | To check the analyzer for error 2 problem | Troubleshooting done | Working fine | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24021381535 | 2024-05-15 | 2024-05-17 | Cust 803 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not updating to CPCB and RSPCB server. | Data is now live on the CPCB and RSPCB servers | Data is now live on the CPCB and RSPCB servers | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24021461536 | 2024-05-15 | 2024-05-27 | Cust 716 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | PH1 & PH2 detector installation as detector down found in previous remote support | MOM attached | MOM attached | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24021881537 | 2024-05-16 | 2024-05-18 | Cust 917 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Aqua4Trans | Display card of Aqua4Trans(1) | Display blank | Display replaced | Working fine | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24021911538 | 2024-05-16 | 2024-05-18 | Cust 532 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is low. | As per discussion with customer data availability was 89.58. But after that Internet cable cut at his end. Once net availability is there they will confirm current status. | As per discussion with customer data availability was 89.58. But after that Internet cable cut at his end. Once net availability is there they will confirm current status. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24022101539 | 2024-05-17 | 2024-05-23 | Cust 918 | DELHI | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | Aqua4Trans | Aqua4 trans-A4TW033000(1) | customer reported for installation of PH/turbidity/chlorine analyzers | installation guidelines provided | Installation guidelines given | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24022241540 | 2024-05-17 | 2024-05-18 | Cust 786 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 2nd CAMC visit | 2nd CAMC visit completed | Working fine | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24022251541 | 2024-05-17 | 2024-05-20 | Cust 904 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | ETP final discharge parameter showing in OCEMS not matching with Forbes Marshall portal on PC . | Now the CPCB data is matching with the actual readings. | Now the CPCB data is matching with the actual readings. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24022301542 | 2024-05-17 | 2024-05-17 | Cust 807 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB Central portal data uploaded issue. | Now the data is live on the Central Portal. | Now the data is live on the Central Portal. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24022571543 | 2024-05-19 | 2024-05-27 | Cust 599 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | SBR-1 DO analyzer not showing reading correct | MOM attached | MOM attached | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24022991544 | 2024-05-20 | 2024-05-24 | Cust 919 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Silica Analyser | Silica 9210(9210) | visit for commissioning of Silica Analyzers Make- SWAN | visit done as per attached mail communication. | visit done as per attached mail communication. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24023511545 | 2024-05-20 | 2024-05-23 | Cust 802 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | HSPCB data uploaded issue. | now the data is live on the HSPCB server. | now the data is live on the HSPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24023861546 | 2024-05-20 | 2024-05-26 | Cust 812 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | pH input card-FM of Multiparameter(1) | AMC visit | Done the AMC visit. | Done the calibration of Multiparameter. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24023861546 | 2024-05-20 | 2024-05-26 | Cust 812 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | pH input card-FM of Multiparameter(1) | AMC visit | Done the AMC visit. | Done the calibration of Multiparameter. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24023911547 | 2024-05-21 | 2024-05-23 | Cust 786 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 4th CAMC visit | 4th CAMC visit completed | Working fine | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24023961548 | 2024-05-21 | 2024-05-22 | Cust 692 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | There is Zero PM data availability in RSPCB server since 15th May 2024 onwards. Although the data is available in the Forbes marshal server. Kindly Upload the data in RSPCB server | Now data is live on CPCB + Local + RSPCB server. | Now data is live on CPCB + Local + RSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24024151549 | 2024-05-21 | 2024-05-22 | Cust 796 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | ETP OUTLET data upload issue. | We took access to the system and now ETP OUTLET data is live on CPCB + Local + RSPCB server. | We took access to the system and now ETP OUTLET data is live on CPCB + Local + RSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24024481550 | 2024-05-22 | 2024-05-26 | Cust 837 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Schedule visit for online Multi Parameter Analyzer in STP against AMC | Attended the Multi-Parameter Analyzer against AMC, Servicing and calibration done. Details are in the attached MOM. | Attended the Multi-Parameter Analyzer against AMC, Servicing and calibration done. Details are in the attached MOM. | EXTENDED | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24024731551 | 2024-05-22 | 2024-05-24 | Cust 803 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | RSPCB data Availability issue. | we took access and made some changes so now the Data availability is almost 90% on the RSPCB server. | we took access and made some changes so now the Data availability is almost 90% on the RSPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24024761552 | 2024-05-22 | 2024-05-29 | Cust 849 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Troubleshooting in Multi-Parameter Analyzer in inlet water line. | Attended the Multiple Analyzers and observed working satisfactorily. Details are in the attached MOM. | Attended the Multiple Analyzers and observed working satisfactorily. Details are in the attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24026581553 | 2024-05-26 | 2024-05-30 | Cust 909 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4 trans-A4TW033000(1) | Commissioning of Chlorine Analysers | Done the commissioning of Chlorine Analyser | Done the commissioning of Chlorine Analyzers. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24027231554 | 2024-05-27 | 2024-07-02 | Cust 849 | DELHI | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | MULTIPARAMETER ANALYSER | Peristaltic sampling pump motor for Multiparameter(AWA-PSPM) | Peristatic sampling pump motor of inlet water multi-Parameter Analyzer was not working due to jam. Please arrange one motor for the same at our Delhi office. | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210075679 LR Dt - 11-JUN-2024 | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210075679 LR Dt - 11-JUN-2024 | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFDL24027571555 | 2024-05-27 | 2024-05-28 | Cust 804 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | 3 MLD CETP manakpur analyser is offline mode .from 17/05/2024. | Now data is live on the CPCB + Local server. | Now data is live on the CPCB + Local server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24027981556 | 2024-05-28 | 2024-05-31 | Cust 804 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to CPCB server | Data is now live on the CPCB server. | Data is now live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24028261557 | 2024-05-28 | 2024-05-30 | Cust 920 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Visit for servicing and calibration of Silica Analyzer. | Attended the Silica Analyzer, noticed mixing pump was not perform, it same to be replaced immediately for best result. Details are in attached MOM. | Attended the Silica Analyzer, noticed mixing pump was not perform, it same to be replaced immediately for best result. Details are in attached MOM. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL24028721558 | 2024-05-29 | 2024-06-01 | Cust 840 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue. | There is no system access from customer after 2 reminders. | There is no system access from customer after 2 reminders. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24028951559 | 2024-05-29 | 2024-05-31 | Cust 902 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | CPCB data upload issue. | As of now we are closing this complaint and will reopen it once the hardware side issue is resolved. | As of now we are closing this complaint and will reopen it once the hardware side issue is resolved. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24029241560 | 2024-05-29 | 2024-06-01 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Schedule visit for ETP#3 Analyzers as COD, Phenol & Sulphide Analyzers against AMC. | Visit for COD, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | Visit for COD, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24029271561 | 2024-05-30 | 2024-06-06 | Cust 868 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration | Calibrated & found working fine. | Calibrated & found working fine. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL24029301562 | 2024-05-30 | 2024-05-31 | Cust 812 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data uploded issue. | Now the issue has been resolved and data is showing live on the CPCB server. | Now the issue has been resolved and data is showing live on the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24029561563 | 2024-05-30 | 2024-06-01 | Cust 804 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software re-configuration and connectivity with cpcb/hspcb . | Data is live now for STP Outlet 2 on the CPCB server. | Data is live now for STP Outlet 2 on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24029781564 | 2024-05-30 | 2024-06-25 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | AMC | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | Visit under AMC | visit done under AMC .MOM attached. | visit done under AMC .MOM attached. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24030781565 | 2024-05-31 | 2024-06-01 | Cust 806 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB + LOCAL + HSPCB DATA UPLOAD ISSUE. | Now data is live on CPCB + Local + HSPCB server. | Now data is live on CPCB + Local + HSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24030791566 | 2024-05-31 | 2024-05-31 | Cust 816 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera streaming issue. | We have created the RTSP link for camera connectivity. | We have created the RTSP link for camera connectivity. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24030801567 | 2024-05-31 | 2024-05-31 | Cust 876 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | UPPCB data uploded issue. | now the data is live on the UPPCB server. | now the data is live on the UPPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24032091568 | 2024-06-03 | 2024-06-05 | Cust 804 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on CPCB + Local + HSPCB. | We took access to the system and now data is live on CPCB + Local + HSPCB server. | We took access to the system and now data is live on CPCB + Local + HSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24032291569 | 2024-06-03 | 2024-06-08 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | AMC | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | Scheduled AMC visit for COD, TOC, Oi l in Water, Sulphide Analyzers, installed at ETP # 1, 2 & TTPRO. | Visit for COD, TOC, Oil in Water, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | Visit for COD, TOC, Oil in Water, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24035191570 | 2024-06-10 | 2024-06-10 | Cust 921 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | Aquamax system(1) | FM Engineer visit require for installation of controller of DO Analyser | FM Engineer visited site and installation of Controller has been done. | FM Engineer visited site and installation of Controller has been done. | NO | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFDL24035881571 | 2024-06-12 | 2024-06-16 | Cust 814 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | FM make RS485 card(1) | Visit for attending MPA and Datcon for cpcb connectivity | visit done for MPA & CPCB connectivity & now both are working fine. | visit done for MPA & CPCB connectivity & now both are working fine. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL24036141572 | 2024-06-12 | 2024-06-14 | Cust 787 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | DATCON | DATCON(DATCON) | DATCon configuration and CPCB and UPPCB connectivity needs to be done. | we have completed the DATCon configuration and now the data is live on the CPCB and UPPCB server. | we have completed the DATCon configuration and now the data is live on the CPCB and UPPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24036411573 | 2024-06-13 | 2024-06-16 | Cust 922 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Visit for MPA | visit done for MPA checking & calibration . | visit done for MPA checking & calibration . | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL24037191574 | 2024-06-14 | 2024-06-14 | Cust 692 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on Local + RSPCB server. | Now data is live on CPCB + Local + RSPCB server. | Now data is live on CPCB + Local + RSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24038871575 | 2024-06-17 | 2024-06-22 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | TOC ANALYSER | FM 101 TOC Analyser(FM 101 TOC) | Visit for TOC Analysers inspection, installed in BS-Vi unit | Inspection visit for TOC Analyzers has been completed, Details are in attached file. | Inspection visit for TOC Analyzers has been completed, Details are in attached file. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24039061576 | 2024-06-18 | 2024-07-16 | Cust 923 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Re-installation of "Enviroconnect Software." | We have accessed the system and successfully reinstalled the "Enviroconnect software." The data is now live on both the CPCB and local servers. However, we need to change the Device ID for Stack 1_35 TPH Boiler station on the HSPCB server. So the customer has raised this issue with HSPCB team, therefore we are closing this complaint. | However, we need to change the Device ID for Stack 1_35 TPH Boiler station on the HSPCB server. So the customer has raised this issue with HSPCB team, therefore we are closing this complaint. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24039251577 | 2024-06-18 | 2024-06-18 | Cust 904 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Engineer visit required for MPA AMC. | Engineer visited site and done the calibration. | Resolved | EXTENDED | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFDL24039371578 | 2024-06-18 | 2024-06-20 | Cust 839 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | RSPCB data upload issue. | Now data is live on the RSPCB server. | Now data is live on the RSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24039761579 | 2024-06-19 | 2024-06-22 | Cust 865 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software re-installation | We have re-installed the software with an updated version and completed local, CPCB and HSPCB connectivity. Now data is live on the same servers. | We have re-installed the software with an updated version and completed local, CPCB and HSPCB connectivity. Now data is live on the same servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24040251580 | 2024-06-20 | 2024-06-25 | Cust 802 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | FM make pH card(1) | Checking and calibration of MPA | Done the calibration of Multiparameter | Done the calibration of Multiparameter | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24040261581 | 2024-06-20 | 2024-06-25 | Cust 799 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | LCD of Multiparameter Display(1) | Display not working. | Checked the MPA display and found optical key alignment issue. | Checked the MPA display unit | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24040561582 | 2024-06-21 | 2024-06-21 | Cust 787 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | DATCON | DATCON(DATCON) | CPCB data uploaded issue. | Data is now live on the CPCB server. | Data is now live on the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24040941583 | 2024-06-22 | 2024-06-24 | Cust 843 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 1st AMC visit | 1st AMC visit completed | Working fine | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24041491584 | 2024-06-23 | 2024-06-26 | Cust 849 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | Peristaltic sampling pump motor for Multiparameter(AWA-PSPM) | Sample pump motor to be replaced in Multi-Parameter Analyzer. | Attended the Multi-Parameter Analyzer, replaced the sampling motor and completed Servicing and calibration. Details are in the attached MOM. | Attended the Multi-Parameter Analyzer, Servicing and calibration done. Details are in the attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24041501585 | 2024-06-23 | 2024-07-01 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | AMC | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | Scheduled AMC visit for COD, TOC, Oi l in Water, Sulphide Analyzers, installed at ETP # 1, 2 & TTPRO. | Visit for COD, TOC, Oil in Water, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | Visit for COD, TOC, Oil in Water, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24041681586 | 2024-06-24 | 2024-06-24 | Cust 692 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on CPCB + RSPCB server. | Now data is live on CPCB + Local + RSPCB server. | Now data is live on CPCB + Local + RSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24042131587 | 2024-06-25 | 2024-07-23 | Cust 924 | DELHI | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | AQUADO Analyzer | DPI-Digital process indicator(1) | after calibration in DPI, reading showing in difference of 0.4 ppm. Calibration issues and reading varied continuously. After commissioning of 5-6 days, display getting off. | Material received in depot for repairs. Depot repair complaint number CCS24D0624 | Material received in depot for repairs. Depot repair complaint number CCS24D0624 | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFDL24043061588 | 2024-06-26 | 2024-06-29 | Cust 816 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | limit updation. | As per the CPCB's mail, we have now updated the Stack Limit on our server. | As per the CPCB's mail, we have now updated the Stack Limit on our server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24043201589 | 2024-06-26 | 2024-07-06 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | visit under AMC | visit done as per attached mom. | visit done as per attached mom. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL24043261590 | 2024-06-27 | 2024-07-05 | Cust 719 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration | MOM attached | MOM attached | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL24045361591 | 2024-06-30 | 2024-07-08 | Cust 709 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC-3rd | work done | work done | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24045371592 | 2024-06-30 | 2024-07-08 | Cust 925 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Conductivity Transmitter Aquacon | Aquacon 2000(AQCN 2000 p) | Conductivity not showing correct | work done | work done | NO | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24045381593 | 2024-06-30 | 2024-07-08 | Cust 926 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Conductivity Transmitter Aquacon | Aquacon 2000(AQCN 2000 p) | Conductivity reading showing 0 | work done | work done | NO | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24045451594 | 2024-06-30 | 2024-07-08 | Cust 927 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 1st AMC | AMC done | AMC done | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24047041595 | 2024-07-01 | 2024-07-03 | Cust 867 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | As per customer complaint that the Silica Analyzer was not performing. | Attended the Silica Analyzer, observed their LED and Detector were not performing. These to be replaced immediately for properly functioning. Details are in attached MOM. | Attended the Silica Analyzer, observed their LED and Detector were not performing. These to be replaced immediately for properly functioning. Details are in attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24047411596 | 2024-07-01 | 2024-07-04 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Schedule visit for ETP#3 Analyzers as COD, Phenol & Sulphide Analyzers against AMC. | Visit for COD, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | Visit for COD, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24047951597 | 2024-07-03 | 2024-07-05 | Cust 800 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data mismatch with actual readings. | we have made some modification from configuration end and now COD data is matching with the actual readings and the same is verified by you. | we have made some modification from configuration end and now COD data is matching with the actual readings and the same is verified by you. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24048041598 | 2024-07-03 | 2024-07-03 | Cust 928 | DELHI | Steam System Division | CLOSED | PRODUCT UPGRADE | CRS | CONDENSATE RECOVERY METER - MODEL CRM-485(CRM-485R) | No isue found, CRM calibrated by Savita | material to be dispatched against RJ OA 5525200121 | material to be dispatched against RJ OA 5525200121 | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFDL24048071599 | 2024-07-03 | 2024-07-16 | Cust 834 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | cal | Calibrated | calibrated and found working fine. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL24048541600 | 2024-07-04 | 2024-07-11 | Cust 810 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | cal | Calibrated and found working fine. | Calibrated and found working fine. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL24048741601 | 2024-07-04 | 2024-07-04 | Cust 806 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | HSPCB data upload issue. | We have modified the details as per the new API and now the data is live on the HSPCB server. | We have modified the details as per the new API and now the data is live on the HSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24048751602 | 2024-07-04 | 2024-07-04 | Cust 905 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | HSPCB data uploading issue. | We have modified the details as per the new API and now the data is live on the HSPCB server. | We have modified the details as per the new API and now the data is live on the HSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24048931603 | 2024-07-04 | 2024-07-06 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | AMC | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | Visit for CX 1000 PM BOARD Field trial in Oil in Water Analyzer. | Attended to the Oil in Water Analyzer, replaced the PM board for testing along with other required parts. Now, the analyzer is working satisfactorily. Details are in the attached MOM. | Attended to the Oil in Water Analyzer, replaced the PM board for testing along with other required parts. Now, the analyzer is working satisfactorily. Details are in the attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24048991604 | 2024-07-05 | 2024-07-05 | Cust 806 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on HSPCB server. since last 15 days . | We have modified the details as per the new API and now the data is live on the HSPCB server. | We have modified the details as per the new API and now the data is live on the HSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24049311605 | 2024-07-05 | 2024-07-05 | Cust 802 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | HSPCB new URL needs to be update. | We have modified the details as per the new API and now the data is live on the HSPCB server. | We have modified the details as per the new API and now the data is live on the HSPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24049321606 | 2024-07-05 | 2024-07-08 | Cust 929 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | HSPCB data upload issue. | We have modified the details as per the new API and now the data is live on the HSPCB server. | We have modified the details as per the new API and now the data is live on the HSPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24049331607 | 2024-07-05 | 2024-07-05 | Cust 786 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | HSPCB new URL needs to be update. | We have modified the details as per the new API and now the data is live on the HSPCB server. | We have modified the details as per the new API and now the data is live on the HSPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24049351608 | 2024-07-05 | 2024-07-05 | Cust 692 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | HSPCB new url needs to be uodate, | We have modified the details as per the new API and now the data is live on the HSPCB server. | We have modified the details as per the new API and now the data is live on the HSPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24049361609 | 2024-07-05 | 2024-07-05 | Cust 806 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | HSPCB new URL needs to be update | We have modified the details as per the new API and now the data is live on the HSPCB server. | We have modified the details as per the new API and now the data is live on the HSPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24049371610 | 2024-07-05 | 2024-07-05 | Cust 806 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | HSPCB new URL needs to be Update. | We have modified the details as per the new API and now the data is live on the HSPCB server. | We have modified the details as per the new API and now the data is live on the HSPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24049381611 | 2024-07-05 | 2024-07-08 | Cust 837 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | HSPCB new URL needs to be update. | As of now we are closing this complaint. We will reopen this complaint once we get access. | As of now we are closing this complaint. We will reopen this complaint once we get access. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24049391612 | 2024-07-05 | 2024-07-08 | Cust 865 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | SCADA Software | AIPL Software for CPCB(AIPL1) | HSPCB new URL needs to be Update. | We have modified the details as per the new API and now the data is live on the HSPCB server. | We have modified the details as per the new API and now the data is live on the HSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24049401613 | 2024-07-05 | 2024-07-08 | Cust 846 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | HSPCB new URL needs to be Update. | As of now we are closing this complaint. We will reopen this complaint once we get access. | As of now we are closing this complaint. We will reopen this complaint once we get access. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24049411614 | 2024-07-05 | 2024-07-08 | Cust 891 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | HSPCB new URL needs to be Update. | We will reopen this complaint once we get sytem access. | We will reopen this complaint once we get system access. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24049421615 | 2024-07-05 | 2024-07-05 | Cust 806 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | HSPCB data upload issue | We took access to the system and now data is live on HSPCB Server. | We took access to the system and now data is live on HSPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24049861616 | 2024-07-07 | 2024-07-13 | Cust 799 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Schedule visit for online Multi Parameter Analyzer in ETP against AMC. | Visit for online Multi-Parameter Analyzers have been completed, observed all the Analyzers are working satisfactorily. Details are in attached MOM. | Visit for online Multi-Parameter Analyzers have been completed, observed all the Analyzers are working satisfactorily. Details are in attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24049941617 | 2024-07-08 | 2024-07-08 | Cust 886 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | HSPCB new URL needs to be update. | We have modified the details as per the new API and now the data is live on the HSPCB server. | We have modified the details as per the new API and now the data is live on the HSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24049951618 | 2024-07-08 | 2024-07-11 | Cust 844 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | HSPCB new URL needs to be update. | We have modified the details as per the new API and now the data is live on the HSPCB server. | We have modified the details as per the new API and now the data is live on the HSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24049961619 | 2024-07-08 | 2024-07-08 | Cust 812 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | HSPCB new URL needs to be update. | We have modified the details as per the new API and now the data is live on the HSPCB server for ETP. | We have modified the details as per the new API and now the data is live on the HSPCB server for ETP. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24050301620 | 2024-07-08 | 2024-07-10 | Cust 806 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | HSPCB new URL needs to be update. | We have modified the details as per the new API and now the data is live on the HSPCB server. | We have modified the details as per the new API and now the data is live on the HSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24050721621 | 2024-07-09 | 2024-07-10 | Cust 839 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on RSPCB server. | Now data is live on RSPCB server. | Now data is live on RSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24050791622 | 2024-07-09 | 2024-07-10 | Cust 891 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | HSPCB new URL needs to be update. | We have modified the details as per the new API and now the data is live on the HSPCB server. | We have modified the details as per the new API and now the data is live on the HSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24051351623 | 2024-07-10 | 2024-07-12 | Cust 930 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | RSPCB data not updating. | We took access to the system and now data is live on rspcb server. | We took access to the system and now data is live on rspcb server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24051381624 | 2024-07-10 | 2024-07-10 | Cust 804 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on HSPCB server. | We took access to the system and completed the re-installation of the software, as well as the configuration for local + HSPCB server. Currently, the data is live on the Local + HSPCB server. | We took access to the system and completed the re-installation of the software, as well as the configuration for local + HSPCB server. Currently, the data is live on the Local + HSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24051491625 | 2024-07-10 | 2024-07-10 | Cust 837 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | HSPCB new URL needs to be update. | We have modified the details as per the new API and now the data is live on the HSPCB server. | We have modified the details as per the new API and now the data is live on the HSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24051981626 | 2024-07-11 | 2024-07-11 | Cust 841 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | HSPCB new URL needs to be update | We have modified the details as per the new API and now the data is live on the HSPCB server. | We have modified the details as per the new API and now the data is live on the HSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24052081627 | 2024-07-11 | 2024-07-12 | Cust 804 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | HSPCB new URL needs to be update. | We have modified the details as per the new API and now the data is live on the HSPCB server | We have modified the details as per the new API and now the data is live on the HSPCB server | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24052101628 | 2024-07-11 | 2024-07-12 | Cust 804 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | HSPCB new URL needs to be update. | We have modified the details as per the new API and now the data is live on the HSPCB server | We have modified the details as per the new API and now the data is live on the HSPCB server | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24052141629 | 2024-07-11 | 2024-07-12 | Cust 804 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | HSPCB new URL needs to be update. | We have modified the details as per the new API and now the data is live on the HSPCB server | We have modified the details as per the new API and now the data is live on the HSPCB server | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24052171630 | 2024-07-11 | 2024-07-12 | Cust 804 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | HSPCB new URL needs to be update. | We have modified the details as per the new API and now the data is live on the HSPCB server | We have modified the details as per the new API and now the data is live on the HSPCB server | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24052221631 | 2024-07-11 | 2024-07-12 | Cust 804 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | HSPCB new URL needs to be update. | We have modified the details as per the new API and now the data is live on the HSPCB server. | We have modified the details as per the new API and now the data is live on the HSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24052241632 | 2024-07-11 | 2024-07-12 | Cust 804 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | HSPCB new URL needs to be update. | We have modified the details as per the new API and now the data is live on the HSPCB server. | We have modified the details as per the new API and now the data is live on the HSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24052771633 | 2024-07-12 | 2024-07-12 | Cust 804 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | HSPCB new URL needs to be update. | We have modified the details as per the new API and now the data is live on the HSPCB server. | We have modified the details as per the new API and now the data is live on the HSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24052811634 | 2024-07-12 | 2024-07-12 | Cust 692 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Localhost not opening | We have re-installed the software with updated version and data is now live on the Local, CPCB and HSPCB servers. | We have re-installed the software with updated version and data is now live on the Local, CPCB and HSPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24053221635 | 2024-07-13 | 2024-07-17 | Cust 862 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Display of Multi-Parameter Analyzer to be replaced, and then calibration to be done. | Attended Multi-Parameter Analyzer, Replaced new display board, Calibration done, observed the Analyzer is working satisfactorily. Details are in attached MOM. | Attended Multi-Parameter Analyzer, Replaced new display board, Calibration done, observed the Analyzer is working satisfactorily. Details are in attached MOM. | EXTENDED | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24053241636 | 2024-07-13 | 2024-07-17 | Cust 862 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | All the data (COD, BOD, TSS & pH) were zero on CPCB server. | The wrong readings of the online Multi-Parameter Analyzer, which were uploading as almost zero on the CPCB server, have been resolved through telephonic support by our Suneet Kumar. Now all the data is uploading with the correct values, hence this complaint is now closed. | Multi-Parameter Analyzer readings was uploading zero on CPCB server, have been resolved through telephonic support by Suneet Kumar. Now all the data is uploading with the correct values. | EXTENDED | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24053781637 | 2024-07-15 | 2024-07-18 | Cust 931 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Not able to see real time data on the RSPCB dashboard. | Issue is from RSPCB server end. hence as of now we are closing this complaint. | Issue is from RSPCB server end. hence as of now we are closing this complaint. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24053881638 | 2024-07-15 | 2024-07-15 | Cust 830 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | HSPCB new URL needs to be update. | We have modified the details as per the new API and now the data is live on the HSPCB server for MATRIX CLOTHING PVT. LTD- MOHAMMADPUR. | We have modified the details as per the new API and now the data is live on the HSPCB server for MATRIX CLOTHING PVT. LTD- MOHAMMADPUR. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24053921639 | 2024-07-15 | 2024-07-15 | Cust 830 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | HSPCB new URL needs to be update | We have modified the details as per the new API and now the data is live on the HSPCB server for MATRIX CLOTHING PVT. LTD -SECTOR 37. | We have modified the details as per the new API and now the data is live on the HSPCB server for MATRIX CLOTHING PVT. LTD -SECTOR 37. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24054001640 | 2024-07-15 | 2024-07-17 | Cust 837 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Contact details need to be updated on FMCPCB server. | we have now also updated the contact details on our website. | we have now also updated the contact details on our website. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24054051641 | 2024-07-15 | 2024-07-17 | Cust 788 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | DATCON | DATCON(DATCON) | HSPCB new URL needs to be update. | We took access of the DATCon and made modifications as per the new API and now the data is live on the HSPCB server . | We took access of the DATCon and made modifications as per the new API and now the data is live on the HSPCB server . | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24054091642 | 2024-07-15 | 2024-07-17 | Cust 786 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 3rd AMC visit | 3rd AMC visit completed | Working fine | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24054481643 | 2024-07-16 | 2024-07-23 | Cust 932 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Transmitter Aquacon | Aquacon 2000(AQCN 2000 p) | As per the customer the conductivity is not working properly. Site address are as below Ind Swift Ltd Jawaharpur Village, OFFNH 21,Derabasi Jawaharpur, Punjab ? 140507 Contact Name. : Mr. Amul (IUS Engineer) 86550 40251 | Checked transmitter with simulator and analyzer taken in process, Reading verified and found ok. | Work done | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24054591644 | 2024-07-16 | 2024-07-23 | Cust 890 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | User reported TSS showing 0 or high for Ecotech, Kapurthala STP | Issue found that analyzer run in no sample condition hence stain found in flowcell assembly. Cleaned and performed calibration. Checked operation in multiple solution, working found ok. | work done | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24054881645 | 2024-07-17 | 2024-07-23 | Cust 849 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | PH1 detector(1) | To check the CETP inlet analyzer | PH1 detector found defective and replaced on 18th July from RO permeate analyzer, replacement arranged sent from Pune on 19th July | Working fine | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24055281646 | 2024-07-17 | 2024-07-19 | Cust 832 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Visit to be organize for Phosphate Analyzer revive. | Attended the Phosphate Analyzer, Now Analyzer is working satisfactorily. Details are in attached MOM. | Attended the Phosphate Analyzer, Now Analyzer is working satisfactorily. Details are in attached MOM. | EXTENDED | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24055701647 | 2024-07-18 | 2024-07-20 | Cust 608 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Checking and calibration of Multiparameter. | Done the calibration of MPA | Done the calibration of Multiparameter. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL24055771648 | 2024-07-18 | 2024-07-19 | Cust 804 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Data upload issue on HSPCB server | Now the data is live on the HSPCB server. | Now the data is live on the HSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24055811649 | 2024-07-18 | 2024-07-19 | Cust 933 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | DATCON | DATCON(DATCON) | CPCB and HSPCB data uploded issue. | We took access of the DATCon and now the data is live on the CPCB and HSPCB server | We took access of the DATCon and now the data is live on the CPCB and HSPCB server | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24056231650 | 2024-07-19 | 2024-07-19 | Cust 846 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Scheduled AMC visit for ETP Multi Parameter Analyzer. | Visit for online Multi-Parameter Analyzer has been completed, observed Analyzer is working satisfactorily. Details are in attached MOM. | Visit for online Multi-Parameter Analyzer has been completed, observed Analyzer is working satisfactorily. Details are in attached MOM. | EXTENDED | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24056261651 | 2024-07-19 | 2024-07-19 | Cust 846 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | HSPCB data uploading issue. | We took access to the system and made modifications as per the new API and now the data is live on the HSPCB server. | We took access to the system and made modifications as per the new API and now the data is live on the HSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24056271652 | 2024-07-19 | 2024-07-23 | Cust 934 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | HSPCB new URL needs to be update. | We have modified the details as per the new API and now the data is live on the HSPCB server. | We have modified the details as per the new API and now the data is live on the HSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24056811653 | 2024-07-21 | 2024-07-26 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | SULPHIDE ANALYSER | SULPHIDE 8810 Single Channel ANALYSER(8810 Sulphide) | visit for checking of sulfide analyzer | visit done as per attached mail communication. | visit done as per attached mail communication. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24056951654 | 2024-07-21 | 2024-07-24 | Cust 858 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Schedule visit for online Multi Parameter Analyzer in ETP against AMC. | Visit for online Multi-Parameter Analyzer has been completed against AMC, observed Now the Analyzer is working satisfactorily. Details are in attached MOM. | Visit for online Multi-Parameter Analyzer has been completed against AMC, observed Now the Analyzer is working satisfactorily. Details are in attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24057141655 | 2024-07-22 | 2024-07-23 | Cust 804 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on CPCB + Local + HSPCB servers. | We took access to the system and successfully completed the software re-installation. Now the data is currently live on the CPCB + Local + HSPCB servers. | We took access to the system and successfully completed the software re-installation. Now the data is currently live on the CPCB + Local + HSPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24057281656 | 2024-07-22 | 2024-07-22 | Cust 839 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on RSPCB server. | Now the data is live on the RSPCB server. | Now the data is live on the RSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24058141657 | 2024-07-23 | 2024-07-29 | Cust 935 | DELHI | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | Ph Electrode | Chemotrode 120(HT238 760) | installation Guideline for MPA | Installation Guideline given to the customer. | Installation Guidelines given to the customer. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24058561658 | 2024-07-24 | 2024-07-26 | Cust 803 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 1st AMC visit for Multiparameter analyzer, dust analyzer and 2nos. mag flow meters | 1st AMC visit completed | Working fine | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24058661659 | 2024-07-24 | 2024-08-01 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Schedule visit for ETP#3 Analyzers as COD, Phenol & Sulphide Analyzers against AMC. | Visit for COD, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | Visit for COD, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24058861660 | 2024-07-24 | 2024-07-26 | Cust 807 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Visit for MPA under 2nd qtr | visit done as per attached mom. | visit done as per attached mom. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24059191661 | 2024-07-25 | 2024-07-27 | Cust 841 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on HSPCB server. | We took access to the system, and found that the EnviroconnectDAS service has been stopped on local PC, which means that data is not being transmitted from the local PC to the CPCB, local, and HSPCB servers. After restarting the same, the data is now live on CPCB + Local + HSPCB servers. | now the data is live on CPCB + Local + HSPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24059801662 | 2024-07-25 | 2024-07-28 | Cust 825 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning of multiparameter Analyzer | MPA commissioning is done successfully. | MPA commissioning is done successfully. | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFDL24059811663 | 2024-07-25 | 2024-07-28 | Cust 826 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Parameter was showing incorrect. | TSS valve is under 25mg/l, it will show when valve will be above 25mg/l | TSS valve is under 25mg/l, it will show when valve will be above 25mg/l | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFDL24060531664 | 2024-07-27 | 2024-08-01 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | AMC | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | Scheduled AMC visit for COD, TOC, Oi l in Water, Sulphide Analyzers, installed at ETP # 1, 2 & TTPRO. | Visit for COD, TOC, Oil in Water, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | Visit for COD, TOC, Oil in Water, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24062181665 | 2024-07-29 | 2024-08-01 | Cust 546 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on RSPCB Server. | We took access to the system and now data is live on RSPCB Server. | We took access to the system and now data is live on RSPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24063031666 | 2024-07-29 | 2024-08-02 | Cust 890 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Reported that TSS reading fluctuated in Goraya STP | Work done | work done | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24063451667 | 2024-07-30 | 2024-08-02 | Cust 858 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | DATCON | DATCON(DATCON) | CPCB and HSPCB data uploaded issue. | Now the internet issue has been resolved from the customer end and data is live on both sites. | Now the internet issue has been resolved from the customer end and data is live on both sites. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24064431668 | 2024-07-31 | 2024-08-02 | Cust 804 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not updating to CPCB and HSPCB server. | Data is now live on the CPCB and HSPCB servers. | Data is now live on the CPCB and HSPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24065421669 | 2024-08-01 | 2024-08-04 | Cust 786 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Schedule visit for ETP online Analyzer against CAMC. | Visit for online Multi-Parameter Analyzer has been completed against CAMC, observed Analyzer is working satisfactorily. Details are in attached MOM. | Visit for online Multi-Parameter Analyzer has been completed against CAMC, observed Analyzer is working satisfactorily. Details are in attached MOM. | EXTENDED | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24065681670 | 2024-08-02 | 2024-08-07 | Cust 936 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | camera streaming is not live on CPCB server. | We accessed the system and found that the customer's PC was frequently freezing, which prevented the camera from staying continuously live. The customer is now working with their IT team to resolve these issues. Therefore, we are closing this complaint. Once we get access we will reopen this complaint. | The customer is now working with their IT team to resolve this issue. Therefore, we are closing this complaint. Once we get access we will reopen this complaint. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24065771671 | 2024-08-02 | 2024-08-02 | Cust 804 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | HSPCB data uploading issue. | We have re-installed the software with the updated version and the data is now live on the CPCB and HSPCB servers. | We have re-installed the software with the updated version and the data is now live on the CPCB and HSPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24066281672 | 2024-08-05 | 2024-08-12 | Cust 918 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Smartpro-con) | Customer reported for commissioning of PH, Conductivity & Chlorine transmitters. | calibration | Commissioning done analyser found working ok | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24066301673 | 2024-08-05 | 2024-08-12 | Cust 824 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Customer reported for calibration | calibration done | Calibration done. | NO | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24066311674 | 2024-08-05 | 2024-08-12 | Cust 866 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Customer reported for calibration | calibration done | Calibration done. | NO | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24067371675 | 2024-08-06 | 2024-08-08 | Cust 546 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on RSPCB Server. | We took access to system and data is updating on RSPCB Server. | We took access to system and data is updating on RSPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24067421676 | 2024-08-06 | 2024-08-09 | Cust 863 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | HSPCB new URL needs to be updated | we have modified the details as per the new API and now the data is live on the HSPCB server. | we have modified the details as per the new API and now the data is live on the HSPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24067641677 | 2024-08-06 | 2024-08-08 | Cust 786 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Emergency AMC visit for ETP Multi Parameter Analyzer due to alert of TSS high readings. | Emergency visit for Multi-Parameter Analyzer has been completed, Observed now the Analyzer is working satisfactorily, Details are in attached MOM. | Emergency visit for Multi-Parameter Analyzer has been completed, Observed now the Analyzer is working satisfactorily, Details are in attached MOM. | EXTENDED | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24067651678 | 2024-08-06 | 2024-08-08 | Cust 849 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | PH1 detector(1) | PH1 Detector to be replaced in RO Feed Water Multi-Parameter Analyzer. | Replaced PH1 Detector in RO Multi-Parameter Analyzer, Calibration done, Now the Analyzer is working satisfactorily. Details are in attached MOM. | Replaced PH1 Detector in RO Multi-Parameter Analyzer, Calibration done, Now the Analyzer is working satisfactorily. Details are in attached MOM. | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24067991679 | 2024-08-07 | 2024-08-09 | Cust 804 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB and HSPCB data uploading issue | Data is now live on the CPCB and HSPCB server. | Data is now live on the CPCB and HSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24068031680 | 2024-08-07 | 2024-08-12 | Cust 804 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | HSPCB data upload issue. | We have modified the details as per the new API and now the data is live on the HSPCB server. | We have modified the details as per the new API and now the data is live on the HSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24068861681 | 2024-08-09 | 2024-08-16 | Cust 832 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | AQUAMAX DOT(1) | Commissioning of DO Analyzer in WTP Plant. | Commissioned DO Analyzer, Observed working satisfactorily. Details are in attached MOM. | Commissioned DO Analyzer, Observed working satisfactorily. Details are in attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24068921682 | 2024-08-09 | 2024-08-09 | Cust 546 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on RSPCB Server. | We took access to the system and upgraded the software version now data is live on RSPCB Server. | We took access to the system and upgraded the software version now data is live on RSPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24068991683 | 2024-08-10 | 2024-08-14 | Cust 790 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 4th AMC visit | 4th AMC visit completed | Working fine | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24069151684 | 2024-08-10 | 2024-08-16 | Cust 571 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | pH Analyzers to be commissioned in Maruti Suzuki India Ltd, Gurgaon unit. | Commissioned Conductivity, pH & Turbidity Analyzers in ZLD Plant, Observed working satisfactorily. Details are in attached MOM. | Commissioned Conductivity, pH & Turbidity Analyzers in ZLD Plant, Observed working satisfactorily. Details are in attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24069171685 | 2024-08-10 | 2024-08-22 | Cust 854 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Multiparameter analyzer tubes needs to be replace and issue in pH parameter. | pH card needs to be replace. | pH card needs to replace offer has been send. | YES | Periodic Maintenance (AMC) | Process | N |  | Analytical - CHAKAN
AFDL24069181686 | 2024-08-10 | 2024-08-22 | Cust 828 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Multiparameter Analyzer datcon is received at site needs to install it and set the parameters. | CPCB connection is to be done. | CPCB connection is to be done. | YES | Periodic Maintenance (AMC) | Process | N |  | Analytical - CHAKAN
AFDL24070611687 | 2024-08-13 | 2024-08-16 | Cust 854 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Data Mismatching issue on CPCB srever. | we took access of the Blackbox and till Blackbox, we are getting 0 readings only for pH and BOD parameters and we have also checked the data on modscan and in modscan also the 4000 counts are coming for both parameters hence this issue need to be checked from the hardware side. | this issue need to be checked from the hardware side. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24070741688 | 2024-08-13 | 2024-08-14 | Cust 804 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on CPCB + Local + HSPCB servers. | We took access to the system and successfully completed the software re-installation. And now the data is live on CPCB + Local + HSPCB server. | We took access to the system and successfully completed the software re-installation. And now the data is live on CPCB + Local + HSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24071171689 | 2024-08-14 | 2024-08-14 | Cust 839 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on RSPCB server. | We took access to the system and now data is live on the RSPCB server. | We took access to the system and now data is live on the RSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24071211690 | 2024-08-14 | 2024-08-26 | Cust 807 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Blackbox hanging issue. | Now data is live on CPCB server with 99 % uptime. | Now data is live on CPCB server with 99 % uptime. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | DCS - CHAKAN
AFDL24071811691 | 2024-08-16 | 2024-08-17 | Cust 937 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data upload issue. | today we took access of the black box, and completed internet connectivity, now internet is coming till the putty and data is live on CPCB server. | today we took access of the black box, and completed internet connectivity, now internet is coming till the putty and data is live on CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24072631692 | 2024-08-19 | 2024-08-20 | Cust 841 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | The EnviroConnect website was not accessible for the local server. | We accessed the system and found that the Apache service was stopped, which was why the EnviroConnect website wasn?t opening for the local server connected to the Boiler OCEMS. We have now restarted the Apache service, and the website is now accessible on that system. | We have now restarted the Apache service, and the website is now accessible on that system. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24072671693 | 2024-08-19 | 2024-08-22 | Cust 846 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | data is not uploading to central portal | data is live on the central portal. | data is live on the central portal. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24072841694 | 2024-08-19 | 2024-08-24 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Emergency visit against AMC for COD Analyzer in ETP#3 | Attended UFM of SO2 Gas. Now COD Analyzer is working satisfactorily. | Attended COD Analyzer. Now COD Analyzer is working satisfactorily. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24073461695 | 2024-08-20 | 2024-08-22 | Cust 844 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | DATCON | DATCON(DATCON) | Data is not uploading to the CPCB and HSPCB servers. | Data is live on the CPCB and HSPCB servers. | Data is live on the CPCB and HSPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24073651696 | 2024-08-20 | 2024-08-23 | Cust 938 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB connectivity needs to be done. | As of now we are closing this complaint, and will reopen it once Black box access is received. | As of now we are closing this complaint, and will reopen it once Black box access is received. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24073671697 | 2024-08-20 | 2024-08-23 | Cust 804 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB readings are not matching with actual readings. | We have accessed the system and found that the data reaching till ModScan is only being uploaded to the CPCB, LOCAL, and HSPCB servers. So this issue is from the hardware side . So we are closing this complaint and will reopen this once the hardware side issue is resolved. | This issue is from the hardware side . So we are closing this complaint and will reopen this once the hardware side issue is resolved. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24074291698 | 2024-08-21 | 2024-08-23 | Cust 939 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data uploaded issue. | now the data is live on the CPCB server. | now the data is live on the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24074601699 | 2024-08-21 | 2024-08-25 | Cust 571 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Smartpro Electronics Assembly(Smartpro-ele) | Visit for commissioning of multiple Analyzer at MSIL, Manesar unit. | Commissioned pH, Conductivity & Turbidity Analyzers, Observed working satisfactorily. Details are in attached MOM. | Commissioned pH, Conductivity & Turbidity Analyzers, Observed working satisfactorily. Details are in attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24074631700 | 2024-08-22 | 2024-08-24 | Cust 809 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 1st AMC visit for Oil in water and multiparameter analyzer | 1st AMC visit completed | Working fine | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24075021701 | 2024-08-22 | 2024-08-24 | Cust 532 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on CPCB + Local + RSPCB servers | We took access to the system and observed that data is not coming till the local system (Checked on MODSCAN). So issue was from the hardware side. And now the issue has been resolved. Now data is live . | Now data is live . | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24075031702 | 2024-08-22 | 2024-08-24 | Cust 940 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB connectivity needs to be done. | As of now we are closing this complaint and will reopen it once we get internet details and BB access. | As of now we are closing this complaint and will reopen it once we get internet details and BB access. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24075081703 | 2024-08-22 | 2024-08-24 | Cust 732 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | EnviroConnect Software installation | We have accessed the system and successfully completed the software installation for new PC. As discussed with the customer since the old PC is no longer operational and customer is unable to trace its name, we have escalated this issue to our development team for further assistance. For now, we are closing this complaint. We will reopen it once we receive an update from our development team. | As discussed with the customer since the old PC is no longer operational and customer is unable to trace its name, we have escalated this issue to our development team for further assistance. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24075101704 | 2024-08-22 | 2024-08-23 | Cust 546 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB and RSPCB Server. | We took access to your system and now data is live on CPCB and RSPCB Server. | We took access to your system and now data is live on CPCB and RSPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24075461705 | 2024-08-23 | 2024-08-28 | Cust 909 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Aqua4Trans | Aqua4 trans-A4TW033000(1) | Checking and calibration of Chlorine and Turbidity Analyser | Done the calibration of Analyser | Checked and Done the calibration of Analyzers. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24075491706 | 2024-08-23 | 2024-08-29 | Cust 941 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | issue with the DO sensor | Do sensors have been checked, sensor cap sr. no. 2898 for 15 KLD STP found faulty and needs to be procured by the customer. Both DO sensors are working fine. | Working fine | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24075511707 | 2024-08-23 | 2024-08-26 | Cust 802 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB and HSPCB data uploded issue. | As of now we are closing this complaint and will reopen it once we get system access. | As of now we are closing this complaint and will reopen it once we get system access. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24075921708 | 2024-08-24 | 2024-08-27 | Cust 532 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | To check the analyzer for Modbus output and other issues | Service and calibration visit completed | Working fine | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL24076491709 | 2024-08-25 | 2024-08-29 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Schedule visit for ETP#3 Analyzers as COD, Phenol & Sulphide Analyzers against AMC. | Visit for COD, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | Visit for COD, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24076851710 | 2024-08-26 | 2024-08-29 | Cust 906 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not updating to CPCB server. | Data is live on server. | Data is live. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24076911711 | 2024-08-26 | 2024-08-29 | Cust 804 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to CPCB server. | Data is not coming till the modscan. As of now we are closing this complaint. We will reopen this once the issue is resolved from the hardware side. | Data is not coming till the modscan. As of now we are closing this complaint. We will reopen this once the issue is resolved from the hardware side. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24077161712 | 2024-08-26 | 2024-08-29 | Cust 940 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB connectivity needs to be done. | we took access of the black box and completed CPCB connectivity, Now data is live on the CPCB server. | we took access of the black box and completed CPCB connectivity, Now data is live on the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24077431713 | 2024-08-26 | 2024-09-01 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | AMC | TOC ANALYSER | FM 101 TOC Analyser(FM 101 TOC) | Scheduled AMC visit for COD, TOC, Oi l in Water, Sulphide Analyzers, installed at ETP # 1, 2 & TTPRO. & BSVI for TOC | Attended the TOC Analyzers against ETP#1 Analyzers AMC, Servicing and calibration done. Details are in the attached MOM. | Attended the TOC Analyzers against ETP#1 Analyzers AMC, Servicing and calibration done. Details are in the attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24077721714 | 2024-08-27 | 2024-08-27 | Cust 732 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | software installation | We took access to the system and now data is live on the CPCB server. | We took access to the system and now data is live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24077941715 | 2024-08-27 | 2024-08-29 | Cust 788 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 1st AMC visit | 1st AMC visit completed | Working fine | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24078341716 | 2024-08-27 | 2024-09-01 | Cust 856 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | SODIUM ANALYSER | SODIUM 9245 Multi Channel ANLAYSER(9245 Sodium) | visit for checking & calibration of Sodium & silica analyzers SWAN make | visit done under AMC for SWAN Make Silica & Sodium Analyzer at SCL | visit done under AMC for SWAN Make Silica & Sodium Analyzer at SCL | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24078391717 | 2024-08-28 | 2024-08-29 | Cust 788 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 4th AMC visit | 4th AMC visit completed | Working fine | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24078451718 | 2024-08-28 | 2024-08-30 | Cust 844 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | DATCON | DATCON(DATCON) | Data is not uploading to CPCB and HSPCB server. | Data is live on the CPCB and HSPCB server. | Data is live on the CPCB and HSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24078591719 | 2024-08-28 | 2024-08-30 | Cust 802 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on HSPCB server. | Now the issue has been resolved and the data is live on both the servers. | Now the issue has been resolved and the data is live on both the servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24079241720 | 2024-08-29 | 2024-08-30 | Cust 802 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB and HSPCB data uploded issue. | Now the issue has been resolved and the data is live on both the servers. | Now the issue has been resolved and the data is live on both the servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24079321721 | 2024-08-29 | 2024-09-01 | Cust 801 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 1st AMC visit | 1st AMC visit completed | Working fine | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24079831722 | 2024-08-30 | 2024-09-02 | Cust 806 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | cpcb and hspcb data uploaded issue. | This issue is from the hardware side.so we are closing this complaint. | This issue is from the hardware side.so we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24079881723 | 2024-08-30 | 2024-09-05 | Cust 908 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Replacement of lens and recommissioning | Replaced the lens,Ph2 detector set and Xenon Lamp Power supply and recommissioned the Analyser | Commissioned the MPA | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24082141724 | 2024-09-01 | 2024-09-06 | Cust 849 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | Peristaltic sampling pump motor for Multiparameter(AWA-PSPM) | Sample pump & Xenon power supply board to be replaced in RO Feed Multi-Parameter Analyzer. | Replaced sample pump motor and Xenon lamp power supply in RO Multi-Parameter Analyzer, Calibration done, Now the Analyzer is working satisfactorily. Details are in attached MOM. | Replaced sample pump motor and Xenon lamp power supply in RO Multi-Parameter Analyzer, Calibration done, Now the Analyzer is working satisfactorily. Details are in attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24082681725 | 2024-09-02 | 2024-09-05 | Cust 894 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | HSPCB data uploaded issue. | Now the Data is live on the HSPCB server. | Now the Data is live on the HSPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24082691726 | 2024-09-02 | 2024-09-05 | Cust 940 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Data mismatch issue and camera issue. | Now data is matching with the actual readings Today we took access to your system and completed Camera configuration, and now Camera is streaming live on CPCB server. | Now data is matching with the actual readings Today we took access to your system and completed Camera configuration, and now Camera is streaming live on CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24082731727 | 2024-09-02 | 2024-09-05 | Cust 803 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera streaming issue | Camera is now live on the CPCB server. | Camera is now live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24083011728 | 2024-09-03 | 2024-09-06 | Cust 842 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Online Data Transfer Problem | commercially not clear | commercially not clear | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24083151729 | 2024-09-03 | 2024-09-05 | Cust 844 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | DATCON | DATCON(DATCON) | HSPCB data uploading issue. | We have changed the internet IP address as provided by you and the data is now live on the HSPCB server. | We have changed the internet IP address as provided by you and the data is now live on the HSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24083211730 | 2024-09-03 | 2024-09-06 | Cust 790 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not updating to RSPCB and CPCB servers. | We have re-installed the software and the data is now live on the Local + CPCB + RSPCB servers. | We have re-installed the software and the data is now live on the Local + CPCB + RSPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24084541731 | 2024-09-06 | 2024-09-15 | Cust 802 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 3rd AMC visit | Done the AMC visit | Done the calibration of Multiparameter. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24085671732 | 2024-09-09 | 2024-09-12 | Cust 905 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to HSPCB server | Data is now live on the Local & HSPCB servers. | Data is now live on the Local &amp;HSPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24085921733 | 2024-09-09 | 2024-09-12 | Cust 844 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | DATCON | DATCON(DATCON) | HSPCB data uploading issue | now the internet issue is resolved from the customer side and the data is live on the HSPCB server. | now the internet issue is resolved from the customer side and the data is live on the HSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24085931734 | 2024-09-09 | 2024-09-12 | Cust 900 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera streaming is not live on CPCB server. | We took access to the system and now camera streaming is live on CPCB server . | We took access to the system and now camera streaming is live on CPCB server . | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24086041735 | 2024-09-09 | 2024-09-12 | Cust 790 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera is not live on CPCB and RSPCB server. | Issue has been resolved and the camera streaming is now live on the CPCB server. For RSPCB, we have sent an RTSP link for camera configuration. | Issue has been resolved and the camera streaming is now live on the CPCB server. For RSPCB, we have sent an RTSP link for camera configuration. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24086271736 | 2024-09-09 | 2024-09-12 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | AMC | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | Emergency AMC visit for COD, TOC, Oi l in Water, Sulphide Analyzers, installed at ETP # 1, 2 & TTPRO. | Emergency visit against ETP#1 AMC for Oil in water Analyzer has been completed, Observed now the Analyzer is working satisfactorily, Details are in attached MOM. | Emergency visit for Oil in water Analyzer has been completed, Observed now the Analyzer is working satisfactorily, Details are in attached MOM. | EXTENDED | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24086341737 | 2024-09-09 | 2024-09-15 | Cust 599 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | DO Transmitters | VISI pro DO(VISI pro DO) | Calibration not done | Visiwater DO Parc FC10 sensor mA output generate but not communicate on modbus hence calibration not done, In open air DO got 3.78mg/l | Work done | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24086831738 | 2024-09-11 | 2024-09-13 | Cust 841 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB and HSPCB DATA UPLOADED ISSUE. | We took access to the system and now data is live on HSPCB server. | We took access to the system and now data is live on HSPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24086931739 | 2024-09-11 | 2024-09-13 | Cust 846 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | OCEMS connected system C drive disk have full and has 0 bytes free space balance | We have made the necessary changes, and the C Drive now has enough free space. | We have made the necessary changes, and the C Drive now has enough free space. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24087081740 | 2024-09-11 | 2024-09-13 | Cust 942 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | DATCON | DATCON(DATCON) | CPCB and HSPCB darta uploaded issue. | Now the issue has been resolved and data is live on the CPCB and HSPCB server. | Now the issue has been resolved and data is live on the CPCB and HSPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24087161741 | 2024-09-11 | 2024-09-18 | Cust 837 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Schedule visit for online Multi Parameter Analyzer in STP against AMC. | Visit for online Multi-Parameter Analyzer has been completed against AMC, observed Analyzer is working satisfactorily. Details are in attached MOM. | Visit for online Multi-Parameter Analyzer has been completed against AMC, observed Analyzer is working satisfactorily. Details are in attached MOM. | EXTENDED | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24087361742 | 2024-09-12 | 2024-09-13 | Cust 804 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to HSPCB server | Data is now live on the Local and HSPCB server. | Data is now live on the Local and HSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24087751743 | 2024-09-13 | 2024-09-14 | Cust 881 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4Trans(1) | Commissioning of 1 FRC, 2 Turbidity and 2 pH | Commissioning completed successfully | Working fine | EXTENDED | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24087791744 | 2024-09-13 | 2024-09-19 | Cust 912 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning | Work done | work done | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24087901745 | 2024-09-13 | 2024-09-16 | Cust 863 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB and HSPCB data uploaded issue. | Now the issue has been resolved as the data is showing live on both sites. | Now the issue has been resolved as the data is showing live on both sites. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24088101746 | 2024-09-14 | 2024-09-17 | Cust 533 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 1st AMC visit | 1st AMC visit completed | Working fine | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24088801747 | 2024-09-16 | 2024-09-23 | Cust 824 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | customer reported for checking of MPA system | Calibration Checked found ok | Work Done | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24088821748 | 2024-09-16 | 2024-09-19 | Cust 915 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB + UPPCB data upload issue. | Data is continuously live on CPCB + UPPCB server from the last 2 days. | Data is continuously live on CPCB + UPPCB server from the last 2 days. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24088881749 | 2024-09-16 | 2024-09-21 | Cust 943 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | COMMISSIONING the Multi Parameter Analyzer at Hisar site. | Commissioned Multi-Parameter Analyzer and its working satisfactorily. Details are in attached MOM. | Commissioned Multi-Parameter Analyzer and its working satisfactorily. Details are in attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24089691750 | 2024-09-17 | 2024-09-20 | Cust 843 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 2nd AMC viisit | 2nd AMC visit completed | Working fine | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24089721751 | 2024-09-17 | 2024-09-24 | Cust 853 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 4th AMC | Work done | work done | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24089731752 | 2024-09-17 | 2024-09-24 | Cust 924 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | VISI pro DO(VISI pro DO) | Commissioning | Work done | work done | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24089781753 | 2024-09-17 | 2024-09-21 | Cust 786 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Schedule visit for ETP online Analyzer against CAMC. | Visit for online Multi-Parameter Analyzer has been completed, observed Analyzer is working satisfactorily. Details are in attached MOM. | Visit for online Multi-Parameter Analyzer has been completed, observed Analyzer is working satisfactorily. Details are in attached MOM. | EXTENDED | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24090101754 | 2024-09-18 | 2024-09-22 | Cust 896 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | UP Housing Board 21.5 MLD STP, Secotr-2, Near UP Jal Nigam Siddharth Vihar, Ghaziabad 201009 Uttar Pradesh India GSTIN 27AICPM5183C1ZE Site Contact person - Mr. Rafique- +91 90013 20128 and Mr. Sonu- +917217783305 1st Point of Contact- Soumitra Deshmukh- 9960162335 | visit done as per attached MOM .All DO analyzers are working fine. | visit done as per attached MOM .All DO analyzers are working fine. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24090161755 | 2024-09-18 | 2024-09-18 | Cust 804 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | EnviroConnect software re-installation. | We accessed the system, successfully re-installed the EnviroConnect software, and completed the local server configuration. Now, the data is live on both the CPCB and local servers. | We accessed the system, successfully re-installed the EnviroConnect software, and completed the local server configuration. Now, the data is live on both the CPCB and local servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24090701756 | 2024-09-19 | 2024-09-19 | Cust 816 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on RSPCB server. & Data availabillity is less than 60% on RSPCB | We accessed the system and noticed an issue with the Stack 1_Boiler_8MW. The analyzer data is not communicating properly with the local PC. We checked the data using Modscan and found that out of 35 polls, we are only receiving 15 responses. The data keeps breaking continuously. which indicates low availability on the RSPCB server. This issue is from the hardware side. So we are closing this complaint. | This issue is from the hardware side. So we are closing this complaint. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24090801757 | 2024-09-19 | 2024-09-20 | Cust 844 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | DATCON | DATCON(DATCON) | Data uploaing issue on HSPCB server | Data is now live on the HSPCB server. | Data is now live on the HSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24090841758 | 2024-09-19 | 2024-09-20 | Cust 830 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB + HSPCB data upload issue. | we took the access to your system and found that EnviroconnectDAS service is stopped in your local PC. After restarting the same, data is now live on the CPCB + HSPCB server. | we took the access to your system and found that EnviroconnectDAS service is stopped in your local PC. After restarting the same, data is now live on the CPCB + HSPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24090851759 | 2024-09-19 | 2024-09-20 | Cust 830 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB + HSPCB data upload issue. | we took access to your system and now Communication PORT is detecting and the data is live on CPCB + HSPCB server and matching with actual readings and the same is verified by you. | we took access to your system and now Communication PORT is detecting and the data is live on CPCB + HSPCB server and matching with actual readings and the same is verified by you. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24090861760 | 2024-09-19 | 2024-09-20 | Cust 818 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Local Host data mismatch issue. | We took access to your system and did some modifications in Local Configuration, now data is matching with actual readings on the Local server and the same is verified by you. | We took access to your system and did some modifications in Local Configuration, now data is matching with actual readings on the Local server and the same is verified by you. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24090871761 | 2024-09-19 | 2024-09-20 | Cust 891 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | SPM data mismatch issue. | we took the access to your system and checked the data on MODSCAN and as per calculations the same data is reflected on the CPCB server for Boiler Stack. | we took the access to your system and checked the data on MODSCAN and as per calculations the same data is reflected on the CPCB server for Boiler Stack. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24091391762 | 2024-09-21 | 2024-09-22 | Cust 786 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Schedule visit for ETP online Analyzer against CAMC. | Visit for online Multi-Parameter Analyzer has been completed, observed Analyzer is working satisfactorily after replacing the spares. Details are in attached MOM. | Visit for online Multi-Parameter Analyzer has been completed, observed Analyzer is working satisfactorily after replacing the spares. Details are in attached MOM.Details are in attached MOM. | EXTENDED | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24091821763 | 2024-09-22 | 2024-09-28 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | AMC | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | Scheduled AMC visit for COD, TOC, Oi l in Water, Sulphide Analyzers, installed at ETP # 1, 2 & TTPRO. | Visit for COD, TOC, Oil in Water, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | Visit for COD, TOC, Oil in Water, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | EXTENDED | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24091891764 | 2024-09-23 | 2024-09-25 | Cust 800 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on CPCB + Local + SPCB servers. | We accessed the system and successfully completed the re-installation of the EnviroConnect software. and completed the configuration for Local + SPCB And the data is now live on the CPCB + Local + SPCB servers. | We accessed the system and successfully completed the re-installation of the EnviroConnect software. And the data is now live on the CPCB + Local + SPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24092431765 | 2024-09-23 | 2024-09-23 | Cust 802 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on CPCB + Local + HSPCB servers. | We took access to the system and now data is live on CPCB + Local + HSPCB servers. | We took access to the system and now data is live on CPCB + Local + HSPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24092991766 | 2024-09-24 | 2024-09-27 | Cust 796 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data availability is low on RSPCB server. | We accessed to the system and noticed an issue with the Stack 1_8MW_Boiler_ESP. The analyzer data is not communicating properly with the local PC. We checked the data using Modscan and found that out of 99 polls, we are only receiving 39 responses. The data keeps breaking continuously. which indicates low availability on the RSPCB server. This issue is from the hardware side. So we are closing this complaint. | This issue is from the hardware side. So we are closing this complaint. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24093271767 | 2024-09-25 | 2024-09-28 | Cust 559 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | PTZ Camera not live on CPCB Server. | As of now closing the complaint once issue resolved from Hardware end will reopen the complaint. | As of now closing the complaint once issue resolved from Hardware end will reopen the complaint. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24093311768 | 2024-09-25 | 2024-09-27 | Cust 831 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB + RSPCB data upload issue. | We took access of the black box and observed that the internet is not coming till the black box Now the same issue is resolved from customer side and data is live on CPCB + RSPCB server. | We took access of the black box and observed that the internet is not coming till the black box. Now the same issue is resolved from customer side and data is live on CPCB + RSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24093991769 | 2024-09-26 | 2024-09-30 | Cust 892 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | To check DO analyzers Sr. nos. 22SDG019, 020 and Flow meters Sr. nos. I22431777, 78 | DO analyzers and flow meters have been checked and working fine as per attached MOM | Working fine | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL24094361770 | 2024-09-27 | 2024-09-28 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | TOC ANALYSER | FM 101 TOC Analyser(FM 101 TOC) | Visit for the TOC Analyzer inspection for review. | Inspection completed against health mapping of TOC Analyzer. Details are in attached MOM. | Inspection completed against health mapping of TOC Analyzer. Details are in attached MOM. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24094411771 | 2024-09-27 | 2024-09-28 | Cust 844 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | DATCON | DATCON(DATCON) | Data is not uploading to HSPCB server. | We have upgraded the version of the DATCon and the data is now live on the CPCB and HSPCB server. | We have upgraded the version of the DATCon and the data is now live on the CPCB and HSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24094421772 | 2024-09-27 | 2024-09-28 | Cust 931 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is uploading to RSPCB server. | Issue has been resolved and the data is now live on the RSPCB server. | Issue has been resolved and the data is now live on the RSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24095491773 | 2024-09-29 | 2024-10-02 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Schedule visit for ETP#3 Analyzers as COD, Phenol & Sulphide Analyzers against AMC. | Visit for COD, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | Visit for COD, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | EXTENDED | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24095511774 | 2024-09-29 | 2024-09-30 | Cust 711 | DELHI | Analytical - CHAKAN | CLOSED | SALES PROMOTION | TOC ANALYSER | TOC 1950 Single StreamANALYSER ANATEL(1950 TOC) | For sales promotion with Deepak Sharma | Sales promotion visit for TOC analyzer | Work done | NO | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24095521775 | 2024-09-29 | 2024-09-30 | Cust 730 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 2 AMC | Work done | Work done | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24095571776 | 2024-09-29 | 2024-09-30 | Cust 924 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | VISI pro DO(VISI pro DO) | Commissioning | Work done | Work done | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24095581777 | 2024-09-29 | 2024-09-30 | Cust 895 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Chlorine Analyzer-MAC 100 | MAC 100(1) | Commissioning | Work done | Work done | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24096251778 | 2024-09-30 | 2024-10-03 | Cust 841 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB AND HSPCB DATA UPLOAD ISSUE. | Commercially not clear | Commercially not clear | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24097661779 | 2024-10-01 | 2024-10-04 | Cust 559 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Installation and Reconfiguration of software and camera configuration on CPCB and RSPCB Server. | We have reinstalled the software and completed the local+rspcb server configuration but not getting data on modscan. As of now we are closing this complaint and will reopen it once the issue is resolved from Hardware side. | We have reinstalled the software and completed the local+rspcb server configuration but not getting data on modscan. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24098371780 | 2024-10-03 | 2024-10-04 | Cust 546 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB Server. | As of now we are closing this complaint and will reopen it once the issue is resolved from the Hardware side. | As of now we are closing this complaint and will reopen it once the issue is resolved from the Hardware side. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24098431781 | 2024-10-03 | 2024-10-04 | Cust 730 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | DAS version needs to be update to the latest version of i.e. 4.6.0.0 | we have now updated the EnviroConnect software DAS version to the latest version i.e. 4.6.0.0 and the data is live on the server. | we have now updated the EnviroConnect software DAS version to the latest version i.e. 4.6.0.0 and the data is live on the server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24098561782 | 2024-10-03 | 2024-10-06 | Cust 850 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Visit for troubling shooting in Multi-Parameter Analyzer. | Visit for online Multi-Parameter Analyzer has been completed, observed Analyzer is working satisfactorily. Details are in attached MOM. | Visit for online Multi-Parameter Analyzer has been completed, observed Analyzer is working satisfactorily. Details are in attached MOM. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL24098671783 | 2024-10-04 | 2024-10-07 | Cust 806 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on CPCB server. | The customer hasn't provided access to the system, so we are closing this complaint. We will reopen it once we receive access. | The customer hasn't provided access to the system, so we are closing this complaint. We will reopen it once we receive access. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24098711784 | 2024-10-04 | 2024-10-07 | Cust 831 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB + RSPCB data upload issue. | we took access of the black box and observed that the internet is not coming till the black box. Now the same issue is resolved from your side and data is live on CPCB + RSPCB server. | we took access of the black box and observed that the internet is not coming till the black box. Now the same issue is resolved from your side and data is live on CPCB + RSPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24099561785 | 2024-10-07 | 2024-10-14 | Cust 944 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Conductivity Tranmitter Smartpro | Display of Smartpro(1) | one of the conductivity analyzer display is not working properly in HPCL WTP site, | Display faded issue and need to replace as it is out of warranty case, Communicated to customer. | Display faded issue and need to replace as it is out of warranty case, Communicated to customer. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFDL24099741786 | 2024-10-07 | 2024-10-10 | Cust 546 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB Server. | We took access to the system and now data is live on CPCB Server. | We took access to the system and now data is live on CPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24099761787 | 2024-10-07 | 2024-10-09 | Cust 945 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | customer wants service support on urgent basiS | visit done as per attached mom. | visit done as per attached mom. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24099861788 | 2024-10-07 | 2024-10-09 | Cust 810 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | VISIT FOR CALIBRATION OF MPA | visit done .MOM attached. | visit done as per attached mom. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL24100181789 | 2024-10-08 | 2024-10-10 | Cust 864 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Chlorine sensor | CHLORINE SENSOR(1) | Visit for Chlorine Analyzer in Moon Beverages Limited, Dhaulana Pilkhuwa road, Dhaulana | Visit for Chlorine Analyzers has been completed, observed Now Analyzer is working satisfactorily. Details are in attached MOM. | Visit for Chlorine Analyzers has been completed, observed Now Analyzer is working satisfactorily. Details are in attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24100461790 | 2024-10-08 | 2024-10-11 | Cust 806 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not getting uploaded to CPCB and HSPCB servers. | As of now, we are closing this complaint. We will reopen this once the issue is resolved from the hardware side. | As of now, we are closing this complaint. We will reopen this once the issue is resolved from the hardware side. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24100571791 | 2024-10-08 | 2024-10-10 | Cust 934 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Visit for Multi-Parameter Analyzer at Bahadurgarh unit. | Visit for online Multi-Parameter Analyzer has been completed against chargeable basis. observed Now Analyzer is working satisfactorily. Details are in attached MOM. | Visit for online Multi-Parameter Analyzer has been completed against chargeable basis. observed Now Analyzer is working satisfactorily. Details are in attached MOM. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL24100751792 | 2024-10-09 | 2024-12-26 | Cust 831 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB + RSPCB data upload issue. | Here is the issue is from RSPCB server side, now the same issue is resolved and data is live on RSPCB server and data availability is above 90%. | Here is the issue is from RSPCB server side, now the same issue is resolved and data is live on RSPCB server and data availability is above 90%. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24100981793 | 2024-10-09 | 2024-10-11 | Cust 559 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on RSPCB Server. | As of now, we are closing this complaint. We will reopen it once we get access to your system. | As of now, we are closing this complaint. We will reopen it once we get access to your system. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24101071794 | 2024-10-09 | 2024-10-14 | Cust 799 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Checking & Calibration of MPA | Done the calibration of MPA | Done the calibration of Multiparameter. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24101071794 | 2024-10-09 | 2024-10-14 | Cust 799 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Checking & Calibration of MPA | Done the calibration of MPA | Done the calibration of Multiparameter. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24101151795 | 2024-10-09 | 2024-10-12 | Cust 608 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Scheduled visit for ETP online Analyzer against AMC. | Visit for online Multi-Parameter Analyzer has been completed against AMC, observed Analyzer is working satisfactorily. Details are in attached MOM. | Visit for online Multi-Parameter Analyzer has been completed against AMC, observed Analyzer is working satisfactorily. Details are in attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24101231796 | 2024-10-10 | 2024-10-14 | Cust 803 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 2nd AMC visit | 2nd AMC visit completed | Working fine | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24101341797 | 2024-10-10 | 2024-10-12 | Cust 848 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Scheduled visit for ETP online Analyzer against AMC. | Visit for online Multi-Parameter Analyzer has been completed against AMC, observed Analyzer is working satisfactorily. Details are in attached MOM. | Visit for online Multi-Parameter Analyzer has been completed against AMC, observed Analyzer is working satisfactorily. Details are in attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24101571798 | 2024-10-10 | 2024-10-12 | Cust 848 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Scheduled visit for ETP online Analyzer against AMC. | Visit for online Multi-Parameter Analyzer has been completed against AMC, observed Analyzer is working satisfactorily. Details are in attached MOM. | Visit for online Multi-Parameter Analyzer has been completed against AMC, observed Analyzer is working satisfactorily. Details are in attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24102211799 | 2024-10-12 | 2024-10-18 | Cust 919 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | visit for commissioning of WQ analyzers. | visit done for PH,Conductivity,ORP,Turbidity& chlorine analyzers commissioning. | visit done for commissioning of PH,Conductivity,ORP,Turbidity& chlorine analyzers | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24102451800 | 2024-10-13 | 2024-10-13 | Cust 854 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | pH input card-FM of Multiparameter(1) | MPA pH and Modbas card replacement. | MPA COD,BOD,pH Parameters are correct but TSS touch is not working due to which its reading is showing wrong | MPA COD,BOD,pH Parameters are correct but TSS touch is not working due to which its reading is showing wrong. | YES | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFDL24102811801 | 2024-10-14 | 2024-10-17 | Cust 559 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on CPCB server and RSPCB server. | As of now closing the complaint as static IP shared by you is not playing out of your network. | As of now closing the complaint as static IP shared by you is not playing out of your network. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24103281802 | 2024-10-14 | 2024-10-17 | Cust 796 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Aqua4Trans | Aqua4Trans(1) | Calibration of pH/TSS Analyzer and SPM | Calibration service visit completed | Working fine | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL24103561803 | 2024-10-15 | 2024-10-17 | Cust 546 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on RSPCB server. | We took access to the system and now data is live on RSPCB server. | We took access to the system and now data is live on RSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24103571804 | 2024-10-15 | 2024-10-16 | Cust 841 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on CPCB portal | The data for the ETP outlet on the FM CPCB server was showing a future date, 07-11-2085. We raised this issue with the AIPL team, and it has now been resolved. The data is currently live on the CPCB server. | The data is live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24103591805 | 2024-10-15 | 2024-10-17 | Cust 571 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Display of Smartpro(1) | Display issue | Problem not solved after changing the display with new display. suggest customer to send transmitter to pune . | closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24103701806 | 2024-10-15 | 2024-10-30 | Cust 790 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to RSPCB server. | Now the data is live on the RSPCB server. | Now the data is live on the RSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24103721807 | 2024-10-15 | 2024-10-15 | Cust 841 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on CPCB + Local + HSPCB. | We took access to the system and now data is live on CPCB + Local + HSPCB servers. for Stack 1_18 TPH Boiler. | We took access to the system and now data is live on CPCB + Local + HSPCB servers. for Stack 1_18 TPH Boiler. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24103811808 | 2024-10-15 | 2024-11-04 | Cust 854 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | PH Input Card for Multiparameter(AWA-PH) | All the parameters of Multipara meter analyzer was showing wrong. | Online support Provided. Failure spares concluded & since spares cost was high suggested buyback & to go for CX2000 | Online support Provided. Failure spares concluded &amp;amp; since spares cost was high suggested buyback &amp;amp; to go for CX2000 | YES | No Trip | Process | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24103821809 | 2024-10-15 | 2024-10-16 | Cust 891 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on CPCB + Local + HSPCB servers. | We took access to the system and observed that the data is not communicating till the Local system.Hence this issue is from the hardware side.so we are closing this complaint. | We took access to the system and observed that the data is not communicating till the Local system.Hence this issue is from the hardware side.so we are closing this complaint. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24103961810 | 2024-10-16 | 2024-10-18 | Cust 533 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB AND RSPCB DATA UPLOADED ISSUE. | now the issue has been resolved and the data is live on the server. | now the issue has been resolved and the data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24104051811 | 2024-10-16 | 2024-11-18 | Cust 816 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue. | This issue is related to the RSPCB and we have already raised this issue with them. Therefore, we are closing this complaint now. | This issue is related to the RSPCB and we have already raised this issue with them. Therefore, we are closing this complaint now. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24104091812 | 2024-10-16 | 2024-10-16 | Cust 839 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on RSPCB server. | Now data is live on the RSPCB server. | Now data is live on the RSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24104491813 | 2024-10-17 | 2024-10-21 | Cust 940 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data upload issue. | we took access of the black box and made some modifications, now data is live on the CPCB server. | we took access of the black box and made some modifications, now data is live on the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24106081814 | 2024-10-20 | 2024-12-23 | Cust 919 | DELHI | Analytical - CHAKAN | CLOSED | ENGINEERING ISSUE | Chlorine sensor | CHLORINE SENSOR(1) | Chlorine Sensor FMSNX-CLD502 product manual ( Documents) issue | Necessary Corrections Done in the documentation as confirmed by engineering | Necessary Corrections Done in the doccumentaion as confirmed by engineering | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFDL24106321815 | 2024-10-21 | 2024-10-24 | Cust 559 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Camera not playing on RSPCB server. | We created the rtsp link and shared with the customer. | We created the rtsp link and shared with the customer. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24106881816 | 2024-10-21 | 2024-10-25 | Cust 786 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Schedule visit for ETP online Analyzer against CAMC. | During visit, observed PH1&PH2 signals are fluctuating, waiting for Detectors, once received from HO, again visit to be done. | During visit, observed PH1&PH2 signals are fluctuating, waiting for Detectors, once received from HO, again visit to be done. | EXTENDED | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24107001817 | 2024-10-22 | 2024-10-25 | Cust 819 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Our Boiler has to be put in operation due to start of Plant. Please start UPPCB and CPCB Server data transfer. | commercially not clear | commercially not clear | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24107301818 | 2024-10-22 | 2024-10-25 | Cust 786 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Schedule visit for ETP online Analyzer against CAMC. | Visit for online Multi-Parameter Analyzer has been completed, observed Analyzer is working satisfactorily. Details are in attached MOM. | Visit for online Multi-Parameter Analyzer has been completed, observed Analyzer is working satisfactorily. Details are in attached MOM. | EXTENDED | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24107331819 | 2024-10-22 | 2024-10-24 | Cust 807 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Visit for checking of MPA | visit done under DEUS | visit done & MOM attached . | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24107471820 | 2024-10-23 | 2024-10-27 | Cust 532 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | To repair and calibrate analyzer with recently purchased spares | Resolved the issues after replacing the spares as per the attached MOM. | Working fine | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24107541821 | 2024-10-23 | 2024-10-24 | Cust 801 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Data not updating on CPCB server | We took access to the system and found data was live on LOCAL+CPCB server. | We took access to the system and found data was live on LOCAL+CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24108041822 | 2024-10-24 | 2024-10-25 | Cust 464 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB server. | Server is accessible and data is live on the CPCB server. | Server is accessible and data is live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24108431823 | 2024-10-25 | 2024-10-28 | Cust 790 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to Local and CPCB server. | data is live on the CPCB server. | data is live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24108551824 | 2024-10-25 | 2024-10-25 | Cust 844 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | DATCON | DATCON(DATCON) | HSPCB data upload issue | the data is live on the HSPCB server. | the data is live on the HSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24108661825 | 2024-10-25 | 2024-10-28 | Cust 723 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation | We have installed the enviroconnect software of the latest version and completed Local and CPCB configuration. Data is now live on the Local and CPCB servers. | We have installed the enviroconnect software of the latest version and completed Local and CPCB configuration. Data is now live on the Local and CPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24108761826 | 2024-10-25 | 2024-10-28 | Cust 946 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera streaming issue. | We have checked and camera streaming is live on the CCPB server. | We have checked and camera streaming is live on the CCPB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24109021827 | 2024-10-26 | 2024-11-01 | Cust 786 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | Peristaltic sampling pump motor for Multiparameter(AWA-PSPM) | Sample pump is not running in automatically during sample interval time. | The sample pump auto-run issue on the Multi-Parameter Analyzer was resolved by our team member, Suneet Kumar, through telephonic support. The online analyzer is now functioning perfectly, and we are closing this complaint. | Sample pump auto-run issue of Multi-Parameter Analyzer was resolved by our Suneet Kumar, through telephonic support. Now analyzer is functioning perfectly, and we are closing this complaint. | EXTENDED | No Trip | Core | N |  | Analytical - CHAKAN
AFDL24109831828 | 2024-10-28 | 2024-10-28 | Cust 816 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | To change the permissible range on CPCB server. | We have now also updated the parameter limit for Stack 1_Boiler_8MW station on our server. | We have now also updated the parameter limit for Stack 1_Boiler_8MW station on our server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24110001829 | 2024-10-28 | 2024-10-30 | Cust 841 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on RTDMS portal. | The current data for the ETP outlet is now being updated live on the FM CPCB server. | The current data for the ETP outlet is now being updated live on the FM CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24110071830 | 2024-10-28 | 2024-10-30 | Cust 816 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on CPCB server from October 28, 2024, at 06:12:41. | We haven't received access to the system from the customer, so we are closing this complaint | We haven't received access to the system from the customer, so we are closing this complaint | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24110081831 | 2024-10-28 | 2024-10-30 | Cust 923 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on CPCB server. | We have not received the AMC purchase order details for this year. Therefore, we are closing this complaint. | We have not received the AMC purchase order details for this year. Therefore, we are closing this complaint. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24110451832 | 2024-10-29 | 2024-10-30 | Cust 803 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera status showing inactive on RSPCB. | We have checked that RTSP link is not working out of network . | We have checked that RTSP link is not working out of network . | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24110551833 | 2024-10-29 | 2024-11-22 | Cust 885 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Checking and calibration of MPA | Checked the MPA | Checked the MPA Analyzer | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL24110631834 | 2024-10-29 | 2024-11-22 | Cust 858 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Checking of MPA | Checked the MPA and replaced the pH PCB | Checked the MPA pH card. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24110801835 | 2024-10-29 | 2024-11-04 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | AMC | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | Visit of Renewal Analyzers for submission of Contract Agreement against AMC. | I visited the site to submit the Contract Agreement for the AMC renewal. However, the submission could not be completed due to the unavailability of the signing authority approval letter. I have informed our Head Office to prepare this approval letter. | Site visited done to submit the Contract Agreement. But it could not be completed due to the unavailability of signing authority letter. I have informed our Head Office to prepare this approval lette | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24110991836 | 2024-10-30 | 2024-10-30 | Cust 839 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | RSPCB data uploading issue. | data is currently live on the RSPCB server. | data is currently live on the RSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24112701837 | 2024-11-04 | 2024-11-07 | Cust 786 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | Flowcell disc of Multiparameter(1) | Visit for spares replacement in online Analyzer against CAMC. | Visit for online Multi-Parameter Analyzer has been completed, observed Analyzer is working satisfactorily. Details are in attached MOM. | Visit for online Multi-Parameter Analyzer has been completed, observed Analyzer is working satisfactorily. Details are in attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24113551838 | 2024-11-06 | 2024-11-28 | Cust 790 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation | Software installation completed and data is live on the Local, CPCB servers. We are facing common issue for RSPCB and taken up this issue with RSPCB team hence as of now closing this complaint. | Software installation completed and data is live on the Local, CPCB servers. We are facing common issue for RSPCB and taken up this issue with RSPCB team hence as of now closing this complaint. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24113581839 | 2024-11-06 | 2024-11-07 | Cust 849 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | Peristaltic sampling pump motor for Multiparameter(AWA-PSPM) | Visit for sample motor replacement in Multi-Parameter Analyzer. | Replaced sample pump motor in RO Multi-Parameter Analyzer, Calibration done, Now the Analyzer is working satisfactorily. Details are in attached MOM. | Replaced sample pump motor in RO Multi-Parameter Analyzer, Calibration done, Now the Analyzer is working satisfactorily. Details are in attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24113751840 | 2024-11-06 | 2024-11-10 | Cust 790 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 1st AMC visit | 1st AMC visit completed | Working fine | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24113961841 | 2024-11-07 | 2024-11-07 | Cust 900 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera streaming issue on CPCB server. | We have accessed the system, and the camera streaming is now live on the CPCB server. | We have accessed the system, and the camera streaming is now live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24114211842 | 2024-11-07 | 2024-11-08 | Cust 885 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | DATCON | DATCON(DATCON) | HSPCB DATA UPLODE ISSUE. | The issue has been resolved and the data is live on the HSPCB server | The issue has been resolved and the data is live on the HSPCB server | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24114501843 | 2024-11-08 | 2024-11-18 | Cust 874 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Customer repoerted for calibartion of MPA | Calibration Done | Calibration done. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24114551844 | 2024-11-08 | 2024-11-09 | Cust 810 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration visit | Completed | Working fine | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL24116191845 | 2024-11-12 | 2024-11-15 | Cust 823 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | PM online data monitoring system site error issue for CPCB, RSPCB server. | Commercially not cleared | Commercially not cleared | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24116781846 | 2024-11-12 | 2024-11-22 | Cust 868 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Visit for calibration of MPA | visit done under MPA checking | visit done as per attached mom. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL24117211847 | 2024-11-13 | 2024-11-15 | Cust 803 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera streaming issue. | We have created RTSP link which is playing outside of customers network. so issue has been resolved | We have created RTSP link which is playing outside of customers network. so issue has been resolved | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24117221848 | 2024-11-13 | 2024-11-15 | Cust 804 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data uploading issue. | Data is now live on the Local, CPCB and HSPCB servers. | Data is now live on the Local, CPCB and HSPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24117841849 | 2024-11-14 | 2024-11-15 | Cust 891 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data mismatch issue. | we took access to your system and now the hardware side issue is resolved and Boiler Stack data is matching with actual readings and the same is verified by you also. | we took access to your system and now the hardware side issue is resolved and Boiler Stack data is matching with actual readings and the same is verified by you also. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24119121850 | 2024-11-18 | 2024-11-21 | Cust 858 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | DATCON | DATCON(DATCON) | HSPCB data upload issue. | We have checked and data is already live on the server. | We have checked and data is already live on the server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24119281851 | 2024-11-18 | 2024-11-21 | Cust 790 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera streaming issue | Camera streaming is live on the CPCB server. | Camera streaming is live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24119631852 | 2024-11-18 | 2024-11-23 | Cust 935 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning visit for Multi-Parameter Analyzer. | Online Multi-Parameter & TSS Analyzers have been successfully commissioned, observed both the Analyzers are working satisfactorily and data's are transferring on cloud server through Black Box. Details are in the attached MOM. | Multi-Parameter & TSS Analyzers have been commissioned, observed both Analyzers are working satisfactorily and data's are transferring on cloud server through BlackBox. Details are in the MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24119901853 | 2024-11-19 | 2024-11-25 | Cust 823 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | PM data not updating on CPCB and RSPCB Server. | We made some some changes from hardware side and now data is live on CPCB Server for STACK 1 and for stack 2 data not continuously updating on modscan need to check from hardware end. | We made some some changes from hardware side and now data is live on CPCB Server for STACK 1 and for stack 2 data not continuously updating on modscan need to check from hardware end. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24119951854 | 2024-11-19 | 2024-11-22 | Cust 947 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera streaming issue | Camera streaming is now live on the CPCB server. | Camera streaming is now live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24119991855 | 2024-11-19 | 2024-11-22 | Cust 802 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | HSPCB data uploaded issue. | Now the issue has been resolved and the data is live on the HSPCB server. | Now the issue has been resolved and the data is live on the HSPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24120871856 | 2024-11-21 | 2024-11-28 | Cust 874 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Customer reported for calibration job | calibration done | Calibration done. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24121081857 | 2024-11-21 | 2024-11-22 | Cust 870 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | RTDAS upgradation needs be done with the latest version. | We have updated the RTDAS version to the Latest version _4.6 and the data is live on the CPCB server. | We have updated the RTDAS version to the Latest version _4.6 and the data is live on the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24121111858 | 2024-11-21 | 2024-11-22 | Cust 809 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | RSPCB data uploading issue | Data is live on the Local and CPCB server. For RSPCB, there is issue from RSPCB server end hence as of now we are closing this complaint. | Data is live on the Local and CPCB server. For RSPCB, there is issue from RSPCB server end hence as of now we are closing this complaint. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24121721859 | 2024-11-22 | 2024-11-27 | Cust 870 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | visit for calibration of MPA | visit done for MPA as per attached MOM . | visit done for MPA as per attached MOM . | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL24121771860 | 2024-11-23 | 2024-11-30 | Cust 895 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Turbidity Meter | DIS/Q(182) | FMTURB analyzer, Chlorine analyzer FM5296H, Aquamon2000pro pH for pabra | MOM attached | MOM attached. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24121781861 | 2024-11-23 | 2024-11-30 | Cust 895 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Turbidity Meter | DIS/Q(182) | FMTURB analyzer, Chlorine analyzer FM5296H, Aquamon2000pro pH for mandouli & Nanowal | MOM attached | MOM attached | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24121791862 | 2024-11-23 | 2024-11-30 | Cust 786 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | commissioning CX-2 | MOM attached | MOM attached | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24122201863 | 2024-11-23 | 2024-11-23 | Cust 938 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | AQUAMAX DOT(1) | DO Meter is not matching lab value. | We will supply new DO set to cross verify the working of present DO meter | We will supply new DO set to cross verify the working of present DO meter | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFDL24122211864 | 2024-11-23 | 2024-11-23 | Cust 938 | DELHI | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Pro-Active visit. | MPA is working fine. | We will supply new DO set to cross verify the working of present DO meter | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFDL24122651865 | 2024-11-25 | 2024-11-26 | Cust 806 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue | We have accessed the system and noticed that the data is not communicating with the local system for CETP Outlet (Checked on MODSCAN)..Additionally, for the CETP Outlet_Flow, it is showing a "connection failed" error.So this issue is from the hardware side.So we are closing this complaint. | This issue is from the hardware side.So we are closing this complaint. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24123631866 | 2024-11-26 | 2024-11-28 | Cust 809 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 2nd AMC visit for Multiparameter and Oil Analyzers | 2nd AMC visit completed | Working fine | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24123901867 | 2024-11-27 | 2024-11-29 | Cust 946 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera intigartion on Kumbh portal/. | Here CPU PLANT Camera streaming is now working hence we are closing this complaint and will reopen it again once we receive any updates from coustomer end. | Here CPU PLANT Camera streaming is now working hence we are closing this complaint and will reopen it again once we receive any updates from coustomer end. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24123961868 | 2024-11-27 | 2024-12-02 | Cust 868 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | visit under troubleshooting | Visit made and found xenon power card and disk faulty hence same will be arranged by client. | visit done .MOM attached . | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24124101869 | 2024-11-27 | 2024-11-29 | Cust 869 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Reconfiguration and installation on software. | We completed the installation and configuration and now data is live on CPCB and UPPCB server. For localhost configuration PC name is invalid asked customer to update old PC name. | We completed the installation and configuration and now data is live on CPCB and UPPCB server. For localhost configuration PC name is invalid asked customer to update old PC name. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24124471870 | 2024-11-27 | 2024-12-02 | Cust 812 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Checking and calibration | Done the calibration of MPA | Done the calibration of Multiparameter. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL24124471870 | 2024-11-27 | 2024-12-02 | Cust 812 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Checking and calibration | Done the calibration of MPA | Done the calibration of Multiparameter. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL24124921871 | 2024-11-28 | 2024-12-02 | Cust 869 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Visit for installation of supplied spares | Visit made mom attached | visit done .MOM attached . | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL24125331872 | 2024-11-28 | 2024-12-05 | Cust 872 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC-4th | Work done | Work done | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24125341873 | 2024-11-28 | 2024-12-05 | Cust 709 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC-4th | Work done | Work done | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24125351874 | 2024-11-28 | 2024-12-05 | Cust 880 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration for dhali MPA | Work done | Work done | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL24126061875 | 2024-11-30 | 2024-12-01 | Cust 564 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | AQUAMAX DOT(1) | Commissioning work of DO analyzers and FRC analyzers at ETP site | Commissioning work completed | Working fine | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24127711876 | 2024-12-02 | 2024-12-24 | Cust 874 | DELHI | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | MULTIPARAMETER ANALYSER | PH Sensor for Multiparameter(AWA-PH-SENS) | Customer reported for fluctuation in PH readings after commissioning. | Replacment material is sent via MARK EXPRESS PRIVATE LIMITED vide LR No - 210124931 LR Dt - 11-DEC-2024 | Replacment material is sent via MARK EXPRESS PRIVATE LIMITED vide LR No - 210124931 LR Dt - 11-DEC-2024 | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFDL24128461877 | 2024-12-02 | 2024-12-05 | Cust 843 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 3rd AMC visit | 3rd AMC visit completed | Working fine | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24128501878 | 2024-12-02 | 2024-12-05 | Cust 869 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | LOCAL configuration need to be done. | We took access to the system and completed the LOCAL configuration and now data is live on LOCAL+CPCB+KSPCB server. | We took access to the system and completed the LOCAL configuration and now data is live on LOCAL+CPCB+KSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24128751879 | 2024-12-03 | 2024-12-06 | Cust 799 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB AND HSPCB DATA UPLODED ISSUE. | Here issue was from the server HSPCB side and same has been resolved now | Here issue was from the server HSPCB side and same has been resolved now | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24129561880 | 2024-12-04 | 2024-12-08 | Cust 922 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Visit for calibration | visit done as per attached mom.MPA is working fine. | visit done as per attached mom.MPA is working fine. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL24130571881 | 2024-12-06 | 2024-12-07 | Cust 802 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB + HSPCB data uploded issue. | Here issue was with the hardware side as the USB port was not detecting and now the same issue has been resolved and the data is live on the server. | Here issue was with the hardware side as the USB port was not detecting and now the same issue has been resolved and the data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24130601882 | 2024-12-06 | 2024-12-08 | Cust 814 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Online support for MPA | online support provided to client on 06/12/2024 & then MPA is working fine. | online support provided to client on 06/12/2024 &amp;amp; then MPA is working fine. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFDL24131671883 | 2024-12-09 | 2024-12-12 | Cust 532 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB and RSPCB server | We took access to the system and data was live on CPCB+LOCAL Server. For RSPCB it is common issue we have taken up the issue with RSPCB Team. | We took access to the system and data was live on CPCB+LOCAL Server. For RSPCB it is common issue we have taken up the issue with RSPCB Team. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24131801884 | 2024-12-09 | 2024-12-12 | Cust 800 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB Server. | We took access to the system and found that RTDAS services were stopped. Data is live now on LOCAL+CPCB server | We took access to the system and found that RTDAS services were stopped. Data is live now on LOCAL+CPCB server | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24131981885 | 2024-12-09 | 2024-12-13 | Cust 871 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Visit for checking and calibration of MPA | visit done .MOM attached . | visit done .MOM attached . | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL24131991886 | 2024-12-09 | 2024-12-21 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | Visit for checking and calibration of 24 analyzers under ETP | visit done for DEC 2024 under ETP analyzers ARC | visit done as per attached mom. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL24132211887 | 2024-12-10 | 2024-12-17 | Cust 837 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Checking of MPA Analyser. | XENON PSU not working ,customer will purchase the same. | checked the PO | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24133241888 | 2024-12-12 | 2024-12-14 | Cust 871 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data uploaded issue. | As of now we are closing this complaint and will reopen it again once we get remort access of the system from the customer end. | As of now we are closing this complaint and will reopen it again once we get remort access of the system from the customer end. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24133251889 | 2024-12-12 | 2025-03-17 | Cust 787 | DELHI | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | DATCON | DATCON(DATCON) | CPCB data uploaded issue. | Material sent as per RF OA no:5525800815. | Material Sent against RF OA no:5525800815 | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24133651890 | 2024-12-13 | 2024-12-14 | Cust 533 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | Xenon Lamp Power supply for Multiparameter(AWA-XLAMPPS) | 2nd AMC visit | 2nd AMC visit completed | Working fine except Xenon supply and offered the same | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24133891891 | 2024-12-14 | 2024-12-17 | Cust 532 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | Flowcell disc of Multiparameter(1) | To replace disc assembly of flow cell and calibrate the analyzer | Disc assembly rand inlet sample tube and strainer replaced, analyzer is working fine. | Working fine | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24134281892 | 2024-12-15 | 2024-12-21 | Cust 948 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning the Multi-Para, Phosphorus, Ammonia & Chlorine Analyzer. | Visit to commissioning of Multi-Parameter, Oil in Water, Ammonia, Phosphorus and Chlorine Analyzer is under progress. Customer will confirm availability of standard solution for next visit. | Visit to commissioning of Multi-Parameter, Oil in Water, Ammonia, Phosphorus and Chlorine Analyzer is under progress. Customer will confirm availability of standard solution for next visit. | NO | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24134851893 | 2024-12-16 | 2024-12-19 | Cust 559 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on CPCB server. | We took access to the system and found streaming window was closed and for rspcb we created link and share with customer. | We took access to the system and found streaming window was closed and for rspcb we created link and share with customer. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24135251894 | 2024-12-17 | 2024-12-27 | Cust 802 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Checking and calibration of Analyzer | AMC visit completed | Done the calibration of Multiparameter. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24136841895 | 2024-12-20 | 2024-12-27 | Cust 850 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Checking and calibration of MPA | Checked The MPA | Done the calibration of Multiparameter. | NO | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24136851896 | 2024-12-20 | 2024-12-21 | Cust 948 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning visit for Multi-Parameter, Oil in water, Ammonia, Phosphorus & Chlorine Analyzers. | All the supplied analyzers have been successfully commissioned, along with site calibration using standard solutions. Currently, all the supplied analyzers are functioning satisfactorily, except for the Phosphorus Analyzer. | All supplied analyzers have been commissioned, alongwith site calibration with standard solutions. Now the supplied analyzers are functioning satisfactorily, except for the Phosphorus Analyzer. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24138371897 | 2024-12-23 | 2024-12-26 | Cust 891 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | HSPCB data upload issue. | today we took access to your system and reinstalled Enviroconnect software related files, now the local host server is accessible and data is live on the HSPCB server. | today we took access to your system and reinstalled Enviroconnect software related files, now the local host server is accessible and data is live on the HSPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24140291898 | 2024-12-27 | 2024-12-28 | Cust 929 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue. | We accessed the system and checked the data on ModScan, but it is showing a "connection failed" message. Additionally, the IP address provided for ETP-1 (192.168.1.99) is not responding to ping requests from the local PC. So this issue is from the customer side.so we are closing this complaint for now and we wlll reopen it once the IP address for ETP-1 (192.168.1.99) starts responding consistently. | We accessed the system and checked the data on ModScan, but it is showing a &amp;quot;connection failed&amp;quot; message. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24140781899 | 2024-12-28 | 2024-12-30 | Cust 946 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera Streaming issue on the CPCB server. | Here was with the customer network end and the same has been resolved data the Camera streaming is live now. | Here was with the customer network end and the same has been resolved data the Camera streaming is live now. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24140821900 | 2024-12-28 | 2024-12-28 | Cust 827 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Line No. 5 (Grain Distillation Feed Inlet Evaporator_MassFlow) and Line No. 9 (Weak Spent wash dual M), Flow transmitters reading is mismatching in Field & Online site. | Here issue was from scanner side and same has been resolved now. | Here issue was from scanner side and same has been resolved now. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24140961901 | 2024-12-28 | 2024-12-31 | Cust 949 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | AQUAMAX DOT(1) | COMMISSIONING | Commissioning has been done for aquamax DO. 1 SMARTPro DO found not display turning up. need to sending pune | Commissioning has been done. 1 SMARTPro DO found not display turning up. need to sending pune | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24142241902 | 2024-12-30 | 2025-01-08 | Cust 950 | DELHI | Steam System Division | CLOSED | PRODUCT FAILURE | BBCS | BBCS - BC3210 - CONTROLLER(BC3210) | BBCS modbus output issue | Work done | Work done | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFDL24142271903 | 2024-12-30 | 2024-12-31 | Cust 610 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Electrode Cable(Ph-Cab) | Commissioning 1712 | work done | Work done | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24142451904 | 2024-12-30 | 2024-12-31 | Cust 951 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | VISI pro DO(VISI pro DO) | Commissioning for 2 DO sensor visiwater pro 1112 | Commissioning done | Commissioning has been done. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24142471905 | 2024-12-30 | 2024-12-31 | Cust 708 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Pre-commissioning 1411 | Pre-commissioning guideline done | Pre-commissioning guideline done | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24142521906 | 2024-12-30 | 2024-12-31 | Cust 951 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | VISI pro DO(VISI pro DO) | commissioning 0910 | Wrong material supply hence commissioning not done | Wrong material supply hence commissioning not done | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24143581907 | 2024-12-31 | 2025-02-28 | Cust 952 | DELHI | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | Turbidity Meter | DIS/Q(182) | STP Turbidity meter half display is not working at the Shri Ram mandir site | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210128144 LR Dt - 11-JAN-2025 | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210128144 LR Dt - 11-JAN-2025 | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFDL24144991908 | 2024-12-31 | 2024-12-31 | Cust 901 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | commissioning at dehra WTP | commissioning has been done | Commissioning has been done. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24145041909 | 2024-12-31 | 2025-01-01 | Cust 895 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Chlorine Analyzer-MAC 100 | MAC 100(1) | commissionig of chlorine and turbidity at pabra WTP | commissionig done with santoshsir | Commissioning has been done. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24145491910 | 2025-01-02 | 2025-01-09 | Cust 862 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Customer reported for Calibration of MPA | Calibration done. | Calibration done. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24145821911 | 2025-01-03 | 2025-01-06 | Cust 905 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to HSPCB server. | Now the HSPCB server is accessible and data is live on the same server. | Now the HSPCB server is accessible and data is live on the same server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24146051912 | 2025-01-03 | 2025-01-03 | Cust 546 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB+RSPCB server. | We took access to the system and now data is live on CPCB+RSPCB server. | We took access to the system and now data is live on CPCB+RSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24146681913 | 2025-01-06 | 2025-01-06 | Cust 923 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on CPCB server. | Now the data is live on the CPCB server. | Now the data is live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24146951914 | 2025-01-06 | 2025-01-10 | Cust 843 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | Xenon Lamp Power supply for Multiparameter(AWA-XLAMPPS) | To check multiparameter analyzer for high reading issue under emergency visit | Checked and found issue in xenon power supply and submitted offer of the same and internal external tubing set | Xenon supply PCB & tubing set ordered by Naraina CETP | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24147081915 | 2025-01-06 | 2025-01-10 | Cust 572 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | AQUADO Analyzer | Visiwater DO sensor(1) | DO sensor (Visiwater DO P Arc FC10, Sr No. 2827) is not working at Tehkhand Waste To Electricity Project Ltd, Maa Anandmayee Marg, Tehkhand, New Delhi. | Visit for DO Analyzer, observed the sensor was not performing. Customer will send sensor to pune for testing. Details are in attached MOM. | Visit for DO Analyzer, observed the sensor was not performing. Customer will send sensor to pune for testing. Details are in attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24147281916 | 2025-01-07 | 2025-01-10 | Cust 869 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | pH data mismatch on UPPCB server. | We have changed the configuration of UPPCB and now correct data is updating on UPPCB server. | We have changed the configuration of UPPCB and now correct data is updating on UPPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24147771917 | 2025-01-08 | 2025-01-10 | Cust 841 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Central portal Data upload issue. | We have updated the Device ID from the server end and now the Data is live. | We have updated the Device ID from the server end and now the Data is live. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24149011918 | 2025-01-10 | 2025-01-10 | Cust 827 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Grain Distillation Feed Inlet Evaporation_MassFlow? is not updating Average date from 31.12.2024 06:04:00. Also, this flow transmitter is not creating Report from 01.01.2025. | We have accessed the system and made the necessary changes from our side. Now, the data for Grain Distillation Feed Inlet Evaporator_MassFlow is live on the CPCB server. Additionally, the station reports are now being generated on the CPCB server. | . Now, the data for Grain Distillation Feed Inlet Evaporator_MassFlow is live on the CPCB server. Additionally, the station reports are now being generated on the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24149021919 | 2025-01-10 | 2025-01-11 | Cust 801 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 2nd AMC visit | 2nd AMC visit completed | Working fine | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24149091920 | 2025-01-10 | 2025-01-10 | Cust 559 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB server. | We took access to the system and checked data is not coming till MODSCAN and same informed to customer to check from hardware end. | We took access to the system and checked data is not coming till MODSCAN and same informed to customer to check from hardware end. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24149281921 | 2025-01-10 | 2025-01-13 | Cust 807 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data uploaded issue. | We have updated the BB with the new patch and now the data is live. | We have updated the BB with the new patch and now the data is live. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24149441922 | 2025-01-11 | 2025-01-12 | Cust 533 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Emergency visit to replace the xenon lamp power supply board and calibration of analyzer | Analyzer is rectified, calibrated and working fine | Working fine | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24150401923 | 2025-01-14 | 2025-01-21 | Cust 874 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Customer reported for calibration of MPA. | calibration done | Calibration done | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24151461924 | 2025-01-15 | 2025-01-21 | Cust 717 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration | Work done | Work done | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL24151481925 | 2025-01-15 | 2025-01-19 | Cust 848 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Scheduled visit for ETP online Analyzer against AMC. | AMC Visit for online Multi-Parameter Analyzer has been completed, observed Analyzer is working satisfactorily. Details are in attached MOM. | AMC Visit for online Multi-Parameter Analyzer has been completed, observed Analyzer is working satisfactorily. Details are in attached MOM. | EXTENDED | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24151491926 | 2025-01-16 | 2025-01-19 | Cust 858 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Scheduled visit for ETP online Analyzer against AMC. | AMC Visit for online Multi-Parameter Analyzer has been completed, observed Analyzer is working satisfactorily. Details are in attached MOM. | AMC Visit for online Multi-Parameter Analyzer has been completed, observed Analyzer is working satisfactorily. Details are in attached MOM. | EXTENDED | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24151951927 | 2025-01-16 | 2025-01-22 | Cust 919 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Visit for commissioning of EQMS Analyzer | visit done for the commissioning of EQMS Analyzers ( COD/BOD,Oil,pH,TSS,Conductivty) . all are working fine .mom attached . | visit done for the commissioning of EQMS Analyzers ( COD/BOD,Oil,pH,TSS,Conductivty) . all are working fine .mom attached . | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24152271928 | 2025-01-17 | 2025-01-20 | Cust 533 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data mismatch issue on CPCB server. | Now the issue has been resolved from the hardware side and the correct data is reflecting to the CPCB server. | Now the issue has been resolved from the hardware side and the correct data is reflecting to the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24153291929 | 2025-01-20 | 2025-01-23 | Cust 831 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | RSPCB data availability issue. | we took access of the black box and updated the new patch, now data is continuously live on the RSPCB server. | we took access of the black box and updated the new patch, now data is continuously live on the RSPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24153921930 | 2025-01-21 | 2025-01-25 | Cust 837 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Schedule visit for online Multi Parameter Analyzer in STP against AMC. | Visit for online Multi-Parameter Analyzer has been completed, observed Analyzer is working satisfactorily. Details are in attached MOM. | Visit for online Multi-Parameter Analyzer has been completed, observed Analyzer is working satisfactorily. Details are in attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24154071931 | 2025-01-21 | 2025-01-30 | Cust 953 | DELHI | Steam System Division | CLOSED | ENGINEERING ISSUE | TRAP - MONITORING SYSTEM | FMTMS(FMTMS) | 5 Nos. Trap Monitoring system ( FMTMS-AA ) not working. Kindly provide solution on urgent basis | material to be supplied against RF OA 5525800726, Complaint is closed for quality review, RF OA is approved | material to be supplied against RF OA 5525800726, Complaint is closed for quality review, RF OA is approved | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFDL24154211932 | 2025-01-21 | 2025-01-25 | Cust 608 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Scheduled visit for ETP online Analyzer against AMC. | Visit for online Multi-Parameter Analyzer has been completed, observed Analyzer is working satisfactorily. Details are in attached MOM. | Visit for online Multi-Parameter Analyzer has been completed, observed Analyzer is working satisfactorily. Details are in attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24154221933 | 2025-01-21 | 2025-01-25 | Cust 786 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Scheduled visit for ETP online Analyzer against AMC. | Visit for online Multi-Parameter Analyzer has been completed, observed Analyzer is working satisfactorily. Details are in attached MOM. | Visit for online Multi-Parameter Analyzer has been completed, observed Analyzer is working satisfactorily. Details are in attached MOM. | EXTENDED | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24154601934 | 2025-01-22 | 2025-01-27 | Cust 954 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | VISI pro DO(VISI pro DO) | visit for commissioning | Visit done for 10 nos DO Analyzers commissioning under blutooth communication .mom attached. | visit done as per attached mail communication. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24154701935 | 2025-01-22 | 2025-01-27 | Cust 799 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | 3-Way Electric Valve for Multiparameter(AWA-3WEV) | checking & calibration of MPA | Done the Calibration | Done the calibration of Multiparameter. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24155461936 | 2025-01-24 | 2025-01-24 | Cust 837 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | DATCON | DATCON(DATCON) | CPCB AND HSPCB DATA UPLODED ISSUE. | Here there was a hardware issue and same has been resolved and data is now live. | Here there was a hardware issue and same has been resolved and data is now live. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24155941937 | 2025-01-25 | 2025-01-25 | Cust 938 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | DO sensor needs to be replace. | DO sensor is replaced and it is under observation. | DO sensor is replaced and it is under observation. | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFDL24156001938 | 2025-01-25 | 2025-01-25 | Cust 955 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | Sample Pump(1) | visit required for troubleshooting of MPA | MPA spare are required offer is send | MPA spare are required offer is send | NO | Periodic Maintenance (AMC) | Process | N |  | Analytical - CHAKAN
AFDL24156011939 | 2025-01-25 | 2025-01-25 | Cust 600 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | Cleaning Pump(1) | MPA is not in operation from long time | MPA maintenance activities are required. | MPA maintenance activities are required. | NO | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFDL24156011939 | 2025-01-25 | 2025-01-25 | Cust 600 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | Cleaning Pump(1) | MPA is not in operation from long time | MPA maintenance activities are required. | MPA maintenance activities are required. | NO | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFDL24161411940 | 2025-01-31 | 2025-02-01 | Cust 843 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Emergency visit for fixing the problem of TSS after replacing xenon power supply card | Resolved all problems and analyzer is working fine | Working fine | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24162181941 | 2025-02-01 | 2025-02-03 | Cust 810 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | DATCON | DATCON(DATCON) | CPCB data mismatch issue for ETP. | Here customer has changed the scanner so we have made modifications as per the new scanner and data the live and uploading correctly on the CPCB server. | Here customer has changed the scanner so we have made modifications as per the new scanner and data the live and uploading correctly on the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24162211942 | 2025-02-01 | 2025-02-03 | Cust 608 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data uploaded issue. | As of now we are closing this complaint and will reopen it again once we get remote access of the BB from coustomer end. | As of now we are closing this complaint and will reopen it again once we get remote access of the BB from coustomer end. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24162401943 | 2025-02-02 | 2025-02-09 | Cust 810 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | Falcon Scanner for Multiparameter(AWA-SCAN) | online support for CPCB data connectivity under falcon mini scanner | Support done over remotly & telephonic ally as per attached mail communication. now system is working fine. | Support done over remotly &amp;amp; telephonic ally as per attached mail communication. now system is working fine. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFDL24162871944 | 2025-02-03 | 2025-02-07 | Cust 943 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Turbidity Meter | DIS/Q(182) | commissioning of Radaar Transmitter and Turbidity meters at Our Rohtak Site. | All the Level Transmitters & Turbidity Analyzer have been successfully commissioned. Details are in the attached MOM. | All the Level Transmitters & Turbidity Analyzer have been successfully commissioned. Details are in the attached MOM.will confirm availability of standard solution for next visit. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24163181945 | 2025-02-04 | 2025-03-17 | Cust 564 | DELHI | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | Conductivity Transmitter Aquacon | Aquacon 2000(AQCN 2000 p) | 1. DO Analyzer needs to calibrate 2. 2 nos pH Transmitter not getting Power ON 3. Display of one pH Transmitter not showing | Material received in depot. repaired & sent back to site | Material received in depot. repaired &amp;amp;amp; sent back to site | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFDL24163551946 | 2025-02-04 | 2025-02-05 | Cust 832 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Visit for Phosphate Analyzers at ETP & STP unit. | Servicing and calibration done of Phosphate Analyzers, Now Both the Analyzers are working satisfactorily. Details are in attached MOM. | Servicing and calibration done of Phosphate Analyzers, Now Both the Analyzers are working satisfactorily. Details are in attached MOM. | EXTENDED | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24163571947 | 2025-02-04 | 2025-02-09 | Cust 919 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Turbidity Meter | Optex turbidity sensor(1) | Visit for commissioning | visit done for water quality analyzers as per attached mail communication. | visit done for water quality analyzers as per attached mail communication. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24163841948 | 2025-02-05 | 2025-03-17 | Cust 730 | DELHI | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | MULTIPARAMETER ANALYSER | FM make RS485 card(1) | FM make RS485 card not communicate, need replacement urgently | Replacement Material sent on 17th Feb 25 | Replacement Material sent on 17th Feb 25 | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFDL24163851949 | 2025-02-05 | 2025-02-07 | Cust 692 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | software re-installation+ CPCB, Local and MPCB configuration | We have re-installed the software and completed Local, CPCB and RSPCB connectivity and the data is now live on the same servers. Also updated Enviroconnect DAS version. | We have re-installed the software and completed Local, CPCB and RSPCB connectivity and the data is now live on the same servers. Also updated Enviroconnect DAS version. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24164051950 | 2025-02-05 | 2025-02-11 | Cust 711 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | TOC ANALYSER | FM 101 TOC Analyser(FM 101 TOC) | Visit for commissioning | Visit done for FMTOC 101 analyzer commissioning | visit done as per attached mom | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24164491951 | 2025-02-07 | 2025-02-07 | Cust 956 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | THERMO SERIES | TSS SENSOR-RT2143(1) | REMOTE SUPPORT TO COMMISSION 1 NO. THERMO SCIENTIFIC SILICA ANALYZER IN PANKI POWER KANPUR DM PLANT. | THERMO SILICA ANALYZER COMMISSIONED ON REMOTE SUPPORT | THERMO SILICA ANALYZER COMMISSIONED ON REMOTE SUPPORT | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24164591952 | 2025-02-07 | 2025-02-07 | Cust 559 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data and camera are not live on CPCB server. | We took access to the system and now data and Camera is live on CPCB server. | We took access to the system and now data and Camera is live on CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24165611953 | 2025-02-10 | 2025-02-10 | Cust 800 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Engineer visit required for AMC of MPA | Engineer visited site and done the calibration of MPA | Resolved | NO | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFDL24166341954 | 2025-02-11 | 2025-02-14 | Cust 810 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Visit done for MPA | visit done for MPA.MOM attached . | visit done as per attached mom. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL24166641955 | 2025-02-12 | 2025-02-15 | Cust 608 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data uploaded issue. | Found internet issue and same has been resolved by their IT team and Data is live on the server. | Found internet issue and same has been resolved by their IT team and Data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24166951956 | 2025-02-12 | 2025-02-16 | Cust 874 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Visit scheduled for the replacement of the Xenon Power Supply and completion of the calibration of the Multi-Parameter Analyzer. | Multi-Parameter Analyzer is working satisfactorily, after replacing the spares and Calibration also done. Details are in the attached MOM. | Multi-Parameter Analyzer is working satisfactorily, after replacing the spares and Calibration also done. Details are in the attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24167791957 | 2025-02-14 | 2025-02-14 | Cust 546 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB+RSPCB server. | We have checked and data is live on both CPCB+RSPCB server. | We have checked and data is live on both CPCB+RSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24169121958 | 2025-02-18 | 2025-03-10 | Cust 957 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | DATCON | DATCON(DATCON) | CPCB and UPPCB data connectivity. | From the division, a new RTU has been arranged and the same has been under dispatch. | From the division, a new RTU has been arranged and the same has been under dispatch. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24169151959 | 2025-02-18 | 2025-02-19 | Cust 823 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on LOCAL+CPCB+RSPCB server. | We took access to the system and found there was time sync issue and CPCB configuration issue. Now data is live on LOCAL+CPCB+RSPCB server. | We took access to the system and found there was time sync issue and CPCB configuration issue. Now data is live on LOCAL+CPCB+RSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24169591960 | 2025-02-19 | 2025-02-21 | Cust 692 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB server. | We took access to the system and hardware side issue has been resolved and data is coming till modscan and is live on CPCB server | We took access to the system and hardware side issue has been resolved and data is coming till modscan and is live on CPCB server | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24169951961 | 2025-02-20 | 2025-02-21 | Cust 802 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | HSPCB data uploaded issue. | Commercially not clear. | Commercially not clear. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24171291962 | 2025-02-24 | 2025-03-01 | Cust 869 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | CPU Display card for Multiparameter(AWA-CPU1) | Visit for calibration | visit done for checking & calibration of MPA | visit done as per attached mom. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL24171431963 | 2025-02-24 | 2025-02-26 | Cust 929 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | Cleaning Pump(1) | AMC | Calibration done | Calibration done | YES | Periodic Maintenance (AMC) | Process | N |  | Analytical - CHAKAN
AFDL24171481964 | 2025-02-24 | 2025-02-28 | Cust 849 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Aqua4Trans | Display card of Aqua4Trans(1) | Vertical lines appearing in display. It will be replaced in 15 transmitters. | Replaced the display and attended the Analyzers, now full Analyzers are working satisfactorily. Details are in the attached NOM. | Replaced the display and attended the Analyzers, now full Analyzers are working satisfactorily. Details are in the attached NOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24171631965 | 2025-02-24 | 2025-02-25 | Cust 790 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 2nd AMC visit | 2nd AMC visit completed | Working fine | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24172381966 | 2025-02-25 | 2025-03-03 | Cust 799 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Reconfiguration of STP Analyzer | As reported that the Analyzers have been reset. Therefore Re-configuration done through telephonic support of both the Multi-Parameter Analyzer (ETP & STP). Now both the Analyzers have been working satisfactorily. So that we are closing this complaint. | We have Re-configured both the Multi-Parameter Analyzers through telephonic support of (ETP and STP). Now both the Analyzers have been working satisfactorily. So that we are closing this complaint. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFDL24172571967 | 2025-02-26 | 2025-03-05 | Cust 554 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | Dr. Reddy's Laboratories Ltd. Plot No 57, Village-Kunjhal, Baddi-174103, Himachal Pradesh End user Contact Person Name :- Mohammad Mubasshir Bijapuri | Work done | Work done | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24172621968 | 2025-02-26 | 2025-02-28 | Cust 786 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Schedule visit for ETP online Analyzer against CAMC. | Visit for online Multi-Parameter Analyzer has been completed against CAMC, observed Analyzer is working satisfactorily. Details are in attached MOM. | Visit for online Multi-Parameter Analyzer has been completed against CAMC, observed Analyzer is working satisfactorily. Details are in attached MOM. | EXTENDED | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24172981969 | 2025-02-26 | 2025-03-01 | Cust 958 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Aqua4Trans | Aqua4 trans-A4TW033000(1) | Site calibration assistance for DO, Turbidity and LT | Site calibration visit completed successfully | Working fine | YES | Paid | Core | N |  | Analytical - CHAKAN
AFDL24176081970 | 2025-03-01 | 2025-03-02 | Cust 803 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 3rd AMC visit | 3rd AMC visit | Working fine | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24176621971 | 2025-03-03 | 2025-03-06 | Cust 931 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to RSPCB server | we have re-installed the software with an updated version and completed Local, CPCB and RSPCB connectivity and the data is now live on the same servers. | we have re-installed the software with an updated version and completed Local, CPCB and RSPCB connectivity and the data is now live on the same servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24176801972 | 2025-03-03 | 2025-05-13 | Cust 959 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | pH Transmitter Aquamon | Ph Transmitter Aquamon 2000(PH2000) | Out of three aquamon 2000 one is working two are not working. | NO issue Noticed | Units were received in Depot But no Issue noticed. Ref. CCS25D0028 | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFDL24176801972 | 2025-03-03 | 2025-05-13 | Cust 959 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | pH Transmitter Aquamon | Ph Transmitter Aquamon 2000(PH2000) | Out of three aquamon 2000 one is working two are not working. | No Issue Noticed | Units were received in Depot But no Issue noticed. Ref. CCS25D0028 | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFDL24177241973 | 2025-03-04 | 2025-03-04 | Cust 960 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(Ph-Trans-Smartpro) | Commissioning the pH, Turbidity, TSS & Chlorine Analyzers. | Customer needs to procure spares. Refer attached MOM | CLosed | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24177411974 | 2025-03-04 | 2025-03-06 | Cust 812 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Checking of MPA | Cheked the MPA Analyzer and found ok. | Checked the MPA Analyzer | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24177451975 | 2025-03-04 | 2025-03-10 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | AMC | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | Gate Pass Formalities to be completed against Analyzer AMC. | Gate Pass Formalities has been completed against Analyzer AMC. | Gate Pass Formalities has been completed against Analyzer AMC. | NO | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24178291976 | 2025-03-06 | 2025-03-10 | Cust 812 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning of FMBB | Done the configuration of FMBB | Commissioned the FMBB | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24178301977 | 2025-03-06 | 2025-03-10 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Schedule visit for ETP#3 online Analyzer against AMC. | Visit for COD, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | Visit for COD, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24178541978 | 2025-03-07 | 2025-03-10 | Cust 812 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Data breakage issue on the CPCB srever. | Found a network issue and the same has been resolved by their IT team and Additionally, we have also configured the new FM BB for the ETP station. now Data is uploading continuously on the CPCB server for both the units. | now Data is uploading continuously on the CPCB server for both the units. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24178621979 | 2025-03-07 | 2025-03-10 | Cust 802 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software re-installation and CPCB & HSPCB connectivity | We have reinstalled the software with an upgraded version and completed Local, CPCB & HSPCB connectivity. For HSPCB, data for all industries is not updating on the HSPCB server and we are currently checking this issue with the HSPCB team. | We have reinstalled the software with an upgraded version and completed Local, CPCB &amp;amp; HSPCB connectivity. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24179241980 | 2025-03-09 | 2025-03-12 | Cust 954 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | DO Transmitters | VISI pro DO(VISI pro DO) | Visit for checking of DO Analyzer | visit done as per attached mail communication . | visit done as per attached mail communication . | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24180331981 | 2025-03-10 | 2025-03-12 | Cust 807 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | LCD of Multiparameter Display(1) | Visit for checking and calibration of multiparameter Analyzer | visit done for MPA checking under EUS | visit done for MPA checking under EUS | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24180451982 | 2025-03-11 | 2025-03-24 | Cust 717 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration job | Work done | Work done | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL24180461983 | 2025-03-11 | 2025-03-18 | Cust 961 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | calibration | Work done | Work done | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL24180471984 | 2025-03-11 | 2025-03-18 | Cust 730 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC | Work done | Work done | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24180641985 | 2025-03-11 | 2025-03-14 | Cust 938 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB and Digital connectivity needs to be done for ETP | Here there is a network issue and our branch engineer is working on that hence as of now we are closing this complaint. | Here there is a network issue and our branch engineer is working on that hence as of now we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24180841986 | 2025-03-11 | 2025-03-14 | Cust 823 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data breaking on CPCB server. | We did changes in the software but today system was not accessible as of now we are closing this complaint. | We did changes in the software but today system was not accessible as of now we are closing this complaint. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24180981987 | 2025-03-11 | 2025-03-14 | Cust 809 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 3rd AMC visit for Multiparameter and Oil in water analyzer | 3rd AMC visit completed | Working fine | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24181351988 | 2025-03-12 | 2025-03-15 | Cust 962 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | visit for commissioning of FM Black box | visit done for FMBB Panel .details in attached MOM. | visit done for FMBB Panel .details in attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24181711989 | 2025-03-13 | 2025-03-14 | Cust 803 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Emergency visit to solve Multiparameter issue | Emergency visit completed and resolved the multiparameter issue | Working fine | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24181721990 | 2025-03-13 | 2025-03-14 | Cust 922 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Blackbox Commissioning | We have completed the Blackbox Configuration and Data is coming till the FMBB | We have completed the Blackbox Configuration and Data is coming till the FMBB | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL24182241991 | 2025-03-15 | 2025-04-21 | Cust 962 | DELHI | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Rs-485 isolator CSCAN RPT 485 S/n- U1212425692 was found not in working . | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210138099 LR Dt - 16-APR-2025 | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210138099 LR Dt - 16-APR-2025 | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFDL24182481992 | 2025-03-16 | 2025-03-22 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | AMC | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | Scheduled AMC visit for COD, TOC, Oi l in Water, Sulphide Analyzers, installed at ETP # 1, 2 & TTPRO. | Visit for COD, TOC, Oil in Water, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | Visit for COD, TOC, Oil in Water, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24182851993 | 2025-03-17 | 2025-03-18 | Cust 484 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data Breaking on RSPCB server. | There was issue from customer end. Now issue has been resolved and data is live on RSPCB server. | There was issue from customer end. Now issue has been resolved and data is live on RSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24182861994 | 2025-03-17 | 2025-03-23 | Cust 532 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data breaking on RSPCB server. | We have updated the patches and delay data file and keep the system under observation. | We have updated the patches and delay data file and keep the system under observation. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24183211995 | 2025-03-18 | 2025-03-21 | Cust 823 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data breaking on CPCB server. | We took access to the system and observed data is breaking on MODSCAN need to check from hardware end. | We took access to the system and observed data is breaking on MODSCAN need to check from hardware end. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24183301996 | 2025-03-18 | 2025-03-20 | Cust 870 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation + Local + CPCB connectivity | We have reinstalled the software with an updated version and completed CPCB connectivity. Also Updated Enviroconnect DAS version to the latest and data is now live on the Local and CPCB servers. | We have reinstalled the software with an updated version and completed CPCB connectivity. Also Updated Enviroconnect DAS version to the latest and data is now live on the Local and CPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24183351997 | 2025-03-18 | 2025-03-20 | Cust 843 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 4th AMC visit | 4th AMC visit completed | Working fine | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24184191998 | 2025-03-19 |  | Cust 963 | DELHI | Steam System Division | ASSIGNED | SYSTEM PERFORMANCE | PRV - MEASURETROL | FM MEASURETROL(SIZE 100NB) | After updating the software of the measuretrol still client is having the issue in the flow reading, as per client comments the flow is coming at lower side. |  |  | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFDL24184481999 | 2025-03-20 | 2025-03-21 | Cust 533 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 3rd AMC visit | 3rd AMC visit completed | Working fine | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24184822000 | 2025-03-21 | 2025-03-26 | Cust 801 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 3rd AMC visit | 3rd AMC visit completed | Working fine | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24186222001 | 2025-03-24 | 2025-03-30 | Cust 935 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | RS485 CARD for Multiparameter(AWA-RS485) | Connectivity to be established of online Analyzers. | Commissioned online Multi-Parameter & TSS Analyzers in ETP Plant, Observed working satisfactorily. Details are in attached MOM. | Commissioned online Multi-Parameter & TSS Analyzers in ETP Plant, Observed working satisfactorily. Details are in attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24186222001 | 2025-03-24 | 2025-03-30 | Cust 935 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4Trans(1) | Commissioned online Multi-Parameter & TSS Analyzers in ETP Plant, Observed working satisfactorily. Details are in attached MOM. | Commissioned online Multi-Parameter & TSS Analyzers in ETP Plant, Observed working satisfactorily. Details are in attached MOM. | Commissioned online Multi-Parameter & TSS Analyzers in ETP Plant, Observed working satisfactorily. Details are in attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24187972002 | 2025-03-28 | 2025-03-31 | Cust 964 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning of MPA with TDS | Commissioned as per the attached MOM | Working fine | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24187972002 | 2025-03-28 | 2025-03-31 | Cust 964 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning of MPA with TDS | Commissioned as per the attached MOM | Working fine | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24188182003 | 2025-03-28 | 2025-03-31 | Cust 692 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to RSPCB server. | We have checked and observed that the data is continuously reflecting to the RSPCB server. | We have checked and observed that the data is continuously reflecting to the RSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL24188782004 | 2025-03-29 | 2025-03-30 | Cust 786 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | commissioning | commissioning has been done | Commissioning has been done. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24188802005 | 2025-03-29 | 2025-03-30 | Cust 872 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 1AMC | AMC done | AMC done | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24188812006 | 2025-03-29 | 2025-03-30 | Cust 709 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 1 AMC | AMC done | AMC done | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24188822007 | 2025-03-29 | 2025-03-30 | Cust 730 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC | Troubleshooted and RS485 card found faulty | Troubleshooted and RS485 card found faulty | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24188832008 | 2025-03-29 | 2025-03-30 | Cust 708 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | commissioning | commissioning done | Commissioning has been done. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24188852009 | 2025-03-29 | 2025-03-30 | Cust 927 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 2AMC | AMC done | AMC done | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24188862010 | 2025-03-29 | 2025-03-30 | Cust 965 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 1 AMC | AMC done | AMC done | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL24189812011 | 2025-03-30 | 2025-03-31 | Cust 912 | DELHI | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | installation and commissioning guideline | commissioning guidelines given to site project manager | commissioning guidelines given to site project manager | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL24189822012 | 2025-03-30 | 2025-03-31 | Cust 723 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | calibration and troubleshooting of 485 o/p | Calibration done & RS485 getting through existing card | Calibration done & RS485 getting through existing card | NO | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL25000102013 | 2025-04-01 | 2025-04-04 | Cust 870 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | localhost login issue | As of now we are closing this complaint. We will reopen this once we get access to the system to proceed further. | As of now we are closing this complaint. We will reopen this once we get access to the system to proceed further. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL25000182014 | 2025-04-01 | 2025-04-03 | Cust 966 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Aqua4Trans | Aqua4 trans-A4TW033000(1) | To check pH and conductivity meters | pH and conductivity meters - 8 nos. checked and resolved | Working fine | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL25000822015 | 2025-04-03 | 2025-04-15 | Cust 548 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | pH Transmitter Aquamon | Ph Transmitter Aquamon 2000(PH2000) | Customer received a pH meter complaint from IOLCP; refer to the below mail. Following is some of the issues observed in Purified water system- High Conductivity at a distribution loop - Consulted the issues with Mr. Chetan ( M/s Aquamech ) and troubleshooting done by following the instructions with help of instrument team, Found that the conductivity controller need to be replaced. Feed water High pH - pH meter need to be calibrated (Required controller manual for calibration) (Refer below mentioned pic for details) Site Address: M/S. IOL CHEMICAL AND PHARMACEUTICAL LTD. Trident Complex, Village Fatehgarh Chana, Mansa Road, Dhaula Barnala - 148101, Panjab Mr. Gagan +91 9875969238 | conductivity analyzer found within accuracy limit but calibrated and pH sensor found broken | conductivity analyzer found within accuracy limit but calibrated and pH sensor found broken | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL25001112016 | 2025-04-04 | 2025-04-07 | Cust 870 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software re-installation | We have reinstalled the software and data is now live on the Local and CPCB servers. | We have reinstalled the software and data is now live on the Local and CPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL25001262017 | 2025-04-04 | 2025-04-17 | Cust 802 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC visit | Done the AMC visit | Done the calibration of Multiparameter. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL25002842018 | 2025-04-08 | 2025-04-11 | Cust 823 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data availability is less. | Issue has been resolved and now data availability has been increased. | Issue has been resolved and now data availability has been increased. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL25002912019 | 2025-04-08 | 2025-04-11 | Cust 931 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data from stack is not uploading to CPCB server. | As of now we are closing this complaint. we will reopen this once hardware side issue is resolved. | As of now we are closing this complaint. we will reopen this once hardware side issue is resolved. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL25003112020 | 2025-04-08 | 2025-04-26 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Schedule visit for ETP#3 online Analyzer against AMC. | Visit for COD, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | Visit for COD, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL25003272021 | 2025-04-09 | 2025-04-10 | Cust 793 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to local and RSPCB server. | We have re-installed the software with an upgraded version and data is now live on the CPCB, Local and RSPCB servers. Also we have updated the Enviroconnect DAS version to the latest. | We have re-installed the software with an upgraded version and data is now live on the CPCB, Local and RSPCB servers. Also we have updated the Enviroconnect DAS version to the latest. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL25003572022 | 2025-04-09 | 2025-04-13 | Cust 814 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | DATCON | DATCON(DATCON) | Visit for commissioning of RTU301 under CPCB/SPCB | visit done for RTU301 under CPCB & UPPCB commissioning . | visit done for RTU301 under CPCB & UPPCB commissioning . | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL25003672023 | 2025-04-10 | 2025-04-11 | Cust 533 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data uploading issue on CPCB and RSPCB server. | Multiparameter connection is through TCP IP and we were not receive a ping response from the configured IP address on your PC. Now the same is pinging and data is live on the CPCB server. | Multiparameter connection is through TCP IP and we were not receive a ping response from the configured IP address on your PC. Now the same is pinging and data is live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL25003862024 | 2025-04-10 | 2025-04-10 | Cust 793 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MUVX Card of Multiparameter(1) | MPA not working due to reset issue in Display. | Engineer visited site and found MUVX card slot faulty. Replaced with spares input slot, now analyzer working fine. MOM attached. | Resolved | NO | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFDL25003992025 | 2025-04-10 | 2025-04-11 | Cust 957 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | Digital Product | RTU301() | RTU Commissioning and CPCB + UPPCB connectivity needs to be done | Successful Commissioning of RTU301- MPA parameters under CPCB & UPPCB Server. and data is live on both the side. | Successful Commissioning of RTU301- MPA parameters under CPCB &amp;amp; UPPCB Server. and data is live on both the side. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL25004042026 | 2025-04-10 | 2025-04-14 | Cust 955 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Windows-based software upgradation needs to be done. | As of now we are closing this complaint and will reopen it once the hardware side issue is resolved. | As of now we are closing this complaint and will reopen it once the hardware side issue is resolved. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL25004062027 | 2025-04-10 | 2025-04-11 | Cust 786 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | RSPCB camera streming issue. | currently, there is a network issue at the site, and the customer is working on that; hence, as of now, we are closing this complaint. | currently, there is a network issue at the site, and the customer is working on that; hence, as of now, we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL25004112028 | 2025-04-10 | 2025-04-11 | Cust 786 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software installation and CPCB + SPCB needs to be done. | We have successfully completed the software installation and established CPCB + HSPCB connectivity today. The data is now live on both sites. | We have successfully completed the software installation and established CPCB + HSPCB connectivity today. The data is now live on both sites. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL25004252029 | 2025-04-10 | 2025-04-13 | Cust 919 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Chlorine sensor | CHLORINE SENSOR(1) | Visit for commissioning of chlorine Analyzer | visit done as per attached mail communication. | visit done as per attached mail communication. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL25004922030 | 2025-04-13 | 2025-05-30 | Cust 919 | DELHI | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | Turbidity Meter | Optex turbidity sensor(1) | 02 nos Turbidity Analyzer Model- FMTC-Mi100,S/n- 2308052970003S & S/n- 2308052970005S checked & found both analyzer deaerator missing & one analyzer S/n- 2308052970003S is not giving 4-20 mA output hence we will check with our team and revert back. | In commercial short supply complaint was logged. New complaint No.- FDL25C0263. | In commercial short supply complaint was logged. New complaint No.- FDL25C0263. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFDL25005062031 | 2025-04-14 | 2025-04-17 | Cust 933 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | DATCON | DATCON(DATCON) | Data is not uploading to CPCB server. | Now the issue has been resolved and data is live on the CPCB server. We have made the required changes in configuration (As per the HSPCB, data uploading URL changed from http to https). Now the data is reflecting continuously on the HSPCB server. | Now the issue has been resolved and data is live on the CPCB and HSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL25005152032 | 2025-04-14 | 2025-04-21 | Cust 730 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc | clean the flow cell and also power pcb . | calibration job done. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL25005192033 | 2025-04-14 | 2025-04-14 | Cust 841 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Analyzer showing zero TSS, need engineer visit. | Engineer visited site and resolved issue. | Resolved | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFDL25005592034 | 2025-04-15 | 2025-04-18 | Cust 806 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to CPCB and HSPCB server. | As of now we are closing this complaint. We will reopen this once the issue is resolved from hardware side. | As of now we are closing this complaint. We will reopen this once the issue is resolved from hardware side. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL25005792035 | 2025-04-15 | 2025-04-22 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | visit for attending 22 nos of special analyzers under ETP ARC | Visit done to attend 22 nos analyzers at site .mom attached. | visit done as per attached mom | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL25006012036 | 2025-04-16 | 2025-04-21 | Cust 954 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | AQUADO Analyzer | DPI-Digital process indicator(1) | Replacement of DO tx,meter | Replace the DO tx'meter with spare & found working fine. | Replace the DO tx'meter with spare & found working fine. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL25006462037 | 2025-04-17 | 2025-04-23 | Cust 832 | DELHI | Analytical - CHAKAN | CLOSED | SALES PROMOTION | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | Visit for spares | Discussed for spares of pH,Conductivity ,DO, Turbidity Analyzer of STP.ETP and WTP plant Analyzer. | Sapres discussion with the customer. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL25006492038 | 2025-04-17 | 2025-04-21 | Cust 799 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Checking and calibration | Done the calibration of Multiparameter. | Done the calibration of Multiparameter. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL25006492038 | 2025-04-17 | 2025-04-21 | Cust 799 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Checking and calibration | Done the calibration of Multiparameter. | Done the calibration of Multiparameter. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL25006522039 | 2025-04-17 | 2025-04-18 | Cust 803 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software re-installation and Local, CPCB and RSPCb connectivity | We have re-installed the software with an updated version and completed Local, CPCB, RSPCB connectivity and data is now live on the same servers. Also updated Enviroconnect DAS version to the latest. | We have re-installed the software with an updated version and data is now live on the Local, CPCB, RSPCB servers. Also updated Enviroconnect DAS version to the latest. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL25006772040 | 2025-04-18 | 2025-04-21 | Cust 806 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB and HSPCB data upload issue. | We took access of the system and identified that the ETP and Flow data are not communicating up to the local PC hence As of now, we are closing this complaint and will reopen this case once the hardware issue is resolved from customer side. | As of now, we are closing this complaint and will reopen this case once the hardware issue is resolved from customer side. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL25007082041 | 2025-04-18 | 2025-04-20 | Cust 803 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 4th AMC visit for multiparameter analyzer, dust analyzer and 2 nos. mag flow meters | 4th AMC visit completed | Working fine | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL25007352042 | 2025-04-20 | 2025-04-26 | Cust 786 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Schedule visit for ETP online Analyzer against CAMC. | Visit for online Multi-Parameter Analyzer has been completed, observed Analyzer is working satisfactorily. Details are in attached MOM. | Visit for online Multi-Parameter Analyzer has been completed, observed Analyzer is working satisfactorily. Details are in attached MOM. | EXTENDED | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL25007562043 | 2025-04-21 | 2025-04-23 | Cust 844 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | 24 VDC SMPS (1) | 2 visit of Amc | Second visit of AMC is completed | calibration job done. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL25007572044 | 2025-04-21 | 2025-04-23 | Cust 858 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | Cleaning Pump(1) | 2nd visit of Annual Amc | second Visit of Amc is completed and analyzer is working fine | calibration job done. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL25007652045 | 2025-04-21 | 2025-04-28 | Cust 608 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | pH input card-FM of Multiparameter(1) | Amc 1st visit | first Visit of Amc is completed. | calibration job done. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL25007872046 | 2025-04-21 | 2025-04-26 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | AMC | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | Scheduled AMC visit for COD, TOC, Oi l in Water, Sulphide Analyzers, installed at ETP # 1, 2 & TTPRO. | Visit for COD, TOC, Oil in Water, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | Visit for COD, TOC, Oil in Water, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL25007972047 | 2025-04-21 | 2025-04-28 | Cust 786 | DELHI | Analytical - CHAKAN | CLOSED | DIGITAL CONEECTIVITY | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | commissioning of MPA CX2000 blackbox | TSS issue resolved and FMBB connectivity with cloud server done. | TSS issue resolved and FMBB connectivity with cloud server done. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL25008022048 | 2025-04-22 | 2025-04-23 | Cust 803 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to Local, CPCB and RSPCB servers. | As of now we are closing this complaint. We will reopen this once we get access to your system after the new CPU is arranged. | As of now we are closing this complaint. We will reopen this once we get access to your system after the new CPU is arranged. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL25008032049 | 2025-04-22 | 2025-04-25 | Cust 806 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to CPCB server. | As of now we are closing this complaint. We will reopen this once we get access to your system. | As of now we are closing this complaint. We will reopen this once we get access to your system. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL25008322050 | 2025-04-22 | 2025-04-28 | Cust 868 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | CPU Display card for Multiparameter(AWA-CPU1) | Visit for checking of MPA | visit done as per attached mom. | visit done as per attached mom. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL25008482051 | 2025-04-23 | 2025-04-25 | Cust 803 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software re-installation and Local + CPCB + RSPCB connectivity | We have re-installed the software with an upgraded version and completed Local, CPCB and RSPCB connectivity. | We have re-installed the software with an upgraded version and completed Local, CPCB and RSPCB connectivity. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL25008492052 | 2025-04-23 | 2025-04-28 | Cust 965 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | Flow cell for Multiparameter(AWA-FLOWCELL) | Amc visit | Amc Visit is completed | calibration job done. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL25009002053 | 2025-04-24 | 2025-04-26 | Cust 905 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | HSPCB url needs to be update. | we have accessed your system and successfully modified the data posting URL from HTTP to HTTPS.The data is now being uploaded through the HTTPS URL to the HPSB server. | we have accessed your system and successfully modified the data posting URL from HTTP to HTTPS.The data is now being uploaded through the HTTPS URL to the HPSB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL25009052054 | 2025-04-24 | 2025-04-26 | Cust 807 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | Digital Product | RTU301() | RTU commissioning for CPCB and UPPCB server. | The RTU configuration has been done and now the Data is uploaded to the CPCB, UPPCB and cloud server from the RTU. | The RTU configuration has been done and now the Data is uploaded to the CPCB, UPPCB and cloud server from the RTU. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL25009152055 | 2025-04-24 | 2025-04-25 | Cust 799 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software re-installation and Local + CPCB + HSPCB connectivity. | We have re-installed the software with an upgraded version and completed Local and CPCB , HSPCB connectivity and successfully modified the HSPCB data uploading URL from HTTP to HTTPS. Data is now live on the same servers | We have re-installed the software with an upgraded version and completed Local and CPCB , HSPCB connectivity and successfully modified the HSPCB data uploading URL from HTTP to HTTPS. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL25009472056 | 2025-04-25 | 2025-04-26 | Cust 793 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | RSPCB data uploaded issue. | We have updated the Delayed Data jar file so now the Data is live on the server. | We have updated the Delayed Data jar file so now the Data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL25010712057 | 2025-04-28 | 2025-04-30 | Cust 803 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software re-installation and Local + CPCB + RSPCB connectivity | We have re-installed the software with an updated version and completed Local, CPCB and RSPCb connectivity and data is now live on the same servers. | We have re-installed the software with an updated version and completed Local, CPCB and RSPCb connectivity and data is now live on the same servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL25011992058 | 2025-04-29 | 2025-05-02 | Cust 532 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on RSPCB server. | Data is now live on RSPCB server. | Data is now live on RSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL25012012059 | 2025-04-29 | 2025-05-02 | Cust 546 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on RSPCB server. | Data is now live on RSPCB server. | Data is now live on RSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL25014242060 | 2025-05-01 | 2025-05-02 | Cust 837 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | Cleaning Pump(1) | Amc peroidic visit | Done the calibration of MPA | Done the calibration of Multiparameter. | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL25014432061 | 2025-05-01 | 2025-05-11 | Cust 962 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Rs-485 Isolator replacement under FMBB Panel. | visit for the replacement of 24 VDC RS-485 Isolator replacement under supplied FMBB Panel. | visit done for replacement of 24 VDC RS-485 Isolator replacement | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL25014552062 | 2025-05-01 | 2025-05-04 | Cust 786 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Calibration of pH Analyzers. | pH Analyzers calibration has been completed and observed pH Analyzers are working satisfactorily. Details are in attached MOM. | pH Analyzers calibration has been completed and observed pH Analyzers are working satisfactorily. Details are in attached MOM. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL25014612063 | 2025-05-02 | 2025-05-02 | Cust 803 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data mismatch issue for Stack 2- 6 TPH | Now the data mismatch issue has been resolved and correct readings are reflecting on the RSPCB server. | Now the data mismatch issue has been resolved and correct readings are reflecting on the RSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL25015532064 | 2025-05-05 | 2025-05-08 | Cust 806 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue. | checked the data on MODSCAN, but data is not coming till the Local PC for CETP Outlet so for now we are closing this complaint and will reopen it once the hardware side issue is resolved. | checked the data on MODSCAN, but data is not coming till the Local PC for CETP Outlet so for now we are closing this complaint and will reopen it once the hardware side issue is resolved. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL25015632065 | 2025-05-05 | 2025-05-08 | Cust 809 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB server. | We took access to the system and completed the installation and configuration of software and now data is live on LOCAL+CPCB+RSPCB server. | We took access to the system and completed the installation and configuration of software and now data is live on LOCAL+CPCB+RSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL25016352066 | 2025-05-06 | 2025-05-10 | Cust 571 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | As per customer that the online reading of pH & Conductivity Analyzers are not matching with the lab reading. | The calibration of pH & conductivity Analyzers have been completed, now all the Analyzers have been working satisfactorily. Details are in the attached MOM. | The calibration of pH & conductivity Analyzers have been completed, now all the Analyzers have been working satisfactorily. Details are in the attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL25016482067 | 2025-05-06 | 2025-05-08 | Cust 806 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to Local, CPCB and HSPCB servers. | Now the data is live on the Local, CPCB and HSPCB servers. | Now the data is live on the Local, CPCB and HSPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL25016852068 | 2025-05-07 | 2025-05-10 | Cust 891 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to HSPCB server. | We accessed the system and observed that some software related files were missing from the PC and localhost was not accessible. Hence to fix this issue we have re-installed the enviroconnect software and completed Local, CPCB and HSPCB connectivity. Also disabled the Antivirus from the PC. Now the localhost is accessible and data is live on the same. Currently, the HSPCB server is not accessible and we are checking this issue with the HSPCB team. | we have re-installed the enviroconnect software and completed Local, CPCB and HSPCB connectivity.Currently, the HSPCB server is not accessible and we are checking this issue with the HSPCB team. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL25017002069 | 2025-05-07 | 2025-05-08 | Cust 692 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to CPCB and RSPCB server. | We have checked and data is showing live on the CPCB and RSPCB server. | We have checked and data is showing live on the CPCB and RSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL25017322070 | 2025-05-08 | 2025-05-10 | Cust 566 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning visit for online Multi-Parameter Analyzer. | MPA Analyzer couldn't commissioned due to unavailability of sample. Details are in the attached MOM. | MPA Analyzer couldn't commissioned due to unavailability of sample. Details are in the attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL25017532071 | 2025-05-08 | 2025-05-14 | Cust 636 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration of the MPA. | Visit for online Multi-Parameter Analyzer has been completed, observed Analyzer is working satisfactorily. Details are in attached MOM. | Visit for online Multi-Parameter Analyzer has been completed, observed Analyzer is working satisfactorily. Details are in attached MOM. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL25017862072 | 2025-05-09 | 2025-05-12 | Cust 786 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | RSPCB data uploded issue. | we took access of the Blackbox and observed that the internet LAN cable connection was wrong because that the internet was not pinging.After making the necessary corrections, the internet starts pinging. However, we have observed that data uploads from all industries are not currently being displayed on the RSPCB server.We have already escalated this matter to the RSPCB team hence as of now we are closing this compalaint. | Here observed that the internet LAN cable connection was wrong .After making the necessary corrections, the internet starts pinging hence as of now we are closing this compalaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL25017882073 | 2025-05-09 | 2025-05-12 | Cust 786 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data uploded issue. | We have accessed the system and identified that the C drive was fulll, hence we have deleted some old log files, so now the issue has been resolved and the data is live on the CPCB server. | We have accessed the system and identified that the C drive was fulll, hence we have deleted some old log files, so now the issue has been resolved and the data is live on the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL25017982074 | 2025-05-10 | 2025-05-14 | Cust 837 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Data is not transferring pf online Multi-Parameter Analyzer with the CPCB & SPCB server. | Visit for online Multi-Parameter Analyzer has been completed, observed Analyzer is working, but connectivity couldn't establish through DATCon. Our Pune team will check and resolve it as soon as possible. | Visit for online Multi-Parameter Analyzer has been completed, Analyzer is working, but connectivity couldn't establish through DATCon. Our Pune team will check and resolve it as soon as possible. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL25018612075 | 2025-05-12 | 2025-05-15 | Cust 732 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB server. | As of now we are closing this complaint. We will reopen once we receive access from customer. | As of now we are closing this complaint. We will reopen once we receive access from customer. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL25018662076 | 2025-05-12 | 2025-05-14 | Cust 801 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 4th AMC visit | 4th AMC visit completed | Working fine | EXTENDED | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL25019742077 | 2025-05-14 | 2025-05-20 | Cust 856 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | SODIUM ANALYSER | SODIUM 9073 single channel ANLAYSER(9073 Sodium) | COPRA Silica reagnet-4 not consuming and sodium swan slope and stability error. | Work done | Work done | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL25019762078 | 2025-05-14 | 2025-05-18 | Cust 843 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 1st AMC visit to check BOD output issue | 1st AMC visit completed | mA output card faulty and needs to be replaced | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL25019822079 | 2025-05-14 | 2025-05-17 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Schedule visit for ETP#3 online Analyzer against AMC. | Visit for COD, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | Visit for COD, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL25020572080 | 2025-05-16 | 2025-05-16 | Cust 732 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB server. | We took access to the system and observed there is issue from hardware end. | We took access to the system and observed there is issue from hardware end. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL25020792081 | 2025-05-17 | 2025-05-20 | Cust 786 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Camera RTSP link is required for RSPCB camera connectivity. | We have created the RTSPB link, but the same has not been playing out of the customer's network; hence, we have asked them to check this with their IT team. As of now, we are closing this complaint. | We have created the RTSPB link, but the same has not been playing out of the customer's network; hence, we have asked them to check this with their IT team. As of now, we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL25021962082 | 2025-05-19 | 2025-05-26 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | AMC | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | Scheduled AMC visit for COD, TOC, Oi l in Water, Sulphide Analyzers, installed at ETP # 1, 2 & TTPRO. | Visit for COD, TOC, Oil in Water, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | Visit for COD, TOC, Oil in Water, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL25022342083 | 2025-05-20 | 2025-05-26 | Cust 810 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration | Attended dust, MPA, GAs analyser agaist AMC and found working fine. | Attended dust, MPA, GAs analyser agaist AMC and found working fine. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL25022912084 | 2025-05-21 | 2025-05-26 | Cust 824 | DELHI | Analytical - CHAKAN | CLOSED | SALES PROMOTION | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Going to gail for promotion | site visit done | site visit done | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL25023012085 | 2025-05-21 | 2025-05-26 | Cust 927 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 3 AMC | AMC done | AMC done | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL25023092086 | 2025-05-22 | 2025-05-26 | Cust 604 | DELHI | Analytical - CHAKAN | CLOSED | SALES PROMOTION | DO Transmitters | optex series(optex series) | Checking of DO sensor-4nos. SBR | 2 transmitter is not powered up. left 2 transmitters keypad not working. Suggested to send to pune for further troubleshooting purpose. | visited with sales person gautam sharma | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL25023112087 | 2025-05-22 | 2025-05-26 | Cust 786 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration and ER issue | Work done, RTU and Hart to modbus converter installed | Work done | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL25023352088 | 2025-05-22 | 2025-05-23 | Cust 840 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | DATCON | DATCON(DATCON) | Data is not uploading to CPCB server. | We have accessed the DATCon and upgraded the same with the latest version. Now the data is live on the CPCB server. | We have accessed the DATCon and upgraded the same with the latest version. Now the data is live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL25023732089 | 2025-05-23 | 2025-06-01 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | visit for ETP analyzers- 22 nos analyzers. | visit done as per attached MOM. | visit done as per attached mom. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL25023922090 | 2025-05-23 | 2025-05-24 | Cust 786 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 1st CAMC visit | 1st CAMC visit completed | Working fine | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL25025002091 | 2025-05-26 | 2025-05-29 | Cust 790 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to local and CPCB server. | Now the data multiparameter and flow is live on the CPCB server. As of now we are closing this complaint. We will reopen this once the camera issue is resolved from your end. | Now the data multiparameter and flow is live on the CPCB server. As of now we are closing this complaint. We will reopen this once the camera issue is resolved from your end. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL25025032092 | 2025-05-26 | 2025-05-28 | Cust 947 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera streaming issue on CPCB server. | Now the camera is streaming live on the CPCB server. | Now the camera is streaming live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL25025192093 | 2025-05-26 | 2025-05-28 | Cust 786 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | CPU with LCD of Multiparameter Display(1) | Display issue | Replaced the PCX1 and resolved the display issue | Working fine | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL25025382094 | 2025-05-26 | 2025-05-31 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | Visit for Oil in Water Analyzer to be commissioned in PNC, Panipat. | Commissioned Oil-In-Water Analyzers, Observed working satisfactorily. Details are in attached MOM. | Commissioned Oil-In-Water Analyzers, Observed working satisfactorily. Details are in attached MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL25025482095 | 2025-05-27 | 2025-05-28 | Cust 841 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Engineer visit required to check MPA not working. | Engineer viisted site and found following problems:- 1. Sample Pump not working properly. 2. Xenon Power Transformer not generating proper volatge 3. Display background light not working. MOM attached | Attended | NO | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFDL25025602096 | 2025-05-27 | 2025-05-28 | Cust 843 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | mA Output card for Multiparameter(AWA-4-20mA) | 1st AMC visit: Replacement of mA output card for COD/BOD | Replaced the mA output card and resolved the BOD output issue | Working fine | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL25026292097 | 2025-05-28 | 2025-05-30 | Cust 732 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB server. | We took access to the system and now data is live on CPCB server. | We took access to the system and now data is live on CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL25026542098 | 2025-05-28 | 2025-05-30 | Cust 803 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camer streaming issue on CPCB and RSPCB server. | Now the camera is live on the CPCB server. We have shared an RTSP link to customer for camera configuration on RSPCB server. | Now the camera is live on the CPCB server. We have shared an RTSP link to customer for camera configuration on RSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL25027042099 | 2025-05-29 | 2025-05-31 | Cust 812 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Scheduled AMC visit for Multi-Parameter Analyzers at ETP & STP Plant. | Visit for online Multi-Parameter Analyzers has been completed, observed both the Analyzers are working satisfactorily. Details are in attached MOM. | Visit for online Multi-Parameter Analyzers has been completed, observed both the Analyzers are working satisfactorily. Details are in attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL25027412100 | 2025-05-29 | 2025-06-01 | Cust 868 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | CPU with LCD of Multiparameter Display(1) | visit for installation of display card and cleaning pump of MPA | visit done for newly supplied display Unit installation as per attached MOM. | visit done as per attached mom. | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL25027602101 | 2025-05-30 | 2025-06-02 | Cust 964 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | RSPCB data upload issue. | Commercially not clear. | Commercially not clear. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL25027612102 | 2025-05-30 | 2025-06-02 | Cust 786 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | The Camera RTSP link needs to be generated for RSPCB | we have completed the Software installation + Local Connectivity, and the Data is live on the Local Enviroconnect server. | we have completed the Software installation + Local Connectivity, and the Data is live on the Local Enviroconnect server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL25029752103 | 2025-06-01 | 2025-06-06 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | DO Sensor Cable(DO Sensor Cable) | Demo Unit-Aqua2trans with Oxysens DO analyzer installation & commissioning for 2-3 months | visit done as per attached MOM. | visit done as per attached mom. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL25030022104 | 2025-06-02 | 2025-06-05 | Cust 967 | DELHI | DCS - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | MICROCON+ER SERIES | CONTROLLER MODULE (ER-201)(ER-201) | data mismatch issue on cpcb server. | As of now we are closing this complaint. We will reopen this once we get access to the system. | As of now we are closing this complaint. We will reopen this once we get access to the system. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL25030422105 | 2025-06-03 | 2025-06-07 | Cust 786 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | TSS reading is higher side due to Xenon power supply not working. | Replaced the new Xenon Power Supply Card, Now Analyzer is working satisfactorily. Details are in attached MOM. | Replaced the new Xenon Power Supply Card, Now Analyzer is working satisfactorily. Details are in attached MOM. | EXTENDED | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL25030702106 | 2025-06-03 | 2025-06-10 | Cust 938 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | PH Input Card for Multiparameter(AWA-PH) | All Parameters are showing incorrect values | Analyzer working Ok. Ph card need to be replaced. | Analyzer working Ok. Analyzer working Ok. Ph card need to be replaced. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFDL25030842107 | 2025-06-04 | 2025-06-06 | Cust 783 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software installation and CPCB connectivity needs to be done. | As of now we are closing this complaint and will reopen it once we get remote access. | As of now we are closing this complaint and will reopen it once we get remote access. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL25030982108 | 2025-06-04 | 2025-06-07 | Cust 786 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Schedule visit for ETP online Analyzer against CAMC. | Visit for online Multi-Parameter Analyzer has been completed against CAMC, observed Analyzer is working satisfactorily. Details are in attached MOM. | Visit for online Multi-Parameter Analyzer has been completed against CAMC, observed Analyzer is working satisfactorily. Details are in attached MOM. | EXTENDED | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL25031442109 | 2025-06-05 | 2025-06-06 | Cust 837 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | New Blackbox Configuration. | we have completed the Blackbox configuration, and now the Data is live on the CPCB server. | we have completed the Blackbox configuration, and now the Data is live on the CPCB server. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFDL25031942110 | 2025-06-06 | 2025-06-06 | Cust 799 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Blackbox Commissioning | we have completed the Blackbox configuration, and now the Data is live on the CPCB server through FM Blackbox | we have completed the Blackbox configuration, and now the Data is live on the CPCB server through FM Blackbox | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL25032182111 | 2025-06-07 | 2025-06-08 | Cust 844 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 3rd AMC visit | 3rd AMC visit completed | Working fine | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL25032272112 | 2025-06-07 | 2025-06-13 | Cust 968 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | commissioning | Commissioned | commissioned | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL25033062113 | 2025-06-10 | 2025-06-11 | Cust 790 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | data mismatch issue between analyzer and CPCB server | As of now we are closing this complaint. We will reopen this once the issue is resolved from the hardware side. | As of now we are closing this complaint. We will reopen this once the issue is resolved from the hardware side. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL25033252114 | 2025-06-10 | 2025-06-13 | Cust 841 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue. | We took access of the system and observed that the software database was crashed, hence we have now restored the Database and the data is now live on the CPCB server for ETP. | We took access of the system and observed that the software database was crashed, hence we have now restored the Database and the data is now live on the CPCB server for ETP. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL25033892115 | 2025-06-11 | 2025-06-15 | Cust 797 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Schedule AMC visit for ETP#3 to be done of online Analyzer against AMC. | Visit for COD, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | Visit for COD, Phenol & Sulphide Analysers against AMC has completed, Details are in attached MOM. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL25033912116 | 2025-06-11 | 2025-06-17 | Cust 898 | DELHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | TOC ANALYSER | TOC 1950 Single StreamANALYSER ANATEL(1950 TOC) | COMMISSIOING of AMIline TOC compact version | site is not ready, we will revisit after confirmation | site is not ready, we will revisit after confirmation | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL25034322117 | 2025-06-12 | 2025-06-13 | Cust 800 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data mismatch issue on CPCB server. | Now the analyzer and server values are matching and reflecting on the CPCB server. | Now the analyzer and server values are matching and reflecting on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL25034492118 | 2025-06-12 | 2025-06-14 | Cust 957 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | xenon card replacement & troubleshooting. | xenon power card replaced & taken in to service & found MPA working fine. | xenon power card replaced & taken in to service & found MPA working fine. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFDL25034622119 | 2025-06-13 | 2025-06-16 | Cust 812 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data break issue. | As of now, we are closing this complaint and will reopen this case once we receive remote access of the system from customer end. | As of now, we are closing this complaint and will reopen this case once we receive remote access of the system from customer end. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL25034752120 | 2025-06-13 | 2025-06-14 | Cust 844 | DELHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | Cleaning Pump(1) | Emergency visit for COD and TSS issue | Emergency visit completed | Working fine | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL25034762121 | 2025-06-13 | 2025-06-15 | Cust 807 | DELHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | checking | CHECKED CAL AND FOUND WORKING FINE. | CHECKED CAL AND FOUND WORKING FINE. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL25034882122 | 2025-06-13 | 2025-06-16 | Cust 841 | DELHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Future Data and time are uploading on the CPCB server. | We raise this issue with the AIPL team and now the issue has been resolved. | We raise this issue with the AIPL team and now the issue has been resolved. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL25035392123 | 2025-06-15 | 2025-06-17 | Cust 803 | DELHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | To check multiparameter analyzer and scanner | Multiparameter analyzer and scanner checked and resolved after necessary calibration | Working fine | NO | Paid | Core | N |  | Analytical - CHAKAN
AFDL25035442124 | 2025-06-15 |  | Cust 799 | DELHI | Analytical - CHAKAN | OPEN | AMC | MULTIPARAMETER ANALYSER | pH input card-FM of Multiparameter(1) | pH card to be replaced in Multi-Parameter Analyzer. |  |  | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFDL25035862125 | 2025-06-16 |  | Cust 823 | DELHI | Analytical - CHAKAN | ASSIGNED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Blackbox configuration for 2dust. |  |  | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFDL25036032126 | 2025-06-16 |  | Cust 837 | DELHI | Analytical - CHAKAN | ASSIGNED | DIGITAL CONEECTIVITY | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | connectivity |  |  | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFDL25036122127 | 2025-06-17 |  | Cust 913 | DELHI | Analytical - CHAKAN | ASSIGNED | QUALITY PROBLEM | MULTIPARAMETER ANALYSER | Peristaltic sampling pump motor for Multiparameter(AWA-PSPM) | PAPL reported on 22May24, Display comms issue and Sample pump motor gear wore out and damaged found. MPA commissioned on 31-May-24 Display comms issue has been resolved |  |  | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFDL25036372128 | 2025-06-17 |  | Cust 799 | DELHI | Analytical - CHAKAN | ASSIGNED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | HSPCB data upload issue. |  |  | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL25036382129 | 2025-06-17 |  | Cust 837 | DELHI | Analytical - CHAKAN | ASSIGNED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | HSPCB data upload issue. |  |  | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFDL25036652130 | 2025-06-17 |  | Cust 913 | DELHI | Analytical - CHAKAN | ASSIGNED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | CPU with LCD of Multiparameter Display(1) | comms issue hence visit make for replacement of cpu and hmi |  |  | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFHD23033712131 | 2023-06-01 | 2023-06-26 | Cust 969 | HYDERABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | Xenon Lamp for Multiparameter(AWA-XLAMP) | Commissioning | Problem Resolve Analyzer Under observation. | Problem resolve, Analyzer Under observation | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFHD23036692132 | 2023-06-08 | 2023-06-14 | Cust 970 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | Calibration | Calibration Completed | Attended | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD23036692132 | 2023-06-08 | 2023-06-14 | Cust 970 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | Calibration | Calibration Completed | Attended | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD23037052133 | 2023-06-09 | 2023-06-14 | Cust 971 | HYDERABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | Calibration | Calibration Completed | Attended | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD23038262134 | 2023-06-12 | 2023-06-16 | Cust 972 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration | Calibration | Attended | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD23038752135 | 2023-06-13 | 2023-06-16 | Cust 973 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Smartpro-con) | Calibration | Calibration Completed | Attended | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD23038752135 | 2023-06-13 | 2023-06-16 | Cust 973 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Smartpro-con) | Calibration | Calibration Completed | Attended | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD23041562136 | 2023-06-19 | 2023-06-24 | Cust 969 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | PH1 filter(1) | Calibration | Dear Sir find the attache of magnetic flow meter calibration certificate format Re-commissioning & Calibration Completed, Analyzer working is ok | Attended | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD23042292137 | 2023-06-19 | 2023-06-22 | Cust 974 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | APPCB data upload issue. | Data is live on APPCB server, hence we are closing this complaint. | Data is live on APPCB server, hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFHD23042612138 | 2023-06-20 | 2023-06-24 | Cust 975 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Re- Commissioning | Re- Commissioning Completed. | Attended | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD23043642139 | 2023-06-21 | 2023-06-26 | Cust 976 | HYDERABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4Trans(1) | commissioning | commissioning Completed under observation | Attended | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD23045322140 | 2023-06-24 | 2023-06-27 | Cust 977 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Localhost not opening in PC + CPCB data upload issue. | As of now we are closing this complaint, we will reopen it after getting the system access from customers end | As of now we are closing this complaint, we will reopen it after getting the system access from customers end | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFHD23047582141 | 2023-06-28 | 2023-06-30 | Cust 977 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Software reinstallation needs to be done. | As of now we are closing this complaint, we will reopen it after getting the correct camera IP address. | As of now we are closing this complaint, we will reopen it after getting the correct camera IP address. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFHD23048582142 | 2023-06-30 | 2023-07-03 | Cust 978 | HYDERABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | Commissioning | Commissioning Completed | Attended | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD23051462143 | 2023-07-05 | 2023-07-07 | Cust 979 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload + camera streaming issue | We took the access to your system and now data is live on the CPCB server and also camera is live on the server. | We took the access to your system and now data is live on the CPCB server and also camera is live on the server. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFHD23051482144 | 2023-07-05 | 2023-07-12 | Cust 952 | HYDERABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | Premier Phase-2 site. 2 Nos instruments display is not working | visit in next week after cust availability | visit in next week after cust availability | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD23051662145 | 2023-07-05 | 2023-07-12 | Cust 952 | HYDERABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Chlorine Analyzer-MAC 100 | MAC 100(1) | chloring analyzer commissioning. An urgent basis. Make Forbes Marshall MO A4TW055000 SR 23A4TO070 Site Engineer ? Rahul Kumar Con No 9760156474 | completed | attended | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD23054212146 | 2023-07-11 | 2023-07-12 | Cust 980 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB Data uploaded issue. | We took access of the system and made some changes in the configuration so now the data is live on the server. | Issue has been resolved and data is live hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFHD23054942147 | 2023-07-11 | 2023-07-15 | Cust 981 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Calibration | Calibration Completed | Attended | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD23057932148 | 2023-07-17 | 2023-07-20 | Cust 641 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | mA Output card for Multiparameter(AWA-4-20mA) | Calibration | Calibration Completed | Attended | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD23058442149 | 2023-07-18 | 2023-07-20 | Cust 982 | HYDERABAD | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Erection work | we recommend 24v dc power supply | attend | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD23066472150 | 2023-07-31 | 2023-08-05 | Cust 981 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Tranmitter Smartpro | Conductivity Electrode Cable(con-cab) | Calibration | Calibration Completed | Attended | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD23069502151 | 2023-08-07 | 2023-08-10 | Cust 603 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Delay data uploading onto the server. | We took the access to the system and successfully updated the RTDAS and now data is continuously uploading onto the server. | We took the access to the system and successfully updated the RTDAS and now data is continuously uploading onto the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFHD23070282152 | 2023-08-08 | 2023-08-19 | Cust 983 | HYDERABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4Trans(1) | commissioing of turbidity analyser | Work Completed | Work Completed | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFHD23070292153 | 2023-08-08 | 2023-08-19 | Cust 983 | HYDERABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | commissioning | Work Completed | Work Completed | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFHD23074182154 | 2023-08-17 | 2023-08-18 | Cust 984 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | software reinstallation and configuration. Note :- This will be the 2nd time installation under AMC. | here is an antivirus installed in PC which can cause enviroconnect software malfunction, so asked customer to arrange the other PC where there will not be any restrictions or antivirus . | here is an antivirus installed in PC which can cause enviroconnect software malfunction, so asked customer to arrange the other PC where there will not be any restrictions or antivirus . | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFHD23077522155 | 2023-08-23 | 2023-08-25 | Cust 984 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | need to do software reinstallation. - 2nd time. | We took access to the system and completed the software reinstallation and configuration. Data is now live on the server. | We took access to the system and completed the software reinstallation and configuration. Data is now live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFHD23078622156 | 2023-08-25 | 2023-09-01 | Cust 983 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | pH Transmitter Aquamon | Ph Transmitter Aquamon 4000(PH-Trans Aqu-4000) | calibration | resolved by phone call | resolved by phone call | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFHD23078802157 | 2023-08-25 | 2023-09-01 | Cust 985 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Tranmitter Smartpro | Conductivity Electrode Cable(con-cab) | calibration | Resolved Over Phone | Resolved Over Phone | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFHD23078972158 | 2023-08-25 | 2023-10-22 | Cust 986 | HYDERABAD | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | Ph Tranmitter Smartpro | Ph Electrode Cable(Ph-Cab) | Wrong Cable Supply | Resolved | Resolved | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFHD23081382159 | 2023-08-29 | 2023-09-04 | Cust 969 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | calibration | calibration | attended | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD23081832160 | 2023-08-30 | 2023-09-04 | Cust 983 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Aqua4Trans | Turbidity Input card of Aqua4Trans(1) | Calibration | Calibration Completed | Attended | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD23081832160 | 2023-08-30 | 2023-09-04 | Cust 983 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Calibration | Calibration Completed | Attended | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD23082982161 | 2023-09-01 | 2023-09-04 | Cust 987 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB configuration needs to be done. | We took access and completed the MPCB connectivity for the CPP boiler . Data is now live on the MPCB server. | We took access and completed the MPCB connectivity for the CPP boiler . Data is now live on the MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFHD23084232162 | 2023-09-04 | 2023-09-09 | Cust 982 | HYDERABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Commissioning | Commissioning Completed | Attended | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD23084562163 | 2023-09-05 | 2023-09-06 | Cust 975 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB camera streaming issue. | We took the access to the system and found that the camera streaming window is not open in local PC, after run the streaming window now camera is live on CPCB server. | We took the access to the system and found that the camera streaming window is not open in local PC, after run the streaming window now camera is live on CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFHD23085602164 | 2023-09-08 | 2023-09-12 | Cust 988 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | calibration | Suggested to spare | Attended | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD23088292165 | 2023-09-13 | 2023-09-16 | Cust 603 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | LCD of Multiparameter Display(1) | Calibration | Calibration Completed | Attended | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD23088812166 | 2023-09-14 | 2023-09-17 | Cust 972 | HYDERABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | MPA REPAIRED electronics comsssiong done | MPA REPAIRED electronics comsssiong done | MPA REPAIRED electronics comsssiong done | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFHD23091062167 | 2023-09-20 | 2023-09-26 | Cust 497 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Aqua4Trans | Aqua4Trans(1) | calibration | CALIBRATION | attended | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD23093332168 | 2023-09-25 | 2023-09-28 | Cust 977 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MySQL crashed need to do new software installation. | We have reinstalled the software and did the local + CPCB connectivity, but we are unable to run the services, so please check with your IT team and revert back to us. Hence as of now we are closing this complaint. We will reopen it once the password is recovered. | We reinstalled the software and did the local+CPCB connectivity, but we are unable to run the services Hence as of now, we are closing this complaint. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD23094692169 | 2023-09-27 | 2023-10-01 | Cust 989 | HYDERABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | Installation Guide line | Installation Guide line | Attended | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD23095472170 | 2023-09-28 | 2023-10-03 | Cust 952 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | pH Transmitter Aquamon | Ph Electrode Cable(PH- cab-4000) | pH Value Variation | pH Sensor ( KCL CONSUMPTION) | Attended | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD23098402171 | 2023-10-02 | 2023-10-05 | Cust 977 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating software Configuration. | Took access of the system and completed software reinstallation and configuration of LOCAL+CPCB+APPCB. | Took access of the system and done software reinstallation and configuration of LOCAL+CPCB+APPCB. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD23099672172 | 2023-10-04 | 2023-10-06 | Cust 618 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on LOCAL+MPCB+CPCB server. | We took access to your system and found that data is not live on modscan hence this issue needs to be checked from hardware side. | We took access to your system and found that data is not live on modscan hence this issue needs to be checked from hardware side. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD23102312173 | 2023-10-10 | 2023-10-16 | Cust 990 | HYDERABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | commissioning | sensor not working | attended | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD23105492174 | 2023-10-17 | 2023-10-22 | Cust 991 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration | Calibration Completed | Attended | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD23105502175 | 2023-10-17 | 2023-10-22 | Cust 992 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MIL card of Multiparameter(1) | Calibration | Calibration | Attended | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD23105932176 | 2023-10-17 | 2023-12-19 | Cust 993 | HYDERABAD | Analytical - CHAKAN | CLOSED | APPLICATION ISSUE | AQUADO Analyzer | Visiferm DO sensor(1) | commisisoning | material reached | material reached | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFHD23106472177 | 2023-10-18 | 2023-10-24 | Cust 994 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | calibration | INSPECTION | attended | NO | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD23106592178 | 2023-10-19 | 2023-10-24 | Cust 618 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | Flowcell disc of Multiparameter(1) | Calibration | Calibration Completed | Attended | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD23106872179 | 2023-10-19 | 2023-10-20 | Cust 618 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Local + CPCB + MPCB connectivity | We took access to the system and data is now live on Local + CPCB + MPCB server. | We took access to the system and data is now live on Local + CPCB + MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD23107452180 | 2023-10-20 | 2023-10-25 | Cust 995 | HYDERABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Conductivity Input card of Aqua4trans(1) | Commissioning | customer asked come in the next week | customer asked come in the next week | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFHD23107612181 | 2023-10-22 | 2023-10-25 | Cust 988 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | ORP transmitter(1) | calibration | resolved over phone | resolved over phone | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFHD23108552182 | 2023-10-25 | 2023-10-30 | Cust 603 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | calibration | calibration has been | attended | NO | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD23108642183 | 2023-10-25 | 2023-10-30 | Cust 995 | HYDERABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4Trans(1) | COMMISSIONING | COMMISSIONING COMPLETED | Attended | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD23109222184 | 2023-10-26 | 2023-10-27 | Cust 977 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | data not uploading to APPCB. | We took access to the system, and data from stack_1 is now reflecting on the APPCB server. | We took access to the system, and data from stack_1 is now reflecting on the APPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFHD23109252185 | 2023-10-26 | 2023-10-30 | Cust 996 | HYDERABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Display card of Aqua4Trans(1) | Commissioning | Commissioning Completed | Attended | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD23109252185 | 2023-10-26 | 2023-10-30 | Cust 996 | HYDERABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4Trans(1) | commissioning | commissioning Completed | Attended | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD23114412186 | 2023-11-01 | 2023-11-03 | Cust 992 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB and CPCB data not updating need to change the configuration from scanner to RS485. | Now the issue has been solved and data is live on both sites. | Now the issue has been solved and data is live on both sites. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD23114962187 | 2023-11-02 | 2023-11-02 | Cust 977 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Stack & Flow meters data not reflecting on APPCB | We took access to the system and data is now live on the APPCB server. | We took access to the system and data is now live on the APPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFHD23118502188 | 2023-11-10 | 2023-11-11 | Cust 994 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | data at analyzer and pc is not matching. | COD data mismatched issue has been resolved | COD data mismatched issue has been resolved | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFHD23118712189 | 2023-11-10 | 2023-11-10 | Cust 977 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Stack data is not reflecting to APPCB Site. | We have made some changes in configuration, now data is live on the APPCB server. | We have made some changes in configuration, now data is live on the APPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD23120422190 | 2023-11-16 | 2023-11-18 | Cust 613 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | APPCB data upload issue. | Did some changes in configuration end and now data is live on APPCB server. | Did some changes in configuration end and now data is live on APPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFHD23121122191 | 2023-11-18 | 2023-11-18 | Cust 613 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | APPCB data upload issue. | Now data is live on the APPCB server. | Now data is live on the APPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFHD23123052192 | 2023-11-22 | 2023-11-25 | Cust 997 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on CPCB Server. | Took access to the system and update camera data and camera both are live on CPCB. | Took access to the system and update camera data and camera both are live on CPCB. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD23123102193 | 2023-11-22 | 2023-11-26 | Cust 998 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration | calibration | attended | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD23123332194 | 2023-11-22 | 2023-11-27 | Cust 999 | HYDERABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | VISI pro DO(VISI pro DO) | commissioning | commissioning Completed | Attended | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD23123522195 | 2023-11-23 | 2023-11-23 | Cust 613 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | APPCB data upload issue | Data is now live on the APPCB server. | Data is now live on the APPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFHD23124392196 | 2023-11-25 | 2023-12-01 | Cust 1000 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | Calibration | Calibration Completed | Attended | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD23126232197 | 2023-11-27 | 2023-11-28 | Cust 1001 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data mismatch issue. | Changed the device ID from the analyzer side. Data is live on CPCB and MPCB. | Changed the device ID from the analyzer side. Data is live on CPCB and MPCB | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFHD23126252198 | 2023-11-27 | 2023-11-28 | Cust 613 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Stack monitor data not updating to APPCB server. | We have made some changes in configuration and data is continuously reflecting to the APPCB server. | We have made some changes in configuration and data is continuously reflecting to the APPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFHD23127012199 | 2023-11-28 | 2023-12-01 | Cust 635 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on CPCB Server. | Took access to the system and Camera is live on CPCB Server. | Took access to the system and Camera is live on CPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD23127022200 | 2023-11-28 | 2023-11-30 | Cust 1002 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on CPCB Server. | Took access to the system and both cameras are live on CPCB now. | Took access to the system and both cameras are live on CPCB now. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD23127252201 | 2023-11-28 | 2023-12-03 | Cust 1003 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | calibration | calibration has done | WORKING SATISFACTORY | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD23129942202 | 2023-12-01 | 2023-12-06 | Cust 1004 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | calibration | calibration has done | WORKING SATISFACTORY | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD23129982203 | 2023-12-01 | 2023-12-01 | Cust 1005 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | All the parameters are not uploading to APPCB server. | Data is now live on the APPCB server. | Data is now live on the APPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFHD23130012204 | 2023-12-01 | 2023-12-01 | Cust 462 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not Playing on RSPCB server. | Shared the rtsp link to the customer and told to taken up with RSPCB. | Shared the rtsp link to the customer and told to taken up with RSPCB. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD23131282205 | 2023-12-04 | 2023-12-05 | Cust 994 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not uploading to CPCB and MPCB | Data is live on CPCB and MPCB servers. | Data is live on CPCB and MPCB servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFHD23132062206 | 2023-12-06 | 2023-12-11 | Cust 1002 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Transmitter Aquacon | Conductivity Transmitter Aquamon 2000(CON-TR-2000) | Transmitter not working | Transmitter not working | Attended | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD23136162207 | 2023-12-14 | 2023-12-15 | Cust 618 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Stack data updating issue | Data is live. | Data is live. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFHD23138572208 | 2023-12-19 | 2023-12-24 | Cust 603 | HYDERABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | ph-1 value showing -ve | working satisfactory after maintenance | WORKING SATISFACTORY | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD23138762209 | 2023-12-19 | 2023-12-22 | Cust 603 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation needs to be done + CPCB + KSPCB | We took access to your system and did software reinstallation + CPCB + KSPCB configuration, and now data is live on both the servers. | We took access to your system and did software reinstallation + CPCB + KSPCB configuration, and now data is live on both the servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFHD23140262210 | 2023-12-22 | 2023-12-22 | Cust 1006 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | APPCB data upload issue. | Now data is live on APPCB server, hence we are closing this complaint. | Now data is live on APPCB server, hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFHD23142292211 | 2023-12-26 | 2023-12-29 | Cust 1007 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB+LOCAL | We took access of the system and now data is live on CPCB+LOCAL | We took access of the system and now data is live on CPCB+LOCAL | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD23142852212 | 2023-12-27 | 2023-12-29 | Cust 603 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera configuration needs to be done. | We have successfully completed the Camera connectivity on the CPCB server.Camera is streaming live on the CPCB portal. | We have successfully completed the Camera connectivity on the CPCB server.Camera is streaming live on the CPCB portal. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFHD23142872213 | 2023-12-27 | 2023-12-29 | Cust 1006 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | APPCB data upload issue. | we took access to your system and did Local + APPCB configuration for Totalizer Flow. And now Totalizer Flow data is live on Local + APPCB server. | we took access to your system and did Local + APPCB configuration for Totalizer Flow. And now Totalizer Flow data is live on Local + APPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFHD23143092214 | 2023-12-27 | 2024-01-09 | Cust 1002 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation | We have reinstalled the software and completed CPCB + MPCB connectivity and data is live on the same servers. | We have reinstalled the software and completed CPCB + MPCB connectivity and data is live on the same servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFHD23144322215 | 2023-12-29 | 2023-12-30 | Cust 603 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB + KSPCB data upload issue for Stack 2_75 TPH Boiler | As of now we are closing this complaint, and will reopen it after the hardware side issue is resolved from our side | As of now we are closing this complaint, and will reopen it after the hardware side issue is resolved from our side | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFHD23147622216 | 2024-01-03 | 2024-01-04 | Cust 613 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not updating to APPCB server. | Data is live on the APPCB server. | Data is live on the APPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFHD23148302217 | 2024-01-04 | 2024-01-05 | Cust 977 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not updating on CPCB and APPCB server. | Data is reflecting on the Local + CPCB + APPCB servers. | Data is reflecting on the Local + CPCB + APPCB servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFHD23148632218 | 2024-01-04 | 2024-01-09 | Cust 1008 | HYDERABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4Trans(1) | Rs485 Modbus Configuration | Rs485 Modbus Configuration | Attended | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD23149042219 | 2024-01-05 | 2024-01-09 | Cust 608 | HYDERABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4Trans(1) | commissioning | commissioning has been done conductivity showing satisfactory but PH electrode was not available on site | attended | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD23150522220 | 2024-01-08 | 2024-01-11 | Cust 497 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | APPCB data upload issue. | Now hardware side issue is resoled and data is live on APPCB server. | Now hardware side issue is resoled and data is live on APPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFHD23152562221 | 2024-01-11 | 2024-01-16 | Cust 1009 | HYDERABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4Trans(1) | HMI CONFIGRATION | HMI CONFIGRATION | Attended | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD23153562222 | 2024-01-15 | 2024-01-16 | Cust 984 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data mismatch issue. | We took the access to the system and now BOD & TSS data is matching to the actual value. | We took the access to the system and now BOD &amp;amp; TSS data is matching to the actual value. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFHD23154102223 | 2024-01-15 | 2024-01-16 | Cust 613 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | APPCB data upload issue. | we took the access to your system and now the issue has been resolved and data is live on the APPCB server. | we took the access to your system and now the issue has been resolved and data is live on the APPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFHD23158052224 | 2024-01-23 | 2024-01-25 | Cust 1007 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on MPCB+CPCB+LOCAL | We took access to the system and now data is updating on LOCAL+CPCB+MPCB | We took access to the system and now data is updating on LOCAL+CPCB+MPCB | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD23159222225 | 2024-01-25 | 2024-01-25 | Cust 979 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on CPCB Server | Took access to the system and cameras are live on CPCB Server for all 3 units. | Took access to the system and cameras are live on CPCB Server for all 3 units. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD23159852226 | 2024-01-28 | 2024-02-02 | Cust 1010 | HYDERABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | Visiferm DO sensor(1) | Commissioning Competed | Commissioning Competed | Attended | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD23160182227 | 2024-01-29 | 2024-02-01 | Cust 979 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on CPCB Server for Unit 1 ,Unit A and Unit B | Both the Camera are playing on Mozilla Firefox but due to the Chrome version issue both the cameras are not playing on Chrome hence as of now we are closing this complaint as cameras are playing properly. | Both the Camera are playing on Mozilla Firefox but due to the Chrome version issue both the cameras are not playing on Chrome hence as of now we are closing this complaint | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD23160492228 | 2024-01-29 | 2024-02-01 | Cust 1002 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data uploading issue to MPCB | Data is now live on the MPCB server. | Data is now live on the MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFHD23162562229 | 2024-01-31 | 2024-02-02 | Cust 977 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | software reinstallation | Data is now live on the APPCB and CPCB server. | Data is now live on the APPCB and CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFHD23170702230 | 2024-02-15 | 2024-02-16 | Cust 613 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Local host login issue | Local host is open now | Local host is open now | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFHD23170922231 | 2024-02-16 | 2024-02-16 | Cust 635 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Customer need CPCB data availability report OF 1 Year from 01-01-23 to 31-12-23. | We have provided customer the Data availability report of 1 Year from 01-01-23 to 31-12-23 | We have provided customer the Data availability report of 1 Year from 01-01-23 to 31-12-23 | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD23172442232 | 2024-02-19 | 2024-02-23 | Cust 979 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data and Camera was not updating on CPCB Portal on 16 Feb 2024 | We took access to the system and now camera and data is live on CPCB Server. | We took access to the system and now camera and data is live on CPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD23172672233 | 2024-02-19 | 2024-02-24 | Cust 1011 | HYDERABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | commissioning | Analyzer found faulty new one analyzer installed and found working satisfactory | attended | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD23174482234 | 2024-02-22 | 2024-02-24 | Cust 969 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration | calibration has been done | attended | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD23179032235 | 2024-02-28 | 2024-02-28 | Cust 972 | HYDERABAD | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | MULTIPARAMETER ANALYSER | Flowcell disc of Multiparameter(1) | TSS Value going above ranges during sampling time | Calibration has been done for the TSS Value and during sampling time PH sample has been checked and now system is working fine its range is under the limit | Calibration has been done for the TSS Value and during sampling time PH sample has been checked and now system is working fine its range is under the limit | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFHD23179992236 | 2024-02-29 | 2024-03-06 | Cust 534 | HYDERABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | VISI pro DO(VISI pro DO) | Commissioning | working satisfactory | WORKING SATISFACTORY | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD23181632237 | 2024-03-04 | 2024-03-07 | Cust 980 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB and SPCB server issue. | now the Flow data is live on the server. | now the Flow data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFHD23181802238 | 2024-03-04 | 2024-03-07 | Cust 979 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on CPCB Server. | Camera was not live on CPCB Server. We had taken up this issue with development team now camera is live on CPCB Server. | Camera was not live on CPCB Server. We had taken up this issue with development team now camera is live on CPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD23181812239 | 2024-03-04 | 2024-03-07 | Cust 979 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not updating on CPCB Server. | Camera was not live on CPCB Server. We had taken up this issue with development team now camera is live on CPCB Server. | Camera was not live on CPCB Server. We had taken up this issue with development team now camera is live on CPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD23182142240 | 2024-03-05 | 2024-03-07 | Cust 979 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not updating on CPCB Server | Camera is live on CPCB Server. There was ICECAST server issue which is resolved. | Camera is live on CPCB Server. There was ICECAST server issue which is resolved. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFHD23182642241 | 2024-03-06 | 2024-03-08 | Cust 991 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | data is not getting uploaded to cpcb central portal | Issue from central portal server. | Issue from central portal server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD23184202242 | 2024-03-11 | 2024-03-30 | Cust 1012 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | Supplied 3 no.s of tds transmitter (0-10micro siemens/cm) for ultra pure water of cell constant 0.01. Among these 2 was commissioned and working fine , one shows deviation while cross checking with the lab solution. Need to calibrate the transmitter against std solution below having 10 micro siemens/ cm. | Pavan savanur from hyd core visited the site and resolved the issue | Pavan savanur from hyd core visited the site and resolved the issue | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFHD23184252243 | 2024-03-11 | 2024-03-15 | Cust 1013 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | AQUADO Analyzer | Optex series DO sensor(1) | Calibration | calibration work done and few sensors are not ok to be replaced | attended | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD23185712244 | 2024-03-14 | 2024-03-15 | Cust 613 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not updating on Local + APPCB | Data is now live on the Local and APPCB servers. | Data is now live on the Local and APPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD23185872245 | 2024-03-14 | 2024-03-15 | Cust 979 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating and camera not playing on CPCB Server. | We took access to the system and camera is live on CPCB Server and there was pc time synch issue data is live now. | We took access to the system and camera is live on CPCB Server and there was pc time synch issue data is live now. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFHD23188122246 | 2024-03-18 | 2024-03-21 | Cust 1014 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not updating to CPCB server. | Data is live on the CPCB portal. | Data is live on the CPCB portal. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFHD23189222247 | 2024-03-19 | 2024-03-24 | Cust 1012 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Transmitter Aquacon | Conductivity Transmitter Aquamon 2000(COND-Trans Aqu-2000) | calibration | completed | attended | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD23189672248 | 2024-03-20 | 2024-03-22 | Cust 1015 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | software re-installation | As of now we are closing this complaint. we will reopen this once the issue is resolved from hardware side. | As of now we are closing this complaint. we will reopen this once the issue is resolved from hardware side. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD23192722249 | 2024-03-25 | 2024-03-29 | Cust 1014 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration | Calibration Completed | Attended | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD23193292250 | 2024-03-26 | 2024-03-29 | Cust 979 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Data not uploading on CPCB. | On 24.03.2024 2:00 am to 24.03.2024 18:00 for this date we did not found any logs, hence as of now we are closing this complaint. | On 24.03.2024 2:00 am to 24.03.2024 18:00 for this date we did not found any logs, hence as of now we are closing this complaint. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD23193452251 | 2024-03-26 | 2024-04-29 | Cust 972 | HYDERABAD | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | THE XENON POWER SUPPLY CARD & MUX CARD ISSUE DUE TO THIS READINGS ARE NOT GETTING PROPERLY | Xenon power supply card replaced with new one | Xenon power supply card replaced with new one | YES | No Trip | Process | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFHD23195832252 | 2024-03-29 | 2024-03-30 | Cust 1015 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not updating to CPCB server | Data is now live on the CPCB and Local servers. | Data is now live on the CPCB and Local servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24004632253 | 2024-04-10 | 2024-04-12 | Cust 979 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on CPCB Server. | Issue has been resolved. | Issue has been resolved. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24004642254 | 2024-04-10 | 2024-04-12 | Cust 979 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera is not live on CPCB Server. | Issue has been resolved. | Issue has been resolved. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24004662255 | 2024-04-10 | 2024-04-12 | Cust 979 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on CPCB Server. | Issue has been resolved. | Issue has been resolved. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24005442256 | 2024-04-12 | 2024-04-13 | Cust 603 | HYDERABAD | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | Flow cell for Multiparameter(AWA-FLOWCELL) | AMC | AMC has been done | attained | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24007572257 | 2024-04-18 | 2024-04-23 | Cust 919 | HYDERABAD | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | DO | guidelines has been given | attended | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24008622258 | 2024-04-20 | 2024-04-25 | Cust 919 | HYDERABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | AQUAMAX DOT(1) | commissioning | working satisfactory | attended | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24008622258 | 2024-04-20 | 2024-04-25 | Cust 919 | HYDERABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | AQUAMAX DOT(1) | commissioning | sensor is not working | attended | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24008622258 | 2024-04-20 | 2024-04-25 | Cust 919 | HYDERABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | AQUAMAX DOT(1) | commissioning | commissioning has been done working satisfactory | attended | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24009582259 | 2024-04-22 | 2024-05-27 | Cust 919 | HYDERABAD | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | AQUADO Analyzer | Visiwater DO sensor(1) | DO sensor not working. HMWSSB people are about to visit the for audit OA 5524108924 | sensor replaced working satisfactory | sensor replaced working satisfactory | YES | No Trip | Core | Y | Product / equipment failure due to which customers plant or part of plant is stopped which leaads to production loss. | Analytical - CHAKAN
AFHD24009752260 | 2024-04-22 | 2024-04-27 | Cust 1016 | HYDERABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Chlorine sensor | CHLORINE SENSOR(1) | sensor problem | checked sensor and found problem in sensor | SUGGEST TO SEND PUNE FOR REPAIR | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24013212261 | 2024-04-29 | 2024-04-30 | Cust 972 | HYDERABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | system performance after replacing the xenon power supply card | we have replaced the xenon power supply cad and observed that the system is working fine | we have replaced the xenon power supply cad and observed that the system is working fine | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFHD24015972262 | 2024-05-01 | 2024-05-03 | Cust 1017 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software re-installation | Issue has been resolved and the data is now live on the Local and CPCB servers. | Issue has been resolved and the data is now live on the Local and CPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24016542263 | 2024-05-03 | 2024-05-08 | Cust 1018 | HYDERABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | Dissolved Oxygen Analyzer Panel mounted(AQUADO4000P) | installation guidelines | Installation guidelines given | attended | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24019032264 | 2024-05-09 | 2024-05-10 | Cust 977 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | APPCB data upload issue | Data is now live on the APPCB server. | Data is now live on the APPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24019542265 | 2024-05-11 | 2024-05-16 | Cust 1019 | HYDERABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | AQUAMAX DOT(1) | commissioning | completed | attended | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24020462266 | 2024-05-13 | 2024-05-19 | Cust 1018 | HYDERABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | Visitrace DO sensor(1) | commissioning | commissioning has been done | WORKING SATISFACTORY | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24020462266 | 2024-05-13 | 2024-05-19 | Cust 1018 | HYDERABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | Visitrace DO sensor(1) | commissioning | commissioning has been done | WORKING SATISFACTORY | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24020462266 | 2024-05-13 | 2024-05-19 | Cust 1018 | HYDERABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | Visitrace DO sensor(1) | commissioning | commissioning has been done | WORKING SATISFACTORY | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24020462266 | 2024-05-13 | 2024-05-19 | Cust 1018 | HYDERABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | Visitrace DO sensor(1) | commissioning | commissioning has been done | WORKING SATISFACTORY | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24020472267 | 2024-05-13 | 2024-05-19 | Cust 1000 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | AQUADO Analyzer | Visitrace DO sensor(1) | Calibration | calibration has been done | WORKING SATISFACTORY | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24020732268 | 2024-05-14 | 2024-05-15 | Cust 641 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Issue for Device Status showing offline | Data is live on CPCB server. | Data is live on CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24020942269 | 2024-05-14 | 2024-05-19 | Cust 919 | HYDERABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | Visiwater DO sensor(1) | commissioning | commissioning has been done | WORKING SATISFACTORY | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24021792270 | 2024-05-16 | 2024-05-18 | Cust 641 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | On line CPCB server is not in connection with system | As per confirmation from customer there is server issue at their end. Once issue resolved customer will inform to reopen the complaint. This is as per telecommunication and email dtd 17/5/24 with customer. | As per confirmation from customer there is server issue at their end. Once issue resolved they will inform to reopen the complaint. This is as per telecommunication and email dtd 17/5 with customer | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24024702271 | 2024-05-22 | 2024-06-05 | Cust 641 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | CPCB server is not connected to the system | Data is live on CPCB + Local server. Also we have completed camera connectivity. | Data is live on CPCB + Local server. Also we have completed camera connectivity. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24026492272 | 2024-05-26 | 2024-05-29 | Cust 972 | HYDERABAD | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | MULTIPARAMETER ANALYSER | RS485 CARD for Multiparameter(AWA-RS485) | COD , BOD AND TSS data is not going on modbus | Modbus communication configuration has been completed and now data is being uploaded on PCB sever | Modbus communication configuration has been completed and now data is being uploaded on PCB sever | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFHD24030262273 | 2024-05-31 | 2024-06-09 | Cust 1019 | HYDERABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Polymetron Dissolved O2 Analyzer | Stainless steel 1.5 m probe extension(08878=C=1500) | commissioning | work resolution out of 4 sensors are working satisfactory | attained | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24032952274 | 2024-06-05 | 2024-06-09 | Cust 610 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Chlorine Analyzer-MAC 100 | MAC 100(1) | CALIB | CALIB DONE | attained | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24032972275 | 2024-06-05 | 2024-06-09 | Cust 610 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Turbidity Meter | Krohne turbidity transmitter(1) | CALIB | CALIB DONE | attained | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24033392276 | 2024-06-06 | 2024-06-09 | Cust 1018 | HYDERABAD | Analytical - CHAKAN | CLOSED | APPLICATION ISSUE | AQUADO Analyzer | Aquamax system(1) | TROUBLESHOOTING | Customer is satisfied with the readings | attained | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24037812277 | 2024-06-16 | 2024-06-21 | Cust 1018 | HYDERABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | AQUAMAX DOT(1) | commissioning | commissioned 2312269 , 2312271 and 2312270 | attended | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24038842278 | 2024-06-17 | 2024-06-21 | Cust 497 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Aqua4Trans | Aqua4Trans(1) | CALIBRATION | CALIBRATION | attained | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24038852279 | 2024-06-17 | 2024-06-21 | Cust 497 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Aqua4Trans | Aqua4Trans(1) | CALIBRATION | CALIB | attained | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24040532280 | 2024-06-21 | 2024-06-26 | Cust 1020 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Transmitter Aquacon | Conductivity Electrode Cable(CON-ELE-2000) | calibration | TRANSMITTER IS FAULTY | attained | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24041192281 | 2024-06-23 | 2024-06-25 | Cust 972 | HYDERABAD | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | MULTIPARAMETER ANALYSER | FM make RS485 card(1) | RS 485 modbus communication work, data is not showing on server | RS 485 modbus setting configuration has been done and data is showing on server. | RS 485 modbus setting configuration has been done and data is showing on server. | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFHD24041202282 | 2024-06-23 | 2024-06-25 | Cust 972 | HYDERABAD | Analytical - CHAKAN | CLOSED | APPLICATION ISSUE | MULTIPARAMETER ANALYSER | Cleaning Pump(1) | Acid pump not taking sample | Cleaning has been done and slowly pump is taking sample | Cleaning has been done and slowly pump is taking sample | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFHD24041292283 | 2024-06-23 | 2024-06-25 | Cust 1016 | HYDERABAD | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | Chlorine sensor | CHLORINE SENSOR(1) | Chlorine sensor installation work | Chlorine sensor installation and commissioning work has been done and values are within the range for MGF and ACF. | Chlorine sensor installation and commissioning work has been done and values are within the range for MGF and ACF. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFHD24043832284 | 2024-06-28 | 2024-07-01 | Cust 1014 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not uploading to PCB sites | Data is live on the CPCB ans TSPCB servers. | Data is live on the CPCB ans TSPCB servers. | YES | No Trip | Core | Y |  | DCS - CHAKAN
AFHD24044182285 | 2024-06-28 | 2024-07-03 | Cust 1021 | HYDERABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | inaccurate readings | informed over phone to calibrate the instrument & analyser is out of warranty. | Informed over phone to calibrate the instrument &amp;amp; Analyser is out of warranty. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFHD24047012286 | 2024-07-01 | 2024-07-05 | Cust 988 | HYDERABAD | Analytical - CHAKAN | CLOSED | DIGITAL CONEECTIVITY | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Site commissioning for MPA & Flow + digital connectivity needs to be done. | We have completed the Black Box configuration and now data is coming till the BB. | We have completed the Black Box configuration and now data is coming till the BB. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFHD24047242287 | 2024-07-01 | 2024-07-03 | Cust 1014 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Stack3 data is not updating to TSPCB server. | Data from Stack3 is now live on the TSPCB server. | Data from Stack3 is now live on the TSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24048262288 | 2024-07-03 | 2024-07-07 | Cust 973 | HYDERABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | COMMISSIONING | work as been done | attained | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24048282289 | 2024-07-03 | 2024-07-08 | Cust 973 | HYDERABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | Aquamax system(1) | COMMISSIONING | commissioning has been done | WORKING SATISFACTORY | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24049002290 | 2024-07-05 | 2024-07-05 | Cust 979 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Delay data not uploading on CPCB Server. | We took access to the system and found due to internet connectivity data was lost on 28th may but later delay data was updating correctly. | We took access to the system and found due to internet connectivity data was lost on 28th may but later delay data was updating correctly. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24050152291 | 2024-07-08 | 2024-07-12 | Cust 1019 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | sensor showing high value | sensor has been calibrated in air | WORKING SATISFACTORY | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24050922292 | 2024-07-09 | 2024-07-12 | Cust 1022 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data breaking issue. | We took access to the system and checked, data is not continuous on the modscan hence as of now we are closing this complaint and will reopen this once the issue is resolved from hardware side. Save | as of now we are closing this complaint and will reopen this once the issue is resolved from hardware side. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24052292293 | 2024-07-11 | 2024-07-16 | Cust 919 | HYDERABAD | Analytical - CHAKAN | CLOSED | APPLICATION ISSUE | AQUADO Analyzer | AQUAMAX DOT(1) | Troubleshooting | Found UPS is faulty, of company APC by Schneider Contacted customer care and booked a site visit of Schneider service engineer. | attained | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24052552294 | 2024-07-12 | 2024-07-19 | Cust 1023 | HYDERABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | AQUADO Analyzer | AQUAMAX DOT(1) | sensor is not working | WORK HAS BEEN COMPLETED | attained | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24052932295 | 2024-07-12 | 2024-07-19 | Cust 993 | HYDERABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Transmitter Aquacon | Aquacon 2000(AQCN 2000 p) | commissioning | FOUND SOME ERROR IN THE INSTRUMENT ,SUGGEST CUSTOMER TO SEND IT PUNE. | attained | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24053462296 | 2024-07-15 | 2024-07-19 | Cust 1002 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | calibration | CALIB HAS BEEN DONE. | attained | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24053692297 | 2024-07-15 | 2024-07-18 | Cust 1022 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Requirement of industry name and logo in station reports. | We have added company logo and industry name to the station reports as requested. | We have added company logo and industry name to the station reports as requested. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24055202298 | 2024-07-17 | 2024-07-18 | Cust 635 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on CPCB Server. | We took access to the system and now camera is live on CPCB server. | We took access to the system and now camera is live on CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24057912299 | 2024-07-22 | 2024-07-27 | Cust 1018 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | DO Transmitters | VISI pro DO(VISI pro DO) | calibration | due to high sludge value sensor showing o ppm after need to clean sensor for every 1 hr | attended | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24058092300 | 2024-07-23 | 2024-07-25 | Cust 977 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | connectivity of totalizer flow on APPCB server. | We have configured totalizer Flow on the APPCB server and the data is live for the same. | We have configured totalizer Flow on the APPCB server and the data is live for the same. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24068822301 | 2024-08-09 | 2024-08-13 | Cust 1020 | HYDERABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Transmitter Aquacon | Aquacon 2000(AQCN 2000 p) | COMMISSIONING | found problem in electronic suggested to send pune | sent for repair to pune | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24070762302 | 2024-08-13 | 2024-08-18 | Cust 781 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Aqua4Trans | Conductivity Input card of Aqua4trans(1) | CALIBRATION | CALIB HAS BEEN DONE | attained | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24071952303 | 2024-08-16 | 2024-08-17 | Cust 979 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB camera streaming issue. | Now the server side issue has been resolved, and the camera is streaming live on CPCB server. | Now the server side issue has been resolved, and the camera is streaming live on CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFHD24072832304 | 2024-08-19 | 2024-08-21 | Cust 1006 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | APPCB data uploading issue. | We took access to the system and now data is live on CPCB + Local + Common cloud server | We took access to the system and now data is live on CPCB + Local + Common cloud server | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24072982305 | 2024-08-19 | 2024-08-22 | Cust 979 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | camera steaming issue. | Now the issue has been resolved, and the camera streaming is live on CPCB server. | Now the issue has been resolved, and the camera streaming is live on CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24073002306 | 2024-08-19 | 2024-08-22 | Cust 1017 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to central portal | live on the central portal. | live on the central portal. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24073772307 | 2024-08-20 | 2024-08-25 | Cust 1024 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | CALIBRATION | calibration has been done | 202954 | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24074282308 | 2024-08-21 | 2024-08-23 | Cust 1024 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on APPCB server. And camera streaming issue on CPCB server. | The issue has been resolved now and the ETP Data is live on the APPCB server. | The issue has been resolved now and the ETP Data is live on the APPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24074782309 | 2024-08-22 | 2024-08-22 | Cust 1006 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on APPCB server. | We took access to the system and now data is live on APPCB server. | We took access to the system and now data is live on APPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24075322310 | 2024-08-23 | 2024-08-24 | Cust 1006 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on APPCB server. | We took access to the system and now data is live on APPCB server. | We took access to the system and now data is live on APPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24076742311 | 2024-08-26 | 2024-08-26 | Cust 1006 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on APPCB server. | We took access to the system and we have changed the output.jar file and the data is now live on the APPCB server. | We took access to the system and we have changed the output.jar file and the data is now live on the APPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24077212312 | 2024-08-26 | 2024-08-29 | Cust 1025 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data upload issue. | We took access of your system, and now the issue has been resolved and data is live on CPCB + MPCB server. | We took access of your system, and now the issue has been resolved and data is live on CPCB + MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFHD24077962313 | 2024-08-27 | 2024-08-30 | Cust 977 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | APPCB data uploading issue. | Stack data is now reflecting to the APPCB server. For Flowmeter issue we will reopen this complaint once we receive access to the system. | As of now we are closing this complaint. We will reopen this once we get access to your system. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24079222314 | 2024-08-29 | 2024-08-30 | Cust 1026 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | SPM data not updating on LOCAL Server. | We took access to the system and completed the installation and configuration of software after which data is live on LOCAL server. | We took access to the system and completed the installation and configuration of software after which data is live on LOCAL server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24083472315 | 2024-09-04 | 2024-09-06 | Cust 977 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Totalizer flow reading mismatch issue. | As of now we are closing this complaint. We will reopen this once we get access to your system. | As of now we are closing this complaint. We will reopen this once we get access to your system. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24084202316 | 2024-09-05 | 2024-09-06 | Cust 464 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Getting High value on CPCB server. | Discussed with customer Need to do changes from hardware end for decimal. | Discussed with customer Need to do changes from hardware end for decimal. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFHD24085372317 | 2024-09-08 | 2024-09-12 | Cust 1013 | HYDERABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | COMMISSIONING OF 4 NO'S OF DO (OXYFERM SENSOR) | commissioning has been done successfully tested in air and found satisfactory result | attended | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24085692318 | 2024-09-09 | 2024-09-09 | Cust 1006 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on APPCB . | Now data is live on APPCB server. | Now data is live on APPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24086872319 | 2024-09-11 | 2024-09-13 | Cust 977 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Flow Data mismatch issue | As of now we are closing this complaint. We will reopen this once we get access to your system. | As of now we are closing this complaint. We will reopen this once we get access to your system. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24087052320 | 2024-09-11 | 2024-09-16 | Cust 991 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration | Calibration has been done | attained | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24087602321 | 2024-09-12 | 2024-09-16 | Cust 641 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | 3-Way Electric Valve for Multiparameter(AWA-3WEV) | calibration | flow through chamber found contaminated | attended | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24089802322 | 2024-09-17 | 2024-09-22 | Cust 994 | HYDERABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | calibration | WORK HAS BEEN DONE | attained | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24090592323 | 2024-09-19 | 2024-09-24 | Cust 618 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | pH input card-FM of Multiparameter(1) | calibration | calibration has been done | attended | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24093122324 | 2024-09-24 | 2024-09-29 | Cust 1020 | HYDERABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Transmitter Aquacon | Conductivity Transmitter Aquamon 2000(COND-Trans Aqu-2000) | commissioning | commissioning has been done | WORKING SATISFACTORY | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24098962325 | 2024-10-04 | 2024-10-09 | Cust 992 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | TROUBLESHOOTING | display was not working after removed 4-20mA Card display start working. hence we suggested to change 4-20mA card | suggested for spare | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24100092326 | 2024-10-07 | 2024-10-10 | Cust 618 | HYDERABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning | WORK HAS BEEN DONE | attained | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24100452327 | 2024-10-08 | 2024-10-11 | Cust 618 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data uploading issue | We have made changes according to the RS485 connection and the data from ETP 1 is now live on the CPCB server. | We have made changes according to the RS485 connection and the data from ETP 1 is now live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24101162328 | 2024-10-09 | 2024-10-13 | Cust 603 | HYDERABAD | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | CALIBRATION | work has been done | attained | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24102532329 | 2024-10-13 | 2024-10-19 | Cust 641 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Tranmitter Smartpro | Conductivity Electrode Cable(Cond-Elec-cbl) | calibration | standard solution not be provided by customer so we do general maintenance. | attended | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24102532329 | 2024-10-13 | 2024-10-19 | Cust 641 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Tranmitter Smartpro | Conductivity Electrode Cable(Cond-Elec-cbl) | calibration | standard solution not be provided by customer so we do general maintenance. | attended | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24102532329 | 2024-10-13 | 2024-10-19 | Cust 641 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Tranmitter Smartpro | Conductivity Electrode Cable(Cond-Elec-cbl) | calibration | standard solution not be provided by customer so we do general maintenance. | attended | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24102532329 | 2024-10-13 | 2024-10-19 | Cust 641 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Tranmitter Smartpro | Conductivity Electrode Cable(Cond-Elec-cbl) | calibration | standard solution not be provided by customer so we do general maintenance. | attended | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24102532329 | 2024-10-13 | 2024-10-19 | Cust 641 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Tranmitter Smartpro | Conductivity Electrode Cable(Cond-Elec-cbl) | calibration | standard solution not be provided by customer so we do general maintenance. | attended | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24102532329 | 2024-10-13 | 2024-10-19 | Cust 641 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Tranmitter Smartpro | Conductivity Electrode Cable(Cond-Elec-cbl) | calibration | standard solution not be provided by customer so we do general maintenance. | attended | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24104382330 | 2024-10-17 | 2024-10-22 | Cust 1000 | HYDERABAD | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | DO Transmitters | VISI pro DO(VISI pro DO) | do SENSOR membrane problem | sensor not working due to membrane failure | SUGGEST TO SEND PUNE FOR REPAIR | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24106022331 | 2024-10-20 | 2024-10-25 | Cust 603 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | CALIBRATION | CALIBRATION HAS BEEN DONE PH CARD FOUND FAULTY | suggested for spare | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24106642332 | 2024-10-21 | 2024-10-24 | Cust 641 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | flow meter showing error reading . | This issue is related to hardware, so we are closing this complaint. | This issue is related to hardware, so we are closing this complaint. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24108842333 | 2024-10-25 | 2024-11-01 | Cust 1027 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Electrode Cable(Ph-Cab) | calibration | Result Found Satisfactory after calibration | Result Found Satisfactory after Calibration | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFHD24108852334 | 2024-10-25 | 2024-11-01 | Cust 1027 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Transmitter Aquacon | Aquacon 2000(AQCN 2000 p) | calibration issue | Result found Satisfactory after calibration | Result Found Satisfactory after Calibration | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFHD24108852334 | 2024-10-25 | 2024-11-01 | Cust 1027 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | calibration issue | Result found satisfactory after Calibration | Result Found Satisfactory after Calibration | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFHD24113742335 | 2024-11-06 | 2024-11-08 | Cust 979 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | CPCB camera streaming issue. | We took access to your system and now the issue has been resolved and the camera is streaming live on CPCB server. | We took access to your system and now the issue has been resolved and the camera is streaming live on CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFHD24115332336 | 2024-11-10 | 2024-11-15 | Cust 1028 | HYDERABAD | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC | Routine maintenance carried out | Routine maintenance carried out | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFHD24116902337 | 2024-11-13 | 2024-11-18 | Cust 1000 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | calibration | Calibration has been done | attained | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24117562338 | 2024-11-13 | 2024-11-18 | Cust 969 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Silica Analyser | Silica 8891(8891) | calibration | Calibration completed successfully | Calibration Completed successfully | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFHD24118152339 | 2024-11-15 | 2024-11-18 | Cust 1029 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | EnviroConnect software re-installation. | We accessed the system and successfully re-installed the EnviroConnect software. However, we noticed that the data is not coming till MODSCAN. So this issue is from the Hardware side.so we are closing this complaint. | We accessed the system and successfully re-installed the EnviroConnect software. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24118292340 | 2024-11-15 | 2024-11-15 | Cust 975 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation + CPCB configuration needs to be done. | we took access to your system and completed software reinstallation + CPCB configuration, now data is live on CPCB server. | we took access to your system and completed software reinstallation + CPCB configuration, now data is live on CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFHD24119102341 | 2024-11-18 | 2024-11-19 | Cust 991 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to CPCB and MPCB server. | Timesync issue has been resolved and data is now live on the CPCB and MPCB server. | Timesync issue has been resolved and data is now live on the CPCB and MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24119562342 | 2024-11-18 | 2024-11-23 | Cust 993 | HYDERABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | ORP transmitter(1) | commissioning | Work completed | work completed | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFHD24119572343 | 2024-11-18 | 2024-11-23 | Cust 1030 | HYDERABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | pH Transmitter Aquamon | Aquamon 2000 plus(AQ2000 plus) | commissioning | Work Completed | work completed | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFHD24120822344 | 2024-11-20 | 2024-11-25 | Cust 1031 | HYDERABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Transmitter Aquacon | Aquacon 2000(AQCN 2000 p) | Commissioning | work has been done | Attained | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24120832345 | 2024-11-20 | 2024-11-25 | Cust 1031 | HYDERABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | pH Transmitter Aquamon | Aquamon 2000 plus(AQ2000 plus) | commissioning | done | Attained | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24120862346 | 2024-11-20 | 2024-11-25 | Cust 1032 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Transmitter Aquacon | Conductivity Transmitter Aquamon 2000(CON-TR-2000) | CALL | work has been done ,directly by the phone call & calibration has been done for the pH meter. | Work has been done , directly from the phone support. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFHD24121052347 | 2024-11-21 | 2024-11-22 | Cust 613 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | APPCB data upload issue | We took access to your system and checked, currently data is live on the APPCB server. | We took access to your system and checked, currently data is live on the APPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24123342348 | 2024-11-26 | 2024-11-29 | Cust 984 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | EnviroConnect software re-installation. | We have not received the AMC PO details for remote support, so we are closing this complaint. | We have not received the AMC PO details for remote support, so we are closing this complaint. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24123362349 | 2024-11-26 | 2024-11-27 | Cust 975 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | camera configuration needs to be done | today we took access to your system and completed Camera configuration, now the camera is streaming live on CPCB server. | today we took access to your system and completed Camera configuration, now the camera is streaming live on CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFHD24123442350 | 2024-11-26 | 2024-11-28 | Cust 613 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | APPCB data uploading issue. | We have checked and data is live on the Local and APPCB servers. | We have checked and data is live on the Local and APPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24123842351 | 2024-11-27 | 2024-12-02 | Cust 603 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | PH1 detector(1) | analyzer showing Err-8 | PH1 and PH2 are hunting, suggest customer to procure new set of it. | PH1 and PH2 are hunting , suggest customer to procure new set of it. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFHD24123862352 | 2024-11-27 | 2024-11-29 | Cust 603 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data uploded issue. | Now the data is live on the MPCB server. | Now the data is live on the MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFHD24123932353 | 2024-11-27 | 2024-11-27 | Cust 618 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on CPCB + Local + MPCB | We have accessed the system and now the data is live on CPCB + Local + MPCB servers. | We have accessed the system and now the data is live on CPCB + Local + MPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24124112354 | 2024-11-27 | 2024-11-29 | Cust 635 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Need to check Camera is live on CPCB server or not. | We checked and camera is live on CPCB Server. | We checked and camera is live on CPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24125682355 | 2024-11-29 | 2024-11-29 | Cust 603 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | The data is live on the KSPCB server for Stack 2_75 TPH Boiler, Stack 1_140 TPH Boiler, and ETP 1_ETP_Outlet. However, when generating the data report, only the values for Stack 2_75 TPH Boiler are displayed. The values for Stack 1_140 TPH Boiler and ETP 1_ETP_Outlet are not appearing in the report | The issue has now been resolved, and the values for Stack 1_140 TPH Boiler and ETP 1_ETP_Outlet have been successfully generated in the data report. | The issue has now been resolved, and the values for Stack 1_140 TPH Boiler and ETP 1_ETP_Outlet have been successfully generated in the data report. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24125712356 | 2024-11-29 | 2024-11-29 | Cust 1002 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on LOCAL+CPCB+MPCB server | We took access to the system and changed the COM port and now data is live on LOCAL+CPCB+MPCB server. | We took access to the system and changed the COM port and now data is live on LOCAL+CPCB+MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24125752357 | 2024-11-29 | 2024-12-02 | Cust 603 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera streaming issue on CPCB server. | We are closing this complaint for now.we will reopen it once we get administrator access to the system. | We are closing this complaint for now.we will reopen it once we get administrator access to the system. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24125982358 | 2024-11-29 | 2024-11-29 | Cust 1033 | HYDERABAD | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | Aqua4Trans | Conductivity Input card of Aqua4trans(1) | CCDS Syatem is supplied in 2019 now customer is planning to commission | Most of the euipments are missed they have to purchase we have in touch with the division for req spares | Most of the euipments are missed they have to purchase we have in touch with the division for req spares | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFHD24128402359 | 2024-12-02 | 2024-12-05 | Cust 1014 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to Local CPCB and TSPCB | There was an issue from the hardware side as the data was not being transmitted till the modscan, which caused the data to be unavailable locally and the TSPCB as well. After replaced the scanner, the issue was resolved and data is coming on modscan. Also we have updated the Enviroconnect DAS version to 4.6. | After replaced the scanner, the issue was resolved and data is coming on modscan. Also we have updated the Enviroconnect DAS version to 4.6. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24128442360 | 2024-12-02 | 2024-12-05 | Cust 975 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | APPCB data upload issue. | we took access to your system and completed APPCB connectivity, now data is live on the APPCB server. | we took access to your system and completed APPCB connectivity, now data is live on the APPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFHD24128802361 | 2024-12-03 | 2024-12-06 | Cust 1034 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not live on CPCB server. | data is now live on the CPCB server. | data is now live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24129072362 | 2024-12-03 | 2024-12-08 | Cust 1011 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | AQUADO Analyzer | Aquamax system(1) | calibration | SUGGESTED CUSTOMER TO SEND ANALZER TO PUNE | SUGGEST TO SEND PUNE FOR REPAIR | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24129072362 | 2024-12-03 | 2024-12-08 | Cust 1011 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | AQUADO Analyzer | Aquamax system(1) | calibration | SUGGESTED TO SEND PUNE FOR FURTHER ANALYSIS | SUGGEST TO SEND PUNE FOR REPAIR | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24129172363 | 2024-12-04 | 2025-01-13 | Cust 603 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera configuration needs to be done | We have accessed the system and completed the camera configuration | We have accessed the system and completed the camera configuration | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24129802364 | 2024-12-05 | 2024-12-06 | Cust 603 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on KSPCB server | Now the correct readings are uploading on the KSPCB server for flow parameter. | Now the correct readings are uploading on the KSPCB server for flow parameter. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24131252365 | 2024-12-08 | 2024-12-14 | Cust 1018 | HYDERABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | OXYFERM(OXYFERM) | commissioning | commissioning has been done | attended | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24131742366 | 2024-12-09 | 2024-12-10 | Cust 603 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Values are not showing up when generating the data report for Distillery to Incineration Boiler Flow Spent Wash Outlet to Compost Yard Flow | The issue has now been resolved, and the values for Distillery to Incineration Boiler Flow and Spent Wash Outlet to Compost Yard Flow have been successfully generated in the data report. | The issue has now been resolved, and the values for Distillery to Incineration Boiler Flow and Spent Wash Outlet to Compost Yard Flow have been successfully generated in the data report. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24132992367 | 2024-12-12 | 2024-12-13 | Cust 975 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | APPCB data upload issue. | we took access to your system, and now the issue has been resolved and data is live on the APPCB server. | we took access to your system, and now the issue has been resolved and data is live on the APPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFHD24133212368 | 2024-12-12 | 2025-02-28 | Cust 564 | HYDERABAD | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | Conductivity Electrode | Forbes Marshall Cell Model 8312, K=1(FM -8312) | 02 Nos. Conductivity Transmitter Display not showing any reading. We have checked the Transmitter Power 24 VDC and also cross check the display with another transmitter but problem identified that the Transmitter is failure. Model: SMARTPRO 8967 Sr. No. - 23SCI815 & 23SCI818 01 No. Conductivity Transmitter not showing any reading after connecting sensor. We have cross checked with another sensor but still not giving response. Model: SMARTPRO 8967 Sr. No. - 23SCI817 Balance 04 nos. pH and 02 nos. ORP transmitter yet to taken inline. Vendor PO No. 225592 Dated: 11-JULY-2023 It is not possible to send this Analyzer Transmitter to Forbes Marshall workshop at PUNE. Hence requesting you to please depute their service engineer with necessary spares at site and rectify the problem. Site Address: TGV SRAAC LTD. Village: Gondiparla, District-Kurnool, Andhra Pradesh Site Contact Person : Vijay Koshe(7999597079), A. Selvamani(9597567606) | Material received in Depot vide complaint number CCS24D1720 repaired & Dispatched on 30th Jan 2025 | Material received in Depot vide complaint number CCS24D1720 repaired &amp;amp;amp; Dispatched on 30th Jan 2025 | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFHD24134482369 | 2024-12-16 | 2024-12-17 | Cust 984 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | EnviroConnect Software re-installation. | We accessed the system and successfully re-installed the EnviroConnect software, as well as completed the CPCB + Local Server configuration. And now the data is live. | We accessed the system and successfully re-installed the EnviroConnect software, as well as completed the CPCB + Local Server configuration. And now the data is live. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24134612370 | 2024-12-16 | 2024-12-20 | Cust 603 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Flow data issue on KSPCB server. cpcb and KSPCB Servers flow data mismatch | We have noticed that the PC is freezing, so this issue is from customer side. | We have noticed that the PC is freezing, so this issue is from customer side. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24134692371 | 2024-12-16 | 2024-12-19 | Cust 1014 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue | we have checked and data is showing live on CPCB and TGPCB server. | we have checked and data is showing live on CPCB and TGPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24136482372 | 2024-12-19 | 2024-12-24 | Cust 998 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | calibration | calibration has been done but found err-8 some time so suggested for spare | suggested for spare | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24138062373 | 2024-12-22 | 2024-12-27 | Cust 497 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(Ph-Trans-Smartpro) | Calib | CALIBRATION HAS BEEN DONE | attended | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24139072374 | 2024-12-24 | 2024-12-26 | Cust 984 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on CPCB server. | As per telephonic conversation we took access to the system and now the data is live on the CPCB Server. | As per telephonic conversation we took access to the system and now the data is live on the CPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24139342375 | 2024-12-24 | 2024-12-29 | Cust 528 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Electrode | CONDUCELL 2DC-PG-120(HT237 610) | calibration | SENSOR NOT RESPONDING TO THE CLO2 | under observation | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24139512376 | 2024-12-25 | 2024-12-29 | Cust 1035 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MUVX Card of Multiparameter(1) | troubleshooting | XEON lamp is not working suggested to send pune for repair | SUGGEST TO SEND PUNE FOR REPAIR | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24143312377 | 2024-12-31 | 2024-12-31 | Cust 972 | HYDERABAD | Analytical - CHAKAN | CLOSED | APPLICATION ISSUE | MULTIPARAMETER ANALYSER | CPU Display card for Multiparameter(AWA-CPU1) | MPA Display controller replacement | required configuration has been done and data is live on PCB. | required configuration has been done and data is live on PCB. | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFHD24143322378 | 2024-12-31 | 2024-12-31 | Cust 972 | HYDERABAD | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | MULTIPARAMETER ANALYSER | CPU Display card for Multiparameter(AWA-CPU1) | new display installation and commissioning | New display controller replaced and commissioning activities are completed system working properly | New display controller replaced and commissioning activities are completed system working properly | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFHD24143452379 | 2024-12-31 | 2025-01-02 | Cust 975 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | APPCB data upload issue. | we took access to your system and now the issue has been resolved and data is live on the APPCB server also Camera is streaming live on CPCB server. | we took access to your system and now the issue has been resolved and data is live on the APPCB server also Camera is streaming live on CPCB server. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFHD24145992380 | 2025-01-03 | 2025-01-08 | Cust 603 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | PH1 detector(1) | Troubleshooting | ph card replaced and calibrated found working ok | attended | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24146012381 | 2025-01-03 | 2025-01-03 | Cust 979 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera streaming issue on CPCB server. | The camera streaming is now live on the CPCB server for Unit-1, Unit-B and Unit-C. | The camera streaming is now live on the CPCB server for Unit-1, Unit-B and Unit-C. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24146612382 | 2025-01-05 | 2025-01-11 | Cust 972 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | display touch problem | found 10 mm difference B/W display and touchscreen issue resolved | WORKING SATISFACTORY | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24147982383 | 2025-01-08 | 2025-01-09 | Cust 992 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Values showing zero on CPCB + Local + MPCB servers for ETP Outlet. | We have accessed the system, and the values are now being uploaded to the CPCB, Local, and MPCB servers for the ETP outlet. As discussed during the call with customer , there is an issue with the DDU for Stack 1_Boiler 135 TPH. customer will be sending it to our Chakan factory for repair. | We have accessed the system, and the values are now being uploaded to the CPCB, Local, and MPCB servers for the ETP outlet. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24148702384 | 2025-01-09 | 2025-01-10 | Cust 992 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on MPCB server for ETP Outlet. | As discussed during the call with customer , we have accessed the system, and resolved the issue . The data is now live on the MPCB server for the ETP Outlet. | As discussed during the call with customer , we have accessed the system, and resolved the issue . The data is now live on the MPCB server for the ETP Outlet. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24150452385 | 2025-01-14 | 2025-01-16 | Cust 992 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Re-installation to the latest version of the EnviroConnect software, | As discussed during the call with customer , we accessed the system, re-installed the latest version of the EnviroConnect software, and completed the configuration for the CPCB, Local, and MPCB servers.The data is now live | We accessed the system, re-installed the latest version of the EnviroConnect software, and completed the configuration for the CPCB, Local, and MPCB servers.The data is now live | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24151662386 | 2025-01-16 | 2025-01-19 | Cust 603 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Mismatch data for flow indication & totalizer on KSPCB and CPCB servers. | Waiting confirmation from Branch Engg. Pavan sir for PO and cost of PO. Once getting confirmation will reopen the complaint. Mail attached for reference | Waiting confirmation from Branch Engg. Pavan sir for PO and cost of PO. Once getting confirmation will reopen the complaint. Mail attached for reference | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24152022387 | 2025-01-17 | 2025-01-23 | Cust 1036 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | calibration | found error during calibration suggested to send pune for further analysis and repair | SUGGEST TO SEND PUNE FOR REPAIR | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24153322388 | 2025-01-20 | 2025-01-23 | Cust 988 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data upload issue. | we took access to your system and now the issue has been resolved and data is live on the MPCB server for all flow meters. | we took access to your system and now the issue has been resolved and data is live on the MPCB server for all flow meters. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFHD24154612389 | 2025-01-22 | 2025-01-24 | Cust 1026 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | re-installation of EnviroConnect software and local + APPCB connectivity needs to be done. | We have accessed the system, reinstalled the EnviroConnect software, and completed the configuration for both the Local and APPCB servers. The data is now live on the Local server. However, we have observed that the APPCB server is showing a "site shutdown" message. Please remove the workflow Once the workflow is removed, the data will start uploading to the APPCB server. | We have accessed the system, reinstalled the EnviroConnect software, and completed the configuration for both the Local and APPCB servers. The data is now live on the Local server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24154952390 | 2025-01-23 | 2025-01-24 | Cust 613 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | APPCB data upload issue | data reflecting on local server and APPCB site also | data reflecting on local server and APPCB site also | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24156882391 | 2025-01-27 | 2025-02-02 | Cust 1037 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Calibration Issue | Issue resolved | Problem resolved | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFHD24157772392 | 2025-01-28 | 2025-02-06 | Cust 1037 | HYDERABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | calibration | transmitters found not working. customer to send all the transmitters and sensors to cust | transmitters found not working. customer to send all the transmitters and sensors to cust | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFHD24160092393 | 2025-01-30 | 2025-01-31 | Cust 603 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Mismatch data for flow & flow totalizer on CPCB + KSPCB server. | The issue has been resolved, and the flow and flow totalizer values for the device ETP 1_ETP_Outlet are now uploading correctly to the KSPCB server. | The issue has been resolved, and the flow and flow totalizer values for the device ETP 1_ETP_Outlet are now uploading correctly to the KSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24160152394 | 2025-01-30 | 2025-02-01 | Cust 979 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on CPCB server for unit-1 | There was server side issue we checked and now camera is live on CPCB server. | There was server side issue we checked and now camera is live on CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24162622395 | 2025-02-03 | 2025-02-10 | Cust 1038 | HYDERABAD | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Smartpro-con) | Erection Work | resolved over phone | resolved over phone | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFHD24162642396 | 2025-02-03 | 2025-02-10 | Cust 1038 | HYDERABAD | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Erection work | resolved over phone | resolved over phone | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFHD24163092397 | 2025-02-04 | 2025-02-07 | Cust 979 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera streaming issue on CPCB server | We have accessed the system, and the issue has now been resolved. The camera streaming is live on the CPCB server. | We have accessed the system, and the issue has now been resolved. The camera streaming is live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24163772398 | 2025-02-05 | 2025-02-10 | Cust 988 | HYDERABAD | Swas - CHAKAN | CLOSED | CALIBRATION JOB | SWAS & Analysers | Steam & Water Analysis System(SWAS & Analysers) | spare replacement and calibration | Services for all analyser has been completed | Attained | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24163782399 | 2025-02-05 | 2025-02-07 | Cust 613 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software re-installation | we installed the software in a new PC and completed Local + APPCB server configuration. Data is live on the same servers. | we installed the software in a new PC and completed Local + APPCB server configuration. Data is live on the same servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24164012400 | 2025-02-05 | 2025-02-07 | Cust 613 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software re-installation + Local and APPCB connectivity | We have re-installed the software with an upgraded version and data is now live on the Local and APPCB server. | We have re-installed the software with an upgraded version and data is now live on the Local and APPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24164622401 | 2025-02-07 | 2025-02-10 | Cust 988 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | 12 V POWER SUPPLY CARD for Multiparameter(AWA-PS1) | Calibration | Service and Calibration for MPA analyzer has been done | Attained | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24165752402 | 2025-02-10 | 2025-02-16 | Cust 969 | HYDERABAD | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration | Calibration done | Attained | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24165772403 | 2025-02-10 | 2025-02-14 | Cust 988 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | AQUADO Analyzer | OXYSENS(OXYSENS) | Calibration Issue | Calibration Done | Attained | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24166752404 | 2025-02-12 | 2025-02-15 | Cust 979 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Camera streaming issue on the CPCB server. | There was an issue with the Icecast streaming server, which the AIPL team resolved. Now, the Camera streaming is live on the CPCB server. | There was an issue with the Icecast streaming server, which the AIPL team resolved. Now, the Camera streaming is live on the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFHD24167392405 | 2025-02-13 | 2025-02-15 | Cust 980 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data uploaded issue. | The Flow data and Camera streaming are now live on the CPCB server. | The Flow data and Camera streaming are now live on the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFHD24167752406 | 2025-02-14 | 2025-02-17 | Cust 603 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data mismatch issue on KSPCB server for Spent Wash Outlet to Compost Yard_Flow. | The issue has been resolved, and the flow values for the Spent Wash Outlet to Compost Yard_Flow are now uploading correctly to the KSPCB server. | The issue has been resolved, and the flow values for the Spent Wash Outlet to Compost Yard_Flow are now uploading correctly to the KSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24169232407 | 2025-02-18 | 2025-02-23 | Cust 1019 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | calibration | completed | attended | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24170222408 | 2025-02-20 | 2025-02-21 | Cust 1014 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | data is not uploading to CPCB server. | Data is now showing live on the CPCB & TGPCB servers. | Data is now showing live on the CPCB &amp;amp; TGPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24170732409 | 2025-02-21 | 2025-02-25 | Cust 603 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration | Carried out Zero calibration | Zero calibration done. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFHD24177842410 | 2025-03-05 | 2025-03-10 | Cust 990 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | AQUADO Analyzer | OXYFERM(OXYFERM) | Calibration | calibration done | attained | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24178532411 | 2025-03-07 | 2025-03-10 | Cust 1017 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | new Windows -11 system as a replacement for the old CEMS system | we have completed the software installation + CPCB connectivity on the new PC and Data is now live on the CPCB server. | we have completed the software installation + CPCB connectivity on the new PC and Data is now live on the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFHD24180942412 | 2025-03-11 | 2025-03-23 | Cust 972 | HYDERABAD | Analytical - CHAKAN | CLOSED | APPLICATION ISSUE | MULTIPARAMETER ANALYSER | pH input card-FM of Multiparameter(1) | For calibration of TSS cod bod | Visited the site and calibration has been done for COD BOD & TSS | Visited the site and calibration has been done for COD BOD &amp;amp; TSS | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFHD24183282413 | 2025-03-18 | 2025-03-21 | Cust 979 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on CPCB server. | We took access to the system and observed camera streaming window was closed. Now camera is live for all units on CPCB server. | We took access to the system and observed camera streaming window was closed. Now camera is live for all units on CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24184202414 | 2025-03-20 | 2025-03-24 | Cust 641 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | Flat cable of Multiparameter Display(1) | Calibration | Calibration performed | Attained | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD24184512415 | 2025-03-20 | 2025-03-21 | Cust 1002 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on LOCAL+CPCB+MPCB server. | We took access to the system and observed there was time sync issue which has been resolved and now data is live on LOCAL+CPCB+MPCB server. | We took access to the system and observed there was time sync issue which has been resolved and now data is live on LOCAL+CPCB+MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD24184542416 | 2025-03-20 | 2025-03-21 | Cust 975 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | APPCB data upload issue. | We took access of the system and now the issue has been resolved and data is live on the APPCB server. (Trunked SQL) | We took access of the system and now the issue has been resolved and data is live on the APPCB server. (Trunked SQL) | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFHD24185162417 | 2025-03-22 | 2025-03-27 | Cust 793 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | AQUADO Analyzer | AQUAMAX DOT(1) | Transmitter Showing Fail Alarm | Transmitter faulty | Checked on phone call and found Transmitter faulty , suggested to send to pune for repair | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFHD25000842418 | 2025-04-03 | 2025-04-07 | Cust 1001 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | MULTIPARAMETER ANALYSER | FM make RS485 card(1) | RS 485 Output Issue | issue resolved | Issue resolved | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFHD25001142419 | 2025-04-04 | 2025-04-07 | Cust 635 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on CPCB server. | As of now we are closing this complaint, we will reopen once we get access to your system. | As of now we are closing this complaint, we will reopen once we get access to your system. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD25002202420 | 2025-04-07 | 2025-04-09 | Cust 613 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to APPCB server. | Now the issue has been resolved and data is live on the APPCB server. | Now the issue has been resolved and data is live on the APPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD25002952421 | 2025-04-08 | 2025-04-14 | Cust 603 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | Flow cell for Multiparameter(AWA-FLOWCELL) | Calibration job | Work completed | Attained | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD25003692422 | 2025-04-10 | 2025-04-15 | Cust 1039 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | calibration | issue resolved | attained | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD25006392423 | 2025-04-17 | 2025-04-22 | Cust 1000 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | CALIBRATION | Sensor sent to pune for repair | Attained | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD25009982424 | 2025-04-26 | 2025-05-01 | Cust 1023 | HYDERABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | AQUAMAX DOT(1) | commissioning | commissioning completed | Attained | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD25010852425 | 2025-04-28 | 2025-05-01 | Cust 1037 | HYDERABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | commisisoning | commissioning completed | Commissioning completed | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFHD25018642426 | 2025-05-12 | 2025-05-15 | Cust 613 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | data not uploaded on APPCB server. | Currently data is reflecting on the APPCB and local server. | data is reflecting on the APPCB and local server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD25018812427 | 2025-05-12 | 2025-05-18 | Cust 1040 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Electrode | Forbes Marshall Cond. Cell Model 8310, K=0.01(FM 8310-1.5) | Calibration issue | issue resolved | issue resolved | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFHD25018922428 | 2025-05-12 | 2025-05-16 | Cust 1000 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | calibration job | CALIBRATION DONE | attained | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD25019712429 | 2025-05-14 | 2025-05-19 | Cust 984 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | DISPLAY not working | Suggested to send analyzer to pune | Attained | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD25021272430 | 2025-05-18 | 2025-05-27 | Cust 1041 | HYDERABAD | Steam System Division | CLOSED | QUALITY PROBLEM | PRV - MEASURETROL | FM MEASURETROL(SIZE 080NB) | Low PRS Measuretrol stopped working | R&D AND QC TEAM VISITED THE SITE AND HAS REPALCED THE ELECTRONICS AND ACTUATOR HOUSING TO THE IP65 FOAM TYPE GASKET | R&amp;amp;amp;D AND QC TEAM VISITED THE SITE AND HAS REPALCED THE ELECTRONICS AND ACTUATOR HOUSING TO THE IP65 FOAM TYPE GASKET | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFHD25021612431 | 2025-05-19 | 2025-05-23 | Cust 1006 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera streaming issue on CPCB server. | As of now we are closing this complaint. We will reopen this once the camera issue is resolved from customer end. | As of now we are closing this complaint. We will reopen this once the camera issue is resolved from customer end. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD25023202432 | 2025-05-22 | 2025-05-26 | Cust 1023 | HYDERABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | AQUAMAX DOT(1) | commissioning | Analyser commissioned successfully. | Analyser commissioned successfully. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFHD25023652433 | 2025-05-23 | 2025-05-26 | Cust 979 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | CPCB camera streaming issue. | After restarting camera streaming window, camera is streaming live on CPCB server. | After restarting camera streaming window, camera is streaming live on CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFHD25026962434 | 2025-05-29 | 2025-05-30 | Cust 969 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to CPCB and APPCB servers. | As of now we are closing this complaint. We will reopen this once we get access to the system. | As of now we are closing this complaint. We will reopen this once we get access to the system. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD25026982435 | 2025-05-29 | 2025-06-02 | Cust 1023 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | AQUADO Analyzer | Aquamax system(1) | Calibration | issue resolved | Attained | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD25028132436 | 2025-05-30 | 2025-05-31 | Cust 1033 | HYDERABAD | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | Conductivity Transmitter Aquacon | Conductivity Transmitter Aquamon 4000(COND-Trans Aqu-4000) | INSTALLATION GUIDELINES FOR THE PH TRANSMITTER | INSTALLATION GUIDELINES GIVEN FOR THE CCDS SYSTEM | INSTALLATION GUIDELINES GIVEN FOR THE CCDS SYSTEM | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFHD25030902437 | 2025-06-04 | 2025-06-06 | Cust 969 | HYDERABAD | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to CPCB and APPCB server. | Now all the cameras are streaming live on the CPCB server. Data from multiparameter and flowmeters is showing live on the CPCB and APPCB servers. | Now all the cameras are streaming live on the CPCB server. Data from multiparameter and flowmeters is showing live on the CPCB and APPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFHD25031202438 | 2025-06-04 | 2025-06-09 | Cust 1042 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Aqua4Trans | Aqua4Trans(1) | Service visit | Calibration performed | attained | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD25033632439 | 2025-06-10 | 2025-06-16 | Cust 641 | HYDERABAD | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | COMMISSIONING | commissioned | attended | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFHD25034712440 | 2025-06-13 | 2025-06-17 | Cust 641 | HYDERABAD | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration | Troubleshooting done analyser working perfectly | attained | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFID23035312441 | 2023-06-05 | 2023-06-09 | Cust 1043 | INDONESIA | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | dismantle and set up configuration for AWA | process dismantle is completed and analyzer has taken out from site. | process completed and analyzer has taken out from site. | NO | No Trip | Intops | N |  | Analytical - CHAKAN
AFID23035352442 | 2023-06-05 | 2023-06-09 | Cust 1044 | INDONESIA | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | check and discuss for material spares. | Problem with peristic pump shaft broken and need to change with new | Offer for new peristic pump for solve issue | NO | No Trip | Intops | N |  | Analytical - CHAKAN
AFID23042852443 | 2023-06-20 | 2023-06-26 | Cust 1045 | INDONESIA | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Display showing is error | Problem with MUVX and Coms Card and need to replace | problem with MUVX and Coms card and this need to change with new. | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFID23049202444 | 2023-06-30 | 2023-06-30 | Cust 1046 | INDONESIA | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | visit for replacement TSS Aqua4trans with new | After replace with NEW Aqua4trans reading communication normal | After visit for replace new Aqua4Trans for TSS, now reading communication normal | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFID23055032445 | 2023-07-12 | 2023-07-19 | Cust 1047 | INDONESIA | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Display showing mode display need to replace for MUVX Module AWA-EA313 and RS-485 CARD & PAIR OF EEPROM RS-485 CARD & PAIR OF EEPROM. | Material dispatch against RF OA 5524600570 | Material dispatch against RF OA 5524600570 | YES | No Trip | Intops | Y |  | Analytical - CHAKAN
AFID23055482446 | 2023-07-12 | 2023-07-19 | Cust 1048 | INDONESIA | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Problem. Motor not rotated [AWA-HE 741 Sampling pump + Adaptor AWA-ME500] Display mode [EAEB PCX2 without modules (Connectors mounted) Communication error [Falcon Mini scanner] | Material dispatch against RF OA 5524600592 | Material dispatch against RF OA 5524600592 | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFID23073282447 | 2023-08-15 | 2023-08-15 | Cust 1048 | INDONESIA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | calibration with third party LAB , INTERTEK | Process calibration has completed | Process calibration has completed as per attach report. | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFID23087052448 | 2023-09-11 | 2023-09-18 | Cust 1047 | INDONESIA | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | replacement for MUVX and RS485 card | replacement for RF material have completed | replacement for RF material have completed | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFID23090312449 | 2023-09-18 | 2023-09-25 | Cust 1049 | INDONESIA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 3rd visit AMC | AMC visit finish | After annual maintenance service - AMC Found issue have signal pH very less 500 and pH2 is 2203 for this need to be replaced. | NO | No Trip | Intops | N |  | Analytical - CHAKAN
AFID23094142450 | 2023-09-27 | 2023-10-03 | Cust 1047 | INDONESIA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Replacement RF material | Working has complete as per MoM report | all completed for solve issue in customer site | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFID23094152451 | 2023-09-27 | 2023-10-03 | Cust 1045 | INDONESIA | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Display backlight dark | problem with LCD and this is process with replace with new LCD | Please see all information in MoM report as per analyze having problem with LCD Display | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFID23094362452 | 2023-09-27 | 2023-12-18 | Cust 1047 | INDONESIA | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | MPA Analyzer DISPLAY Backlight | Replacement parts received and installed in customer. result problem solved. | Replacement parts received and installed in customer. result problem solved. | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFID23109982453 | 2023-10-27 | 2023-10-31 | Cust 1046 | INDONESIA | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Connectivity analyzer to KLHK | Process connectivity is finish. | Process connectivity is finish with online support | NO | No Trip | Intops | N |  | Analytical - CHAKAN
AFID23151182454 | 2024-01-09 | 2024-01-16 | Cust 1047 | INDONESIA | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | Ph Electrode | FM1104(1) | Guide for customer replacement spares pH sensor for MPA | Support provided by Zaenal | Support provided by Zaenal | NO | No Trip | Intops | N |  | Analytical - CHAKAN
AFID23182432455 | 2024-03-05 | 2024-03-11 | Cust 1048 | INDONESIA | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Aquat4trans RS485 card have burn | Material dispatch against RF OA 5524601970 | Material dispatch against RF OA 5524601970 | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFID24018392456 | 2024-05-08 | 2024-05-15 | Cust 1047 | INDONESIA | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Installation and calibration for TSS sensor after customer bought sensor | After replacing with new TSS Sensor, now system running. | system running a good after replacing with new TSS sensor. | NO | No Trip | Intops | N |  | Analytical - CHAKAN
AFID24057952457 | 2024-07-23 | 2024-07-29 | Cust 1050 | INDONESIA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Electrode | Forbes Marshall Cell Model 8311, K=0.1(FM -8311) | reading not accurate | sensor no response after check with buffer solution, need to replace sensor | Sensor has damage, and offer for new sensor send to customer. | NO | No Trip | Intops | N |  | Analytical - CHAKAN
AFID24057952457 | 2024-07-23 | 2024-07-29 | Cust 1050 | INDONESIA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Electrode | Forbes Marshall Cond. Cell Model 8310, K=0.01(FM -8310) | Reading not accurate | sensor no response after check with buffer solution, need to replace sensor | Sensor has damage, and offer for new sensor send to customer. | NO | No Trip | Intops | N |  | Analytical - CHAKAN
AFID24058012458 | 2024-07-23 | 2024-07-29 | Cust 1044 | INDONESIA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Turbidity Meter | Optex turbidity sensor(1) | Sensor reading value is zero | have leaked water inside sensor, need to replace with new sensor | Quotation new sensor has been sending to customer | NO | No Trip | Intops | N |  | Analytical - CHAKAN
AFID24097732459 | 2024-10-01 | 2024-10-07 | Cust 1051 | INDONESIA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Electrode | FM1104(1) | Support calibration pH sensor atter buy new spares | process calibration has successful used new sensor | process calibration done | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFID24097742460 | 2024-10-01 | 2024-10-07 | Cust 1052 | INDONESIA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | Photo Detector for Multiparameter(AWA-PD) | replacement pH1 & pH2 set with light guide consist of. PH1 detector, PH2 detector, Light guide Application at COD analyzer. | process replacement spares support by online given to customer | replacement with new spares is done | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFIN24097672461 | 2024-10-01 | 2024-10-04 | Cust 1053 | INDORE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue. | we took access of the black box and observed that the internet is not coming till the FMBB, now the same issue has been resolved from your side and data is live on CPCB + MPPCB server. | we took access of the black box and observed that the internet is not coming till the FMBB, now the same issue has been resolved from your side and data is live on CPCB + MPPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFIN24103422462 | 2024-10-15 | 2024-10-19 | Cust 1054 | INDORE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB and Mppcb data uploaded issue. | Data is live on the server. | Data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFIN24186542463 | 2025-03-25 | 2025-03-27 | Cust 528 | INDORE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | MPPCB data upload issue. | We have made changes in configuration as per MPPCB mail, now data is live on MPPCB server. | We have made changes in configuration as per MPPCB mail, now data is live on MPPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFIN25018782464 | 2025-05-12 | 2025-05-27 | Cust 1055 | INDORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Electrode | ORP Electrode- Polilyte RX 120(1) | 1. ORP Sensor Malfunction | Replacement Material has been booked through MARK EXPRESS PRIVATE LIMITED, with LR No. 210141269 dated 24-May-2025. | Replacement Material has been booked through MARK EXPRESS PRIVATE LIMITED, with LR No. 210141269 dated 24-May-2025. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFIN25018802465 | 2025-05-12 | 2025-05-27 | Cust 1055 | INDORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Electrode | Chemotrode ORP 200(HT238 744) | 2.ORP Sensor Display Dim | Replacement Material has been booked through MARK EXPRESS PRIVATE LIMITED, with LR No. 210141269 dated 24-May-2025. | Replacement Material has been booked through MARK EXPRESS PRIVATE LIMITED, with LR No. 210141269 dated 24-May-2025. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFJP23034052466 | 2023-06-02 | 2023-06-02 | Cust 1056 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC | AMC completed | Attended | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP23050742467 | 2023-07-03 | 2023-07-03 | Cust 623 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | Water Detector for Multiparameter(AWA-WD) | stuck on cleaning cycle | CHECKED AND FOUND WATER DETECTOR NOT WORKING, SAME BYPASS START THE ANALYSER, NOW CUSTOMER IS ORDERING THE SENSOR | attended | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP23059242468 | 2023-07-20 | 2023-07-21 | Cust 1057 | JAMSHEDPUR | DCS - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | MICROCON+ER SERIES | CONTROLLER MODULE (ER-201)(ER-201) | Black Box hang issue. | We took the access to the black box and changed the polling frequency, and now data is live on all the servers. | We took the access to the black box and changed the polling frequency, and now data is live on all the servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFJP23064662469 | 2023-07-30 | 2023-08-07 | Cust 1058 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration and maintenance | Calibration and Maintenance done | Calibration and Maintenance done | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFJP23064662469 | 2023-07-30 | 2023-08-07 | Cust 1058 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration and maintenance | Calibration and Maintenance done | Calibration and Maintenance done | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFJP23064662469 | 2023-07-30 | 2023-08-07 | Cust 1058 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration and maintenance | Calibration and Maintenance done | Calibration and Maintenance done | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFJP23064662469 | 2023-07-30 | 2023-08-07 | Cust 1058 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration and maintenance | Calibration and Maintenance done | Calibration and Maintenance done | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFJP23064672470 | 2023-07-30 | 2023-08-07 | Cust 1059 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration and Maintenance | AMC completed | AMC COMPLETED | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFJP23068342471 | 2023-08-04 | 2023-08-07 | Cust 1060 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera configuration needs to be done. | we took access of the system and found that the Camera IP is not working on the VLC media player hence their IT team is working on it and as of now we are closing this complaint. | we took access of the system and found that the Camera IP is not working on the VLC media player hence their IT team is working on it and as of now we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFJP23069062472 | 2023-08-07 | 2023-08-12 | Cust 1056 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration | Calibration has been perfomed successfully. | Attended | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP23069062472 | 2023-08-07 | 2023-08-12 | Cust 1056 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration | Calibration has been perfomed successfully. | Attended | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP23069062472 | 2023-08-07 | 2023-08-12 | Cust 1056 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration | Calibration has been perfomed successfully. | Attended | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP23069062472 | 2023-08-07 | 2023-08-12 | Cust 1056 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration | Calibration has been perfomed successfully. | Attended | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP23071462473 | 2023-08-10 | 2023-08-18 | Cust 650 | JAMSHEDPUR | DCS - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | MICROCON+ER SERIES | CONTROLLER MODULE (ER-201)(ER-201) | BB hang issue. | We took access of the system and made some changes in the configuration of the Black Box and also in the Router configuration so now the data is live on the server. | We took access of the system and made some changes in the configuration of the Black Box and also in the Router configuration so now the data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFJP23076642474 | 2023-08-21 | 2023-08-28 | Cust 1059 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration job | Calibration Done | Work completed | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP23080642475 | 2023-08-28 | 2023-08-31 | Cust 1061 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | DATCON | DATCON(DATCON) | CPCB Connectivity need to be done. | CPCB Connectivity has been completed and data is live on the server. | CPCB Connectivity has been completed and data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFJP23083142476 | 2023-09-01 | 2023-09-01 | Cust 1061 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | DATCON | DATCON(DATCON) | MPPCB connectivity needs to be done. | we have completed the MPPCB Connectivity and data is live on the server. | we have completed the MPPCB Connectivity and data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFJP23083242477 | 2023-09-03 | 2023-09-03 | Cust 1062 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | Commissioning | commissioning done | Work completed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFJP23083242477 | 2023-09-03 | 2023-09-03 | Cust 1062 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | Commissioning | Commissioning Done | Work completed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFJP23086582478 | 2023-09-11 | 2023-09-11 | Cust 1063 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC VISIT | COMPLETED PREVENTIVE MAINTENANCE AND CALIBRATION | attended | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP23086582478 | 2023-09-11 | 2023-09-11 | Cust 1063 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC VISIT | COMPLETED PREVENTIVE MAINTENANCE AND CALIBRATION | attended | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP23086582478 | 2023-09-11 | 2023-09-11 | Cust 1063 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC VISIT | COMPLETED PREVENTIVE MAINTENANCE AND CALIBRATION | attended | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP23086582478 | 2023-09-11 | 2023-09-11 | Cust 1063 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC VISIT | COMPLETED PREVENTIVE MAINTENANCE AND CALIBRATION | attended | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP23091162479 | 2023-09-20 | 2023-09-28 | Cust 1059 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Maintenance of Multiparameter Analyser | AMC activity has been completed. | AMC activity has been completed. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFJP23093362480 | 2023-09-25 | 2023-09-28 | Cust 1061 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB and MPPCB data uploaded issue. | we took the access to the system and did some changes from configuration end and now data is live on CPCB server. | we took the access to the system and did some changes from configuration end and now data is live on CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFJP23100582481 | 2023-10-06 | 2023-10-06 | Cust 1056 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC VISIT FOR 4NOS OF MULTIPARAMETER ANALYSER | AMC VISIT COMPLETED | attended | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP23100582481 | 2023-10-06 | 2023-10-06 | Cust 1056 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC VISIT FOR 4NOS OF MULTIPARAMETER ANALYSER | AMC VISIT COMPLETED | attended | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP23100582481 | 2023-10-06 | 2023-10-06 | Cust 1056 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC VISIT FOR 4NOS OF MULTIPARAMETER ANALYSER | AMC VISIT COMPLETED | attended | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP23100582481 | 2023-10-06 | 2023-10-06 | Cust 1056 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC VISIT FOR 4NOS OF MULTIPARAMETER ANALYSER | AMC VISIT COMPLETED | attended | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP23100682482 | 2023-10-06 | 2023-10-09 | Cust 623 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc visit and flow cell installation | amc visit completed, need to replace ph2 detector | attended | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP23106282483 | 2023-10-18 | 2023-10-19 | Cust 1059 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc visit for preventive maintenance | AMC VISIT COMPLETED, MULTIPARAMETER ANALYSER WORKING SATISFACTORY | attended | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP23112142484 | 2023-10-30 | 2023-10-30 | Cust 1058 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | THERMO SERIES | TORRIDAL SENSOR(1) | Analyser showing constant value. | cleaning and calibration has been carried out. | Attended | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFJP23112162485 | 2023-10-30 | 2023-10-30 | Cust 1056 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC | AMC completed | Attended | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP23112162485 | 2023-10-30 | 2023-10-30 | Cust 1056 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC | AMC completed mom attached. | Attended | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP23112162485 | 2023-10-30 | 2023-10-30 | Cust 1056 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC | AMC completed mom attached. | Attended | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP23112162485 | 2023-10-30 | 2023-10-30 | Cust 1056 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC | AMC completed mom attached. | Attended | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP23116582486 | 2023-11-06 | 2023-11-09 | Cust 1061 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | DATCON | DATCON(DATCON) | Configuration changes need to be done as per the new Scanner. | we have made changes in the configuration as per the New Scanner and now the data is live on both servers | we have made changes in the configuration as per the New Scanner and now the data is live on both servers | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFJP23117742487 | 2023-11-08 | 2023-11-12 | Cust 1056 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration Job | Calibration Done | Work completed | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP23117742487 | 2023-11-08 | 2023-11-12 | Cust 1056 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration Job | Calibration Done | Work completed | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP23117742487 | 2023-11-08 | 2023-11-12 | Cust 1056 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration Job | Calibration Done | Work completed | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP23117742487 | 2023-11-08 | 2023-11-12 | Cust 1056 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration Job | Calibration Done | Work completed | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP23118682488 | 2023-11-10 | 2023-11-13 | Cust 1064 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Calibration of pH sensor | Calibration has been carried out successfully through remote support. | Calibration has been carried out successfully through remote support. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFJP23118912489 | 2023-11-11 | 2023-11-13 | Cust 1060 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Blackbox Hanging issue in ETP Blackbox. | Updated patch in the router. Mail attached. | Attended | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP23144952490 | 2023-12-30 | 2023-12-30 | Cust 1058 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | LOCAL DISPLAY READING AND CLODE READING NOT MATCHING | same resolved remotely by doing mA calibration | attended | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP23147362491 | 2024-01-02 | 2024-01-02 | Cust 1060 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Aqua4Trans | Aqua4Trans(1) | calibration of pH sensor and 4-20mA configuration | calibration of pH sensor and 4-20mA configuration completed remotely | attended | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFJP23147392492 | 2024-01-02 | 2024-01-03 | Cust 1056 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc call | calibration and preventive maintenance completed | attended | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP23147392492 | 2024-01-02 | 2024-01-03 | Cust 1056 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc call | calibration and preventive maintenance completed | attended | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP23147392492 | 2024-01-02 | 2024-01-03 | Cust 1056 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc call | calibration and preventive maintenance completed | attended | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP23147392492 | 2024-01-02 | 2024-01-03 | Cust 1056 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc call | calibration and preventive maintenance completed | attended | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP23151632493 | 2024-01-10 | 2024-01-12 | Cust 1059 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | monthly amc visit for preventive maintenance and calibration | preventive maintenance and calibration of multiparameter done | attended | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP23155012494 | 2024-01-16 | 2024-01-17 | Cust 1065 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Aqua4Trans | PH input card(1) | ph, turbidity and conductivity reading fluctuating | completed the maintenance make analyzer operational. | attended | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFJP23155012494 | 2024-01-16 | 2024-01-17 | Cust 1065 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Aqua4Trans | PH input card(1) | ph, turbidity and conductivity reading fluctuating | completed the maintenance make analyzer operational. | attended | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFJP23155012494 | 2024-01-16 | 2024-01-17 | Cust 1065 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Aqua4Trans | PH input card(1) | ph, turbidity and conductivity reading fluctuating | completed the maintenance make analyzer operational. | attended | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFJP23155012494 | 2024-01-16 | 2024-01-17 | Cust 1065 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Aqua4Trans | PH input card(1) | ph, turbidity and conductivity reading fluctuating | completed the maintenance make analyzer operational. | attended | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFJP23155012494 | 2024-01-16 | 2024-01-17 | Cust 1065 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Aqua4Trans | PH input card(1) | ph, turbidity and conductivity reading fluctuating | completed the maintenance make analyzer operational. | attended | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFJP23155012494 | 2024-01-16 | 2024-01-17 | Cust 1065 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Aqua4Trans | PH input card(1) | ph, turbidity and conductivity reading fluctuating | completed the maintenance make analyzer operational. | attended | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFJP23162442495 | 2024-01-31 | 2024-03-25 | Cust 895 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | 4-20 mA input card of Aqua4trans(1) | COMMISSIONING OF ANALYSERS | COMMISSIONING COMPLETED | attended | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFJP23167302496 | 2024-02-08 | 2024-02-08 | Cust 895 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4Trans(1) | installation and commissioning | installation completed due to site not ready and sample not available commissioning pending | attended | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFJP23169362497 | 2024-02-13 | 2024-02-20 | Cust 1056 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration Job | Calibration Done | Work completed | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP23169362497 | 2024-02-13 | 2024-02-20 | Cust 1056 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration Job | Calibration Done | Work completed | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP23169362497 | 2024-02-13 | 2024-02-20 | Cust 1056 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration Job | Calibration Done | Work completed | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP23169362497 | 2024-02-13 | 2024-02-20 | Cust 1056 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration Job | Calibration Done | Work completed | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP23170822498 | 2024-02-15 | 2024-02-15 | Cust 1059 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | monthly amc | preventive maintenance and calibration of multiparameter analyser done | attended | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP23174052499 | 2024-02-21 | 2024-03-01 | Cust 1056 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration Job | Calibration Done | Work completed | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP23174052499 | 2024-02-21 | 2024-03-01 | Cust 1056 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration Job | Calibration Done | Work completed | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP23174052499 | 2024-02-21 | 2024-03-01 | Cust 1056 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration Job | Calibration Done | Work completed | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP23174052499 | 2024-02-21 | 2024-03-01 | Cust 1056 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration Job | Calibration Done | Work completed | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP23180792500 | 2024-03-01 | 2024-03-11 | Cust 1056 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | Aqua4Trans | Aqua4Trans(1) | Calibration | Calibration Done | Work completed | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP24003202501 | 2024-04-08 | 2024-04-22 | Cust 480 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning of CX2000 Analyser, Site address as below. RSP SAIL site ROURKELA Odisa Pin code 769001 | Commissioning Done | Commisssioning done | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFJP24003202501 | 2024-04-08 | 2024-04-22 | Cust 480 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning of CX2000 Analyser, Site address as below. RSP SAIL site ROURKELA Odisa Pin code 769001 | Commissioning Done | Commisssioning done | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFJP24003202501 | 2024-04-08 | 2024-04-22 | Cust 480 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning of CX2000 Analyser, Site address as below. RSP SAIL site ROURKELA Odisa Pin code 769001 | Commissioning Done | Commisssioning done | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFJP24003202501 | 2024-04-08 | 2024-04-22 | Cust 480 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning of CX2000 Analyser, Site address as below. RSP SAIL site ROURKELA Odisa Pin code 769001 | Commissioning Done | Commisssioning done | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFJP24016422502 | 2024-05-03 | 2024-05-03 | Cust 1056 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC VISIT | AMC COMPLETED | attended | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP24016422502 | 2024-05-03 | 2024-05-03 | Cust 1056 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC VISIT | AMC COMPLETED | attended | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP24016422502 | 2024-05-03 | 2024-05-03 | Cust 1056 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC VISIT | AMC COMPLETED | attended | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP24016422502 | 2024-05-03 | 2024-05-03 | Cust 1056 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC VISIT | AMC COMPLETED | attended | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP24016462503 | 2024-05-03 | 2024-05-03 | Cust 480 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | COMMISSINING OF 4 NO'S OF MULTIPARAMETER | COMMISSINING COMPLETED SUCCESFULLY | attended | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFJP24016462503 | 2024-05-03 | 2024-05-03 | Cust 480 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | COMMISSINING OF 4 NO'S OF MULTIPARAMETER | COMMISSINING COMPLETED SUCCESFULLY | attended | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFJP24016462503 | 2024-05-03 | 2024-05-03 | Cust 480 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | COMMISSINING OF 4 NO'S OF MULTIPARAMETER | COMMISSINING COMPLETED SUCCESFULLY | attended | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFJP24016462503 | 2024-05-03 | 2024-05-03 | Cust 480 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | COMMISSINING OF 4 NO'S OF MULTIPARAMETER | COMMISSINING COMPLETED SUCCESFULLY | attended | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFJP24016592504 | 2024-05-03 | 2024-05-03 | Cust 1066 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | Chlorine Analyzer-MAC 100 | MAC 100(1) | COMMISSIONING OF CHLORINE ANALYSER, PH ANALYSER AND TURBIDITY ANALYSER | COMMISSIONING OF CHLORINE ANALYSER, PH ANALYSER AND TURBIDITY ANALYSER COMPLETED | attended | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFJP24016722505 | 2024-05-03 | 2024-05-06 | Cust 1067 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue. | we took access to your system and changed configuration for NOx, now data is live on CPCB server and matching with actual readings | we took access to your system and changed configuration for NOx, now data is live on CPCB server and matching with actual readings | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFJP24020152506 | 2024-05-12 | 2024-05-20 | Cust 823 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Display Got Blanck | Display found faultiy | Work completed | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP24020712507 | 2024-05-14 | 2024-05-20 | Cust 526 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Tranmitter Polymetron | Conductivity Tranmitter Polymetron 9125(9125) | calibration of the Analysers Site contact person - Prakash Number - +917587796051 Address - vedanta limited Jharsuguda odisha. | SITE NOT READY AS CONFIRM BY CUSTOMER IT WILL TAKE ANOTHER 15 DAYS | attended | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFJP24027462508 | 2024-05-27 | 2024-09-20 | Cust 1058 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | THERMO SERIES | AquaPro Analyzer(1) | Thermo scientific AquaPro transmitter are getting Hang, rebooting itself and stuck on constant reading. | As it is 2018 model need to purchase a new AP main board door assembly, model AP0DA upgrade the software. | attended | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFJP24027462508 | 2024-05-27 | 2024-09-20 | Cust 1058 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | THERMO SERIES | AquaPro Analyzer(1) | Thermo scientific AquaPro transmitter are getting Hang, rebooting itself and stuck on constant reading. | software upgraded successfully. | attended | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFJP24035742509 | 2024-06-11 | 2024-06-17 | Cust 1060 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Aqua4Trans | Aqua4Trans(1) | pH value is not showing correct, it measuring like 0.0 pH to 2.2 pH which is not correct. | pH, tss and temp. was working fine, issue in the data uploading same has been resolved. | attended | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP24036522510 | 2024-06-13 | 2024-06-14 | Cust 1068 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | ETP Data mismatch issue on CPCB server. | The problem was that the CEMS data was interfering with the ETP data, causing incorrect readings such as temperature fluctuating between 0 to 440, pH between 0 to 2.2, and TSS between 0 to 540, which were not possible. We resolved the issue by connecting a separate cable in the black box for the ETP. | The problem was that the CEMS data was interfering with the ETP data, causing incorrect readings | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFJP24042992511 | 2024-06-26 | 2024-06-29 | Cust 1069 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | JSPCB ISSUE. | Now the hardware issue has been resolved and data is live on the server. | Now the hardware issue has been resolved and data is live on the server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFJP24043612512 | 2024-06-27 | 2024-06-29 | Cust 1070 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB + GSPCB ISSUE | Now the issue has been resolved and data is live on both the sites | Now the issue has been resolved and data is live on both the sites | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFJP24049242513 | 2024-07-05 | 2024-07-05 | Cust 623 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC schedule visit | completed schedule visit done calibration & preventing maintenance | Attended | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP24049242513 | 2024-07-05 | 2024-07-05 | Cust 623 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC schedule visit | completed schedule visit done calibration & preventing maintenance | Attended | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP24053282514 | 2024-07-14 | 2024-07-14 | Cust 1071 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | THERMO SERIES | AquaPro Analyzer(1) | COMMISSIONING OF HCL CONCENTRATION ANALYSER | COMMISSIONING COMPLETED. | attended | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFJP24054982515 | 2024-07-17 | 2024-07-17 | Cust 1072 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | DIGITAL CONEECTIVITY | Aqua4Trans | Aqua4Trans(1) | DIGITAL CONNECTIVITY TUNING | PROGRAM MODIFICATION DONE | attended | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFJP24055932516 | 2024-07-18 | 2024-07-22 | Cust 823 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | pH and RS485 card was faulty | pH and RS485 card replaced successfully | attended | NO | Paid | Core | N |  | Analytical - CHAKAN
AFJP24065972517 | 2024-08-04 | 2024-08-10 | Cust 1056 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration job | calibration done | work completed | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP24065972517 | 2024-08-04 | 2024-08-10 | Cust 1056 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration job | calibration done | work completed | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP24065972517 | 2024-08-04 | 2024-08-10 | Cust 1056 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration job | calibration done | work completed | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP24067112518 | 2024-08-05 | 2024-08-05 | Cust 823 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | reading fluctuating after replacement of motherboard | reprogram and completed the calibration and issue resolved | attended | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP24071402519 | 2024-08-15 | 2024-08-20 | Cust 1060 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | Aqua4Trans | Aqua4Trans(1) | Healthiness checkup visit | Healthiness checked and found OK. | work completed | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP24077622520 | 2024-08-27 | 2024-08-27 | Cust 1073 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | THERMO SERIES | AquaPro Analyzer(1) | visit required for the Replacement of REF. electrode. | sensor replaced successfully and completed the calibration. | attended | NO | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFJP24089262521 | 2024-09-16 | 2024-10-14 | Cust 480 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Analyzer display touch is not working. Site person - Mr. Manoj : 9460205071 (RSP Site) | Analyzer display touch is working after led alignment. | attended | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFJP24091312522 | 2024-09-20 | 2024-09-20 | Cust 1074 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | DIGITAL CONEECTIVITY | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Cloud server connectivity for ETP device. | we completed the cloud server connectivity for the ETP device and data is live on the cloud server. | we completed the cloud server connectivity for the ETP device and data is live on the cloud server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFJP24097872523 | 2024-10-02 | 2024-10-29 | Cust 650 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | BLACK BOX hanging issue, | Data is now live continuously. | Data is now live continuously. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | DCS - CHAKAN
AFJP24098442524 | 2024-10-03 | 2024-10-04 | Cust 1060 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on JSPCB. | we made some modifications to the BLACK BOX configuration and the data is l | we made some modifications to the BLACK BOX configuration and the data is l | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFJP24100102525 | 2024-10-07 | 2024-10-07 | Cust 1072 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | DIGITAL CONEECTIVITY | Aqua4Trans | RS485 card- of Aqua4Trans(1) | Data uploading stop due to modbus communication not working properly | changed communication from modbus to 4-20mA issue resolved | attended | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFJP24100102525 | 2024-10-07 | 2024-10-07 | Cust 1072 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | DIGITAL CONEECTIVITY | Aqua4Trans | RS485 card- of Aqua4Trans(1) | Data uploading stop due to modbus communication not working properly | changed communication from modbus to 4-20mA issue resolved | attended | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFJP24103622526 | 2024-10-15 | 2024-10-18 | Cust 1061 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | DATCON | DATCON(DATCON) | CPCB and MPPCB data update issue for ETP device. | The issue was from the network side and the same has been resolved and the data is now live on both servers. | The issue was from the network side and the same has been resolved and the data is now live on both servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFJP24104562527 | 2024-10-17 | 2024-10-21 | Cust 623 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC | Visited the site and done calibration of the analyzer with the respective standard solutions. | Attended | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP24104562527 | 2024-10-17 | 2024-10-21 | Cust 623 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC | Visited the site and done calibration of the analyzer with the respective standard solutions. | Attended | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP24110562528 | 2024-10-29 | 2024-10-29 | Cust 1075 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | commissioning and calibration | successfully completed the commissioning and calibration. | attended | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFJP24112502529 | 2024-11-04 | 2024-11-10 | Cust 1056 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration Job | Calibration done | work completed | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP24112502529 | 2024-11-04 | 2024-11-10 | Cust 1056 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration Job | calibration done | work completed | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP24113912530 | 2024-11-07 | 2024-11-13 | Cust 1060 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | calibration job | Calibration done | work completed | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP24116932531 | 2024-11-13 | 2024-11-19 | Cust 823 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Recommissioning | COMMISSIONING DONE | work completed | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFJP24123982532 | 2024-11-27 | 2024-12-03 | Cust 1056 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | Silica Analyser | Silica 8891(8891) | Sample SOV is not working | sample sov replace with new one | work completed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFJP24134302533 | 2024-12-15 | 2024-12-21 | Cust 1060 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4Trans(1) | Commissioning of pH and TSS analyzer | Commissioning done | commissioning done | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFJP24134312534 | 2024-12-15 | 2024-12-21 | Cust 1056 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration Job | Calibration done | AMC Done | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP24134312534 | 2024-12-15 | 2024-12-21 | Cust 1056 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration Job | Calibration Done | AMC Done | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP24136402535 | 2024-12-19 | 2024-12-26 | Cust 1056 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | Aqua4Trans | Aqua4Trans(1) | Calibration Job | Calibration Done | calibration done | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP24148242536 | 2025-01-08 | 2025-01-10 | Cust 1076 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | ETP data is not updating on the JSPCB server. | We took access of the BB and found that the internet was not coming till the BB and now the has been resolved and the data is live . | We took access of the BB and found that the internet was not coming till the BB and now the has been resolved and the data is live . | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFJP24166912537 | 2025-02-12 | 2025-02-18 | Cust 1076 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration job | Calibration done | work completed | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP24170292538 | 2025-02-20 | 2025-02-27 | Cust 1077 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration job | calibration done | work completed | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP24170292538 | 2025-02-20 | 2025-02-27 | Cust 1077 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration job | Calibration done | work completed | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP24174002539 | 2025-02-27 | 2025-03-05 | Cust 895 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Chlorine Analyzer-MAC 100 | MAC 100(1) | Low value showing on the DDU | Sensor Found Faulty | work completed | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFJP24174142540 | 2025-02-27 | 2025-04-06 | Cust 1059 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC CALL | amc call completed | attended | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP24181872541 | 2025-03-13 | 2025-03-20 | Cust 1060 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | calibration Job | Healthiness checked | work completed | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP25012912542 | 2025-04-30 | 2025-05-13 | Cust 1076 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | JSPCB data upload issue for ETP. | Here there was a network issue so after switching the SIM setting to 4G only and updating the router version, the data is now uploading continuously without any issues. | Here there was a network issue so after switching the SIM setting to 4G only and updating the router version, the data is now uploading continuously without any issues. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFJP25019382543 | 2025-05-13 | 2025-05-20 | Cust 1058 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | MULTIPARAMETER ANALYSER | pH input card-FM of Multiparameter(1) | Spare Installation | Spare Installation done | work completed | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFJP25027082544 | 2025-05-29 | 2025-06-02 | Cust 564 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | pH Transmitter Monec | 8965 CPU Card(CPU-card-8965) | Forbes Marshall make pH Meters are Not Functioning; accordingly | Ph sensor found faulty. Offer also submitted to customer. | Ph sensor found faulty. Offer also submitted to customer. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFJP25027082544 | 2025-05-29 | 2025-06-02 | Cust 564 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | pH Transmitter Monec | 8965 CPU Card(CPU-card-8965) | Forbes Marshall make pH Meters are Not Functioning; accordingly | Ph sensor found faulty. Offer also submitted to customer. | Ph sensor found faulty. Offer also submitted to customer. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFJP25027792545 | 2025-05-30 | 2025-06-04 | Cust 1056 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration Job | Calibration Done | work completed | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP25027832546 | 2025-05-30 | 2025-06-04 | Cust 1059 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration Job | Calibration Done | work completed | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFJP25031602547 | 2025-06-05 | 2025-06-07 | Cust 823 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | DIGITAL CONEECTIVITY | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | MODBUS COMMUNICATION ISSUE | took system online and resolved the issue, now modbus communication stabilized | attended | NO | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFJP25033932548 | 2025-06-11 | 2025-06-14 | Cust 1071 | JAMSHEDPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | pH commissioning | I provided remote support and successfully completed the pH transmitter commissioning, including wiring and calibration. | attended | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFJP25036622549 | 2025-06-17 |  | Cust 1060 | JAMSHEDPUR | Analytical - CHAKAN | ASSIGNED | CALIBRATION JOB | Aqua4Trans | Aqua4 trans-A4TW033000(1) | checking for dual channel Aqua4trans pH and tss heath checking |  |  | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFKK23033592550 | 2023-06-01 | 2023-06-02 | Cust 1078 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | DATCON | DATCON(DATCON) | For ETP Station, CPCB connectivity needs to be done. | For ETP Station, CPCB connectivity has been completed. | For ETP Station, CPCB connectivity has been completed. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFKK23035032551 | 2023-06-05 | 2023-06-08 | Cust 1079 | CALCUTTA | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Product Servicing on paid service basis. | Visit done. calibration report attached. | Visit done. calibration report attached. | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFKK23035612552 | 2023-06-06 | 2023-06-09 | Cust 915 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data uploaded issue. | we took access of the system and observed that ETP data is not communicating with the local system and we have checked data on Modscan but not getting any counts. Hence please this issue needs to be checked from the hardware side. | As of now, we are closing this complaint and will reopen this case once the hardware issue has been resolved. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFKK23036742553 | 2023-06-08 | 2023-06-12 | Cust 1080 | CALCUTTA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | 12 V POWER SUPPLY CARD for Multiparameter(AWA-PS1) | 0.3V o/p from SMPS-- analyzer is not working | Visited. Job done. MOM is attached. | Visited. Job done. MOM is attached. | NO | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK23037432554 | 2023-06-11 | 2023-06-11 | Cust 797 | CALCUTTA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC JOB | AMC VISIT MADE | AMC VISIT HAS BEEN MADE | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFKK23037432554 | 2023-06-11 | 2023-06-11 | Cust 797 | CALCUTTA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC JOB | AMC VISIT MADE | AMC VISIT HAS BEEN MADE | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFKK23037432554 | 2023-06-11 | 2023-06-11 | Cust 797 | CALCUTTA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC JOB | AMC VISIT MADE | AMC VISIT HAS BEEN MADE | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFKK23039502555 | 2023-06-14 | 2023-06-17 | Cust 476 | CALCUTTA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | 3-Way Electric Valve for Multiparameter(AWA-3WEV) | COD VALUES OF THE SAMPLE TOO HIGH AS COMPARED TO LAB RESULTS. | Calibration done. Report attached. System under observation. | Calibration done. Report attached. System under observation. | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFKK23039612556 | 2023-06-14 | 2023-06-17 | Cust 603 | CALCUTTA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | Flow cell for Multiparameter(AWA-FLOWCELL) | AMC call | VISIT MADE. | Site visited, MOM done. | NO | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK23043042557 | 2023-06-20 | 2023-06-27 | Cust 797 | CALCUTTA | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | MULTIPARAMETER ANALYSER | pH input card-FM of Multiparameter(1) | Panel of analyser to be shifted to another location under amc | Site visit made and found erection is going on. | Site visit made and found erection is going on. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFKK23044682558 | 2023-06-23 | 2023-06-26 | Cust 1081 | CALCUTTA | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Turbidity Input card of Aqua4Trans(1) | installation and commissioning | commissioned | commissioning has been done | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK23046252559 | 2023-06-26 | 2023-06-29 | Cust 844 | CALCUTTA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | pH input card-FM of Multiparameter(1) | AMC | Visit made. | Site visited, MOM done. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFKK23047122560 | 2023-06-27 | 2023-06-30 | Cust 797 | CALCUTTA | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | MULTIPARAMETER ANALYSER | Flowcell disc of Multiparameter(1) | Erection | Site visit made. MOM is attached | Site visit made. MOM is attached | NO | Paid | Core | N |  | Analytical - CHAKAN
AFKK23047552561 | 2023-06-28 | 2023-06-29 | Cust 528 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | The local Host is not working. | We took access of the system and restart the local services and now the data is live on the local server.. | Issue has been resolved hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFKK23047592562 | 2023-06-28 | 2023-07-03 | Cust 1082 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | local host not opening in local PC. | Today we took the access to the system and successfully did the eversense software reinstallation + local + CPCB configuration and now data is live on both the servers. | Today we took the access to the system and successfully did the eversense software reinstallation + local + CPCB configuration and now data is live on both the servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFKK23047782563 | 2023-06-28 | 2023-06-30 | Cust 528 | CALCUTTA | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | MULTIPARAMETER ANALYSER | Xenon Lamp Power supply for Multiparameter(AWA-XLAMPPS) | MPA not giving proper reading . Need to check and fix. | Visit done. Report attached. | Visit done. Report attached. | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFKK23048592564 | 2023-06-30 | 2023-07-04 | Cust 915 | CALCUTTA | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | The analyser to be removed from panel to ensure the panel fabrication modification at site | Site visit is made and job report is sent | Site visit made. Report is attached | NO | Paid | Core | N |  | Analytical - CHAKAN
AFKK23054352565 | 2023-07-11 | 2023-07-17 | Cust 1083 | CALCUTTA | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | VISI pro DO(VISI pro DO) | 3 nos DO meters at KMDA Halisahar | Visited. Commissioned. MOM is made. | Visited. Commissioned. MOM is attached. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK23055082566 | 2023-07-12 | 2023-07-17 | Cust 1080 | CALCUTTA | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | RS485 CARD for Multiparameter(AWA-RS485) | Recommissioning | Visited. Job done. MOM is attached | Visited. Job done. MOM is attached. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK23056832567 | 2023-07-15 | 2023-07-19 | Cust 915 | CALCUTTA | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Analyser to be installed in refurbished panel and installation of Scanner and RS 485 converter and reinstating CPCB connectivity | Site visit made. Job report is attached | Site visit made. Job report is attached | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFKK23057772568 | 2023-07-17 | 2023-07-20 | Cust 915 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | scanner is changed , so need to do configuration . | We have done the required configuration as pe new scanner . Data is now live on the CPCB server for the mentioned site. Save | We have done the required configuration as pe new scanner . Data is now live on the CPCB server for the mentioned site. Save | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFKK23058212569 | 2023-07-18 | 2023-07-22 | Cust 1081 | CALCUTTA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC VISIT | The Site was visited by Nikhil. MOM done. | The Site was visited by Nikhil. MOM done. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFKK23059002570 | 2023-07-19 | 2023-07-22 | Cust 1081 | CALCUTTA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC VISIT | The Site was visited by Nikhil. MOM done. | The Site was visited by Nikhil. MOM done. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFKK23071622571 | 2023-08-10 | 2023-08-11 | Cust 1084 | CALCUTTA | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | SULPHIDE ANALYSER | Power supply board(359110,20XXX) | Discussion | SITE VISITED, DISCUSSION DONE. | Site visited, Discussion done. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK23072432572 | 2023-08-14 | 2023-08-20 | Cust 663 | CALCUTTA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Electrode | CONDUCELL 2DC-PG-120(HT237 610) | Calibration | Site is attended & mom is attached | Site is attended & mom is attached | NO | Paid | Core | N |  | Analytical - CHAKAN
AFKK23074452573 | 2023-08-17 | 2023-08-24 | Cust 1084 | CALCUTTA | Analytical - CHAKAN | CLOSED | COMMISSIONING | SULPHIDE ANALYSER | SULPHIDE 8810 Single Channel ANALYSER(8810 Sulphide) | FM SULPHIDE ANALYZER COMMISSIONING | The site visit was made . MOM done. | The site visit was made . MOM done. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFKK23074452573 | 2023-08-17 | 2023-08-24 | Cust 1084 | CALCUTTA | Analytical - CHAKAN | CLOSED | COMMISSIONING | SULPHIDE ANALYSER | SULPHIDE 8810 Single Channel ANALYSER(8810 Sulphide) | FM SULPHIDE ANALYZER COMMISSIONING | The site visit was made . MOM done. | The site visit was made . MOM done. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFKK23074452573 | 2023-08-17 | 2023-08-24 | Cust 1084 | CALCUTTA | Analytical - CHAKAN | CLOSED | COMMISSIONING | SULPHIDE ANALYSER | SULPHIDE 8810 Single Channel ANALYSER(8810 Sulphide) | FM SULPHIDE ANALYZER COMMISSIONING | The site visit was made . MOM done. | The site visit was made . MOM done. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFKK23075022574 | 2023-08-18 | 2023-08-18 | Cust 719 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | LocalHost is not accessible | We took access of the system and restart the local services so now the data is live on the server. | We took access of the system and restart the local services so now the data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFKK23080462575 | 2023-08-28 | 2023-09-04 | Cust 1074 | CALCUTTA | Analytical - CHAKAN | CLOSED | DIGITAL CONEECTIVITY | DATCON | DATCON(DATCON) | EUS data is not updated | Issue resolved | Job done. Report attached. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFKK23081812576 | 2023-08-30 | 2023-09-01 | Cust 681 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue. | We took the access to the system and found that the data is not communicating with the local system (checked on MODSCAN), hence as of now we are closing this complaint. | We took the access to the system and found that the data is not communicating with the local system (checked on MODSCAN), hence as of now we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFKK23083292577 | 2023-09-03 | 2023-09-05 | Cust 1085 | CALCUTTA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | CPU Display card for Multiparameter(AWA-CPU1) | AMC | Site visited, MOM done. | Site visited, MOM done. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFKK23084042578 | 2023-09-04 | 2023-09-06 | Cust 681 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue. | We took the access to the system and found that the CPCB server URL block in customers PC, after allowed the same URL now data is live on CPCB server. | We took the access to the system and found that the CPCB server URL block in customers PC, after allowed the same URL now data is live on CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFKK23086702579 | 2023-09-11 | 2023-09-18 | Cust 1086 | CALCUTTA | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | MULTIPARAMETER ANALYSER | FM make current output card(1) | MFA | Visited. Job done. Report made. | Visited. Job done. Report made. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK23090122580 | 2023-09-18 | 2023-09-25 | Cust 604 | CALCUTTA | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Electrode | Chemotrode 120(HT238 760) | pH electrode-Hamilton | Visited. Job done. MOM is attached. | Visited. Job done. MOM made. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK23091252581 | 2023-09-20 | 2023-09-22 | Cust 1087 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software Installation and CPCB Connectivity | Took access of the system and successfully updated DAS. | Took access of the system and successfully updated DAS. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFKK23091782582 | 2023-09-21 | 2023-09-22 | Cust 1088 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation needs to be done. | Took access of the system and successfully reinstalled the software. | Took access of the system and successfully reinstalled the software. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFKK23093192583 | 2023-09-25 | 2023-09-28 | Cust 1089 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Display board files are not generating on local PCB. | The issue has been resolved and now the Display Board files are generating on the Local PC. | The issue has been resolved and now the Display Board files are generating on the Local PC. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFKK23093532584 | 2023-09-25 | 2023-09-25 | Cust 1084 | CALCUTTA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC OF TOC, OIL, H2S, MULTIPARAMETER | AMC VISIT HAS BEEN DONE | AMC VISIT HAS BEEN MADE | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFKK23098462585 | 2023-10-02 | 2023-10-05 | Cust 600 | CALCUTTA | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | PH Input Card for Multiparameter(AWA-PH) | Change of ph sensor and calibration. | Ph sensor changed and calibrated. Report attached. | Ph sensor changed and calibrated. Report attached. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFKK23101892586 | 2023-10-09 | 2023-10-16 | Cust 1090 | CALCUTTA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | AQUADO Analyzer | AQUAMAX DOT(1) | Service engineer visit required for newly supplied CLO2 Analyser calibration and commissioning. | Arunava visited site ,MOM done. | Arunava visited site ,MOM done. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFKK23102712587 | 2023-10-10 | 2023-10-16 | Cust 564 | CALCUTTA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Smartpro-con) | Breakdown | RF raised. | RF raised. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFKK23102932588 | 2023-10-11 | 2023-10-12 | Cust 1091 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | DATA NOT GETTING TRANSFERED TO CPCB | there is internet side issue, we took the access to the system and now data is liv on CPCB server. | there is internet side issue, we took the access to the system and now data is liv on CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFKK23103802589 | 2023-10-13 | 2023-10-19 | Cust 1081 | CALCUTTA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | Flat cable of Multiparameter Display(1) | Calibration during AMC visit | Visit done and MOM is attached | Visit done and MOM is attached | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFKK23103822590 | 2023-10-13 | 2023-10-16 | Cust 1092 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB connectivity | We successfully did the CPC + UPPCB configuration now data is live on CPCB server for UPPCB server url is block in customers network, we have send a mail to customer regarding this issue, | We successfully did the CPC + UPPCB configuration now data is live on CPCB server for UPPCB server url is block in customers network, we have send a mail to customer regarding this issue, | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFKK23106172591 | 2023-10-18 | 2023-10-25 | Cust 604 | CALCUTTA | Analytical - CHAKAN | CLOSED | COMMISSIONING | Turbidity Meter | Optex(FMSC-P3) | Commissioning TSS | Visited. Job done. MOM is made. | Visited. Job done. MOM made. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK23115222592 | 2023-11-02 | 2023-11-06 | Cust 604 | CALCUTTA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Turbidity Meter | Optex turbidity sensor(1) | Non response of turbidity sensor as per process change | Visited. Job done. MOM is attached. | Visited. Job done. MOM made. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK23118382593 | 2023-11-09 | 2023-11-11 | Cust 591 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data upload issue. | we have changed the configuration of ETP Analyser as per the scanner and the data is live on the server. | we have changed the configuration of ETP Analyser as per the scanner and the data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFKK23118462594 | 2023-11-09 | 2023-11-12 | Cust 1093 | CALCUTTA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | 100 Series Control Valve | 100 Sereies Standard top flange and pressure balancing(107 STD Pressure Balanced) | Calibration and issues related with all new supply ( Control valve , analytical, mass flow meter etc) | Site is attended & report is attached | Site is attended & report is attached | YES | Warranty (Free) | Core | N |  | Arca - MIDC
AFKK23120392595 | 2023-11-16 | 2023-11-18 | Cust 1082 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | CPCB data upload issue. | here is PC hang issue, after restart the PC now data is live on CPCB server. | here is PC hang issue, after restart the PC now data is live on CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFKK23120522596 | 2023-11-16 | 2023-11-18 | Cust 603 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation needs to be done. | here is hardware side issue hence as of now we are closing this complaint. | here is hardware side issue hence as of now we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFKK23121942597 | 2023-11-20 | 2023-11-23 | Cust 1094 | CALCUTTA | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Installation Guideline of CX2000 Analyzer with Flow Meter and Black box. | Installation guideline given in details. Report attached. | Installation guideline given in details. Report attached. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFKK23122222598 | 2023-11-20 | 2023-11-23 | Cust 1078 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | WBPCB link for Alpine Distilleries Pvt. Ltd. | We have created the WBPCB link for Alpine Distilleries Pvt. Ltd | We have created the WBPCB link for Alpine Distilleries Pvt. Ltd | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFKK23131412599 | 2023-12-04 | 2023-12-10 | Cust 797 | CALCUTTA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | Peristaltic sampling pump motor for Multiparameter(AWA-PSPM) | Abnormal sound of sampling pump | Site visit made. Issue are attended & MOM is attached | Site visit made. Issue are attended & MOM is attached | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFKK23131592600 | 2023-12-05 | 2023-12-07 | Cust 1095 | CALCUTTA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | FM make pH card(1) | AMC activity | Site visited, MOM done. Resolved everything. | Site visited, MOM done. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFKK23132242601 | 2023-12-06 | 2023-12-10 | Cust 603 | CALCUTTA | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | MULTIPARAMETER ANALYSER | RS485 CARD for Multiparameter(AWA-RS485) | Version issue and subsequently the issues in MODBUS communication | Site visit made. Issues are attended & MOM is attached | Site visit made. Issues are attended & MOM is attached | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFKK23134012602 | 2023-12-11 | 2023-12-18 | Cust 1086 | CALCUTTA | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | TEPL-KMDA STP site | Visited. Job done. MOM made. | Visited. Job done. MOM made. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK23134732603 | 2023-12-12 | 2023-12-12 | Cust 1096 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Issue regarding data uptdation on BSPCB OCEMS portal | Data is now live on BSPCB + Local + CPCB server. | Data is now live on BSPCB + Local + CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFKK23134822604 | 2023-12-12 | 2023-12-15 | Cust 719 | CALCUTTA | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | MULTIPARAMETER ANALYSER | PH2 detector(1) | PH 2 ERROR. FLOW CELL DISCS NOT PROPER. | Offer given for required spares after doing the troubleshooting. | Offer given for required spares after doing the troubleshooting. | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFKK23138602605 | 2023-12-19 | 2024-01-04 | Cust 1074 | CALCUTTA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Electrode | Chemotrode 120(HT238 760) | Need PH sensor | The material reached to the site. The visit will be made accordingly. | The material reached to the site. The visit will be made accordingly. | EXTENDED | No Trip | Core | N |  | Analytical - CHAKAN
AFKK23140422606 | 2023-12-22 | 2023-12-27 | Cust 1097 | CALCUTTA | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | DO transmitter installation & calibration (Qty: 3 nos) for 01 day. Model No: DPI-201 Sensor: VisiWaterDO P Arc FC10 (Plastic). Kolkata?OA: 5524104887 Contact person- Mr. Murali Krishna ? 8096801283. Site: NINL Orissa. We need to attend this on Priority | The said DO meter has been commissioned remotely. The meter is under observation. | The said DO meter has been commissioned remotely. The meter is under observation. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFKK23149792607 | 2024-01-07 | 2024-01-15 | Cust 1074 | CALCUTTA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | pH Transmitter Aquamon | Ph Transmitter Aquamon 4000(PH-Trans-4000) | Analyzer is not powered on. | RF material received. Now visit is organised. | RF material received. Now visit is organised. | EXTENDED | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFKK23150612608 | 2024-01-08 | 2024-01-15 | Cust 650 | CALCUTTA | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | INSTALLATION GUIDE LINE. | Installation Guideline given in details. Report attached. | Installation Guideline given in details. Report attached. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFKK23152162609 | 2024-01-11 | 2024-01-15 | Cust 604 | CALCUTTA | Analytical - CHAKAN | CLOSED | COMMISSIONING | Turbidity Meter | Optex(FMSC-P3) | Commissioning CW Turbidity | Visited. Commissioned. MOM is attached. | Visited. Job done. MOM made. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK23152452610 | 2024-01-11 | 2024-01-16 | Cust 1098 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | DATCON | DATCON(DATCON) | DATCON configuration needs to be done. | here the data is not coming till the DATCON, hence as of now we are closing this complaint, and will reopen iot after the hardware side issue is resolved. | here the data is not coming till the DATCON, hence as of now we are closing this complaint, and will reopen iot after the hardware side issue is resolved. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFKK23155822611 | 2024-01-18 | 2024-01-23 | Cust 1083 | CALCUTTA | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | optex series(optex series) | commissioning | Re-commissioned. MOM is made. | Visited. Job done. MOM made. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK23157732612 | 2024-01-22 | 2024-01-25 | Cust 844 | CALCUTTA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | pH input card-FM of Multiparameter(1) | AMC | Visit made | Site visited, MOM done. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFKK23158672613 | 2024-01-24 | 2024-01-31 | Cust 547 | CALCUTTA | Analytical - CHAKAN | CLOSED | COMMISSIONING | Turbidity Meter | Krohne turbidity transmitter(1) | Commissioning for tubidity, OA - 5524103627 | Erection guidelines were given. Once the erection part is over, the customer will call us for commissioning over the phone. | Erection guidelines were given. Once the erection part is over, the customer will call us for commissioning over the phone. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFKK23159792614 | 2024-01-28 | 2024-01-30 | Cust 1099 | CALCUTTA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | CPU with LCD of Multiparameter Display(1) | AMC | Site visited,MOM done. | Site visited, MOM done. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFKK23160592615 | 2024-01-29 | 2024-02-02 | Cust 1099 | CALCUTTA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC | Site visited,MOM done. | Site visited,MOM done. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFKK23161002616 | 2024-01-30 | 2024-02-02 | Cust 623 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Reinstallation of software. | We took access to the system and successfully completed the reinstallation of software and now data is live on CPCB. | We took access to the system and successfully completed the reinstallation of software and now data is live on CPCB. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFKK23161102617 | 2024-01-30 | 2024-01-31 | Cust 1099 | CALCUTTA | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Smartpro Electronics Assembly(Smartpro-Elc-cbl) | Commissioning | Visit made, MOM done. | Site visited, MOM done. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK23161162618 | 2024-01-30 | 2024-07-23 | Cust 1090 | CALCUTTA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Chlorine Analyzer-MAC 100 | MAC 100(1) | Newly supplied CLO2 analyser is not responding and reading mismatches . | Visit Done. Calibration of all transmitters done. All are working OK. Few are showing slightly sluggish response as per MoM attached | Visit Done. Calibration of all transmitters done. All are working OK. Few are showing slightly sluggish response as per MoM attached | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFKK23162412619 | 2024-01-31 | 2024-02-02 | Cust 1100 | CALCUTTA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | pH input card-FM of Multiparameter(1) | AMC | Site visited,MOM done. | Site visited, MOM done. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFKK23162422620 | 2024-01-31 | 2024-02-02 | Cust 1099 | CALCUTTA | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | Aquamax system(1) | DO analyzer commissioning | Commissioned. | Site visited,commissioned. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK23163382621 | 2024-01-31 | 2024-01-31 | Cust 797 | CALCUTTA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc job | AMC JOB HAS BEEN DONE | AMC VISIT HAS BEEN MADE | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFKK23164952622 | 2024-02-02 | 2024-02-09 | Cust 1101 | CALCUTTA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Chlorine sensor | CHLORINE SENSOR(1) | Issue in CL sensor | Material dispatched via RF OA No:5524801446 /6704975 | Material dispatched via RF OA No:5524801446 /6704975 | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFKK23165452623 | 2024-02-04 | 2024-02-05 | Cust 1102 | CALCUTTA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | Xenon Lamp for Multiparameter(AWA-XLAMP) | Replacement of Newly received spares at site | Attended by Indranil Giri and resolved. MOM is attached. | Attended by Indranil Giri and resolved. MOM is attached. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFKK23165792624 | 2024-02-05 | 2024-02-07 | Cust 650 | CALCUTTA | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | CX-2000 Analyser commissioning. | Commissioning successfully completed. Report attached. Signed hard copy is kept at branch. | Commissioning successfully completed. Report attached. Signed hard copy is kept at branch. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFKK23167412625 | 2024-02-08 | 2024-02-09 | Cust 1103 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Addition of new device on CPCB | we have sent documents to customer. we will reopen this complaint once receipt all documents. | we have sent documents to customer. we will reopen this complaint once receipt all documents. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFKK23167992626 | 2024-02-09 | 2024-02-11 | Cust 1085 | CALCUTTA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | 12 V POWER SUPPLY CARD for Multiparameter(AWA-PS1) | reading fluctuation | Resolved,MOM done. | Site visited, MOM done. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFKK23168482627 | 2024-02-11 | 2024-02-11 | Cust 1084 | CALCUTTA | Analytical - CHAKAN | CLOSED | COMMISSIONING | OIL ANALYSER | PCX1 Microprocessor with Dispay Board(PCX1) | DISPLAY REPLACEMENT | DISPLAY REPLACED | OIL DISPLAY REPLACED | NO | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK23170142628 | 2024-02-14 | 2024-02-16 | Cust 797 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | they have the proxy address so data is not uploading to the CPCB server. | We have changed the Proxy IP to 10.18.86.69 and now the data started uploading to the server. | We have changed the Proxy IP to 10.18.86.69 and now the data started uploading to the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFKK23171022629 | 2024-02-16 | 2024-02-23 | Cust 1104 | CALCUTTA | Analytical - CHAKAN | CLOSED | AMC | Chlorine Analyzer-MAC 100 | MAC 100(1) | AMC activity | Site visited by Arunava. | Site visited by Arunava. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFKK23171362630 | 2024-02-17 | 2024-02-21 | Cust 1105 | CALCUTTA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Electrode | Polilyte HT 120(HT238 431) | POLILYTE PRO VP (PT 1000 VERSION) | Site visit made. MOM is attached | Site visit made. MOM is attached | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK23171942631 | 2024-02-18 | 2024-03-01 | Cust 1014 | CALCUTTA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Maintenance | Visit & job done | Site Visit done MOM is attached | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK23172642632 | 2024-02-19 | 2024-02-26 | Cust 1106 | CALCUTTA | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | COD BOD TSS analyser commissioning | Rajib Sen is in Site , commissioning job is going On. | Rajib Sen is in Site , commissioning job is going On. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFKK23173182633 | 2024-02-20 | 2024-02-27 | Cust 1107 | CALCUTTA | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Electrode | Chemotrode 120(HT238 760) | pH conductivity ORP TDS-- DM plant | Commissioned. MOM is attached | Visited. Commissioned. MOM made. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK23176992634 | 2024-02-26 | 2024-03-03 | Cust 1108 | CALCUTTA | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning job | Rajib Sen visited Site ,Mom done. | Rajib Sen visited Site ,Mom done. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFKK23177582635 | 2024-02-27 | 2024-03-05 | Cust 1107 | CALCUTTA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | issue in Trx and senseo | Material dispatched via-- REF : OUR RF OA NO.: 5524801569 DT: 28-FEB-2024 OUR RF INV. NO : FC/RF552411738 DT. 05-MAR-24 | Material dispatched via-- REF : OUR RF OA NO.: 5524801569 DT: 28-FEB-2024 OUR RF INV. NO : FC/RF552411738 DT. 05-MAR-24 | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFKK23183802636 | 2024-03-11 | 2024-09-26 | Cust 1109 | CALCUTTA | Analytical - CHAKAN | CLOSED | PRODUCT DESIGN FAILURE | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | READING NOT SHOWING | Replacement material sent vide docket no 210091871. Received in May 24. Installed & working fine. | Replacement material sent vide docket no 210091871. Received in May 24. Installed &amp;amp;amp; working fine. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFKK23183802636 | 2024-03-11 | 2024-09-26 | Cust 1109 | CALCUTTA | Analytical - CHAKAN | CLOSED | PRODUCT DESIGN FAILURE | Turbidity Meter | Optex(FMSC-P3) | SENSOR NOT WORKING | Replacement material sent vide docket no 210091871. Received in May 24. Installed & working fine. | Replacement material sent vide docket no 210091871. Received in May 24. Installed &amp;amp;amp; working fine. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFKK23183802636 | 2024-03-11 | 2024-09-26 | Cust 1109 | CALCUTTA | Analytical - CHAKAN | CLOSED | PRODUCT DESIGN FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(Ph-Trans-Smartpro) | SENSORS WERE NOT RESPONDING | Replacement material sent vide docket no 210091871. Received in May 24. Installed & working fine. | Replacement material sent vide docket no 210091871. Received in May 24. Installed &amp;amp;amp; working fine. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFKK23183802636 | 2024-03-11 | 2024-09-26 | Cust 1109 | CALCUTTA | Analytical - CHAKAN | CLOSED | PRODUCT DESIGN FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(Ph-Trans-Smartpro) | SENSORS WERE NOT RESPONDING | Replacement material sent vide docket no 210091871. Received in May 24. Installed & working fine. | Replacement material sent vide docket no 210091871. Received in May 24. Installed &amp;amp;amp; working fine. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFKK23183802636 | 2024-03-11 | 2024-09-26 | Cust 1109 | CALCUTTA | Analytical - CHAKAN | CLOSED | PRODUCT DESIGN FAILURE | Ph Electrode | Polilyte Plus 120(HT242 431) | SENSORS WERE NOT RESPONDING | Replacement material sent vide docket no 210091871. Received in May 24. Installed & working fine. | Replacement material sent vide docket no 210091871. Received in May 24. Installed &amp;amp;amp; working fine. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFKK23184372637 | 2024-03-11 | 2024-03-13 | Cust 1094 | CALCUTTA | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning of CX-2000 Analyser. | PRODUCT COMMISSIONING COMPLETED. REPORT ATTACHED. | PRODUCT COMMISSIONING COMPLETED. REPORT ATTACHED. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFKK23184522638 | 2024-03-12 | 2024-03-15 | Cust 797 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data Breakage issue on CPCB server. | We took access of your system and observed that someone had switched off the scanner supply and after restarting the same the data started uploading. | We took access of your system and observed that someone had switched off the scanner supply and after restarting the same the data started uploading. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFKK23186432639 | 2024-03-15 | 2024-03-19 | Cust 1094 | CALCUTTA | Analytical - CHAKAN | CLOSED | COMMISSIONING | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Black Box connectivity for Digital Services of CX-2000. | Job completed. Report attached. | Job completed. Report attached. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFKK23186462640 | 2024-03-15 | 2024-03-15 | Cust 1094 | CALCUTTA | Analytical - CHAKAN | CLOSED | DIGITAL CONEECTIVITY | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Digital Connectivity needs to be done. | Today we have completed the Black Box configuration and Cloud server connectivity for Emami Agrotech Limited and data is also live on the server. | Today we have completed the Black Box configuration and Cloud server connectivity for Emami Agrotech Limited and data is also live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFKK23186472641 | 2024-03-15 | 2024-03-24 | Cust 1107 | CALCUTTA | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Electrode | Polilyte Plus 225(HT242 432) | Ph Cond in RO | Visited. Job done. MOM is attached. | Visited. Commissioned. MOM made. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK23186832642 | 2024-03-16 | 2024-03-16 | Cust 797 | CALCUTTA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC VISIT | AMC VISITED | AMC VISIT HAS BEEN MADE | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFKK23188862643 | 2024-03-19 | 2024-03-24 | Cust 1086 | CALCUTTA | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | VISI pro DO(VISI pro DO) | 4 DO-- TEPL site | Commissioned | Visited. Commissioned. MOM made. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK23193252644 | 2024-03-26 | 2024-03-29 | Cust 1110 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | data not uploading on CPCB. | we took the access to your system and observed that the data communication cable is not properly connected to the local PC, hence the data is breaking on to the CPCB server. After reconnecting the same, now data is live on CPCB server. | Data is live now hence we are closing this compliant. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFKK23193412645 | 2024-03-26 | 2024-04-01 | Cust 1106 | CALCUTTA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | Flow cell for Multiparameter(AWA-FLOWCELL) | Calibration of ETP Multiparameter and Dissolve analyser | The Site has been visited, and Commissioning done. MOM made. | The Site has been visited, and Commissioning done. MOM made. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFKK23193422646 | 2024-03-26 | 2024-03-31 | Cust 1078 | CALCUTTA | Analytical - CHAKAN | CLOSED | DIGITAL CONEECTIVITY | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Reported CPCB data are frequently getting not available | Site visit made. MOM is attached | Site visit made. MOM is attached | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK23195892647 | 2024-03-29 | 2024-03-31 | Cust 650 | CALCUTTA | Analytical - CHAKAN | CLOSED | COMMISSIONING | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | ONLINE CONNECTIVITY FOR DIGITAL SERVICES OF CX2000 | Connectivity successfully established. | Connectivity successfully established. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFKK23196262648 | 2024-03-29 | 2024-04-01 | Cust 1108 | CALCUTTA | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | pH input card-FM of Multiparameter(1) | commissioning | (Model-CX2000-3932,A4TW130000) MOM done. | Site visited, MOM done. Commissioning sucessful. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFKK24001012649 | 2024-04-02 | 2024-04-05 | Cust 1111 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue. | we took the access to your system and completed CPCB connectivity for Stack_8_AFBC_2_36TPH gas and dust. Currently data is live on CPCB server for dust. | we took the access to your system and completed CPCB connectivity for Stack_8_AFBC_2_36TPH gas and dust. Currently data is live on CPCB server for dust. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFKK24002532650 | 2024-04-05 | 2024-04-06 | Cust 1099 | CALCUTTA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | ERROR IN MPA ANALYZER | The site visited, MOM done. | The site visited, MOM done. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFKK24002542651 | 2024-04-05 | 2024-04-06 | Cust 1100 | CALCUTTA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | FM make current output card(1) | AMC | The site visited, MOM done. | The site visited, MOM done. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFKK24004062652 | 2024-04-09 | 2024-04-11 | Cust 1112 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | New industry added to CPCB connectivity. | We took access to the system and successfully completed the configuration of LOCAL+CPCB+RTDMS | We took access to the system and successfully completed the configuration of LOCAL+CPCB+RTDMS | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFKK24005152653 | 2024-04-11 | 2024-04-12 | Cust 1113 | CALCUTTA | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning job | Site visited, Commissioned, MOM done. | Site visited, Commissioned, MOM done. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK24006582654 | 2024-04-15 | 2024-04-21 | Cust 915 | CALCUTTA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | CPU Display card for Multiparameter(AWA-CPU1) | calibration | Site visit made. MOM is attached | Site visit made. MOM is attached | NO | Paid | Core | N |  | Analytical - CHAKAN
AFKK24006912655 | 2024-04-16 | 2024-04-17 | Cust 1112 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB Server. | Job done. | Job done. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFKK24010362656 | 2024-04-23 | 2024-04-30 | Cust 1103 | CALCUTTA | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | VISI pro DO(VISI pro DO) | Optical DO | Visited. Commissioned. MOM is made. | Visited. Commissioned. MOM made. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK24011602657 | 2024-04-25 | 2024-04-27 | Cust 797 | CALCUTTA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | Flow cell for Multiparameter(AWA-FLOWCELL) | Multiple errors come in the analyzer, and for that analyzer was not serviced. | Resolved the issue,MOM done. | Site visited, MOM done. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFKK24016222658 | 2024-05-02 | 2024-05-02 | Cust 1081 | CALCUTTA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Maintenance | AMC Job done & analyzer is working healthy in service | Job Done MOM is attached | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFKK24018052659 | 2024-05-07 | 2024-05-10 | Cust 1107 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | reinstallation of software. | We took access to you system and successfully completed the reinstallation and configuration of software now data is live on LOCAL and CPCB server. | We took access to you system and successfully completed the reinstallation and configuration of software now data is live on LOCAL and CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFKK24019702660 | 2024-05-11 | 2024-05-13 | Cust 1114 | CALCUTTA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Turbidity Meter | Optex(FMSC-P3) | 0-500NTU cal | Visited. Job done. MOM is made. | Visited. Job done. MOM made. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFKK24019722661 | 2024-05-11 | 2024-05-13 | Cust 1114 | CALCUTTA | Analytical - CHAKAN | CLOSED | AMC | Conductivity Tranmitter Smartpro | Conductivity Electrode Cable(Cond-Elec-cbl) | 10 nos of ,water analyzer | Visited. Job done. MOM is attached. | Visited. Job done. MOM made. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFKK24021332662 | 2024-05-15 | 2024-05-17 | Cust 600 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Roter configuration needs to be done. | Today we have completed the Router configuration and now the data is live on the CPCB server. | Today we have completed the Router configuration and now the data is live on the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFKK24021542663 | 2024-05-15 | 2024-05-21 | Cust 636 | CALCUTTA | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 2 nos @ CX2000 | Visited. Commissioned. MOM made. SR NO-- 231023 23024 | Visited. Commissioned. MOM made. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK24024212664 | 2024-05-21 | 2024-05-26 | Cust 604 | CALCUTTA | Analytical - CHAKAN | CLOSED | COMMISSIONING | Turbidity Meter | Optex turbidity sensor(1) | 3000NTU ANALYZER | Sr no-- 23280490014S-- Attended. Commissioned. MOM is made. | Visited. Commissioned. MOM made. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK24024542665 | 2024-05-22 | 2024-05-23 | Cust 844 | CALCUTTA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | CPU with LCD of Multiparameter Display(1) | AMC | Site visited,MOM done | Site visited, MOM done. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFKK24025032666 | 2024-05-23 | 2024-05-27 | Cust 1099 | CALCUTTA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | FM make current output card(1) | AMC | Site visited,MOM done. | Site visited, MOM done. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFKK24025352667 | 2024-05-23 | 2024-05-27 | Cust 1099 | CALCUTTA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | DATCON | DATCON(DATCON) | RS485 -SPLITTER FALIURE | Site visited, MOM done. Resolved. | Site visited, MOM done. Resolved. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK24027142668 | 2024-05-27 | 2024-06-02 | Cust 1115 | CALCUTTA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC JOB | AMC VISITED | AMC VISIT HAS BEEN MADE | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFKK24027452669 | 2024-05-27 | 2024-05-29 | Cust 1100 | CALCUTTA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | pH input card-FM of Multiparameter(1) | ANALYZER RESET,DISPLAY ISSUE | Resolved over telephone. | Resolved over telephone. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFKK24030362670 | 2024-05-31 | 2024-08-15 | Cust 650 | CALCUTTA | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | Ph Electrode | Polilyte PRO 120(HT238 411) | Giving erratic reading. Supplied along with CX 2000. Need to give replacement. | Material dispatched. Invoice attached. | Material dispatched. Invoice attached. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFKK24033362671 | 2024-06-06 | 2024-06-10 | Cust 1081 | CALCUTTA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Telephonic | AMC Job Done | Site Visit done MOM is attached | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFKK24033372672 | 2024-06-06 | 2024-06-10 | Cust 1081 | CALCUTTA | Analytical - CHAKAN | CLOSED | SALES PROMOTION | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Telephonic | Display Changed, working fine | Site Visit done MOM is attached | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK24033632673 | 2024-06-06 | 2024-06-10 | Cust 650 | CALCUTTA | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Data not going to cloud although configuration and internet connection is okay. | Issue discussed with the customer . detailed report attached. | Issue discussed with the customer . detailed report attached. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFKK24033652674 | 2024-06-06 | 2024-06-09 | Cust 797 | CALCUTTA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC visit | Site visit made. MOM is attached | Site visit made. MOM is attached | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFKK24034182675 | 2024-06-07 | 2024-06-07 | Cust 1116 | CALCUTTA | DCS - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | MICROCON+ER SERIES | CONTROLLER MODULE (ER-201)(ER-201) | Data not updating on CPCB Server. | We took access to the system and there was time synchronization issue. Now issue has been resolved and data is live on CPCB Server. | We took access to the system and there was time synchronization issue. Now issue has been resolved and data is live on CPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFKK24034212676 | 2024-06-07 | 2024-06-10 | Cust 1107 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Therefore to restore the online data uploading to CPCB through enviroconnect server, system configuration for enviroconnect is required at newly installed PC. | Data is live on CPCB server. | Data is live on CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFKK24035902677 | 2024-06-12 | 2024-06-16 | Cust 1095 | CALCUTTA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC | Site visited,MOM done. | Site visited, MOM done. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFKK24036112678 | 2024-06-12 | 2024-06-13 | Cust 1080 | CALCUTTA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Sr No--2208147, KMDA Kanchrapara | Visited. Job done. MOM made. | Visited. Job done. MOM made. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFKK24036832679 | 2024-06-13 | 2024-06-19 | Cust 1117 | CALCUTTA | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Site servery for MPA | Visited. Job done. Report made. | Visited. Job done. Report made. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK24037362680 | 2024-06-14 | 2024-06-16 | Cust 1090 | CALCUTTA | Analytical - CHAKAN | CLOSED | COMMISSIONING | Chlorine Analyzer-MAC 100 | MAC 100(1) | Commissioning | Visit made, Discussion done. | Partly visit made. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK24042102681 | 2024-06-24 | 2024-06-26 | Cust 1118 | CALCUTTA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Transmitter Aquacon | Aquacon Electronics Assembly(CON-ELE-set-2000) | Calibration activity | Arunava visited the site and the MOM (Minutes of Meeting) was completed.e MOM (Minutes of Meeting) was completed. | Arunava visited the site and the MOM (Minutes of Meeting) was completed.e MOM (Minutes of Meeting) was completed. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFKK24043022682 | 2024-06-26 | 2024-06-29 | Cust 1119 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | BLACKBOX HANG ISSUE. | we took access of the BB and made some modifications so The data is live now. | we took access of the BB and made some modifications so The data is live now. | YES | No Trip | Core | Y |  | DCS - CHAKAN
AFKK24044272683 | 2024-06-28 | 2024-07-03 | Cust 1120 | CALCUTTA | Analytical - CHAKAN | CLOSED | AMC | TOC ANALYSER | FM 101 TOC Analyser(FM 101 TOC) | AMC | AMC visit made. | Site visited, MOM done. | EXTENDED | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFKK24047052684 | 2024-07-01 | 2024-07-10 | Cust 1121 | CALCUTTA | Steam System Division | CLOSED | PRODUCT FAILURE | CRS | FJP - FLASH JET PUPM - SIZE 40NB(040FJP) | CRM meter not getting powered on. Fuse getting burnt everytime | Kindly send the CRM to chakan | Kindly send the CRM to chakan | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFKK24047092685 | 2024-07-01 | 2024-07-02 | Cust 1088 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on CPCB server. | We took access to the system and now data is live on the CPCB server. | We took access to the system and now data is live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFKK24047482686 | 2024-07-02 | 2024-07-03 | Cust 1113 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation + CPCB connectivity needs to be done. | we took access of the system and completed reinstallation of enviroconnect software + CPCB connectivity, Now data is live on CPCB server. | we took access of the system and completed reinstallation of enviroconnect software + CPCB connectivity, Now data is live on CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFKK24047562687 | 2024-07-02 | 2024-07-05 | Cust 1122 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB configuration need to be done. | We took access to the system and successfully completed the CPCB Configuration and now data is live on CPCB Server. | We took access to the system and successfully completed the CPCB Configuration and now data is live on CPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFKK24048022688 | 2024-07-03 | 2024-07-04 | Cust 1123 | CALCUTTA | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | pH input card-FM of Multiparameter(1) | Commissioning | The site visited by Indranil Giri, MOM done. | The site visited by Indranil Giri, MOM done. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFKK24050952689 | 2024-07-09 | 2024-07-12 | Cust 1123 | CALCUTTA | Analytical - CHAKAN | CLOSED | DIGITAL CONEECTIVITY | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Digital connectivity | Digital connectivity has been complted . | Digital connectivity has been complted . | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFKK24052472690 | 2024-07-12 | 2024-07-14 | Cust 1084 | CALCUTTA | Analytical - CHAKAN | CLOSED | COMMISSIONING | Chlorine Analyzer-MAC 100 | MAC 100(1) | RECOMMISSIONING | The site visited, Resolved the issue. | The site visited, Resolved the issue. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFKK24052472690 | 2024-07-12 | 2024-07-14 | Cust 1084 | CALCUTTA | Analytical - CHAKAN | CLOSED | COMMISSIONING | Chlorine Analyzer-MAC 100 | MAC 100(1) | RECOMMISSIONING | The site visited, Resolved the issue. | The site visited, Resolved the issue. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFKK24052472690 | 2024-07-12 | 2024-07-14 | Cust 1084 | CALCUTTA | Analytical - CHAKAN | CLOSED | COMMISSIONING | Chlorine Analyzer-MAC 100 | MAC 100(1) | RECOMMISSIONING | The site visited, Resolved the issue. | The site visited, Resolved the issue. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFKK24052472690 | 2024-07-12 | 2024-07-14 | Cust 1084 | CALCUTTA | Analytical - CHAKAN | CLOSED | COMMISSIONING | Chlorine Analyzer-MAC 100 | MAC 100(1) | RECOMMISSIONING | The site visited, Resolved the issue. | The site visited, Resolved the issue. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFKK24052472690 | 2024-07-12 | 2024-07-14 | Cust 1084 | CALCUTTA | Analytical - CHAKAN | CLOSED | COMMISSIONING | Chlorine Analyzer-MAC 100 | MAC 100(1) | RECOMMISSIONING | The site visited, Resolved the issue. | The site visited, Resolved the issue. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFKK24052852691 | 2024-07-12 | 2024-07-17 | Cust 534 | CALCUTTA | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | CX2000 - install the online monitoring system supplied against PO No.1248/18. Please send your site engineer for the Installation, Commissioning & connectivity of the online monitoring system Site Address: ITC LIMITED PCPB EAST-Uluberia, Howrah West Bengal - 711303. Site person contact details: Mr. Abhishek Sukla: 7987435584 | Once the assembly work is finished, please send the assembly guidelines to the client. We will arrange a visit after the work is completed. | Once the assembly work is finished, please send the assembly guidelines to the client. We will arrange a visit after the erection work is completed. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFKK24053092692 | 2024-07-13 | 2024-07-16 | Cust 1074 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data uploaded issue. | Now the data is live on the server. | Now the data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFKK24053202693 | 2024-07-13 | 2024-07-18 | Cust 1074 | CALCUTTA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Turbidity Meter | Optex(FMSC-P3) | EUS service-- pH TSS data is not updated Issue in U3 SPM | Site visited, MOM done. | Site visited, resolved the issue. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFKK24057612694 | 2024-07-22 | 2024-07-26 | Cust 823 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | software installation + display board configuration needs to be done. | We have completed software reinstallation and display board configuration now data is available on CPCB + Local + display board. | We have completed software reinstallation and display board configuration now data is available on CPCB + Local + display board. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFKK24063632695 | 2024-07-30 | 2024-08-01 | Cust 823 | CALCUTTA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | Flow cell for Multiparameter(AWA-FLOWCELL) | Issues in Analyser | Site visit made. MOM is attached | Site visit made. MOM is attached | NO | Paid | Core | N |  | Analytical - CHAKAN
AFKK24066132696 | 2024-08-04 | 2024-08-10 | Cust 1093 | CALCUTTA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Tranmitter Polymetron | Conductivity Electrode Cable(cable-9125) | Installationis is done by customer as per guidance. Reading needs to be authenticated | Site visit done, conductivity transmitter is checked for calibration and found ok. It is found that during on line condition ( Thru ss flow thru) reading is 1100 microS/cm. It is decided that we will send PP flow thru chamber. RF initiated and despatched the material. Mass flow meter is found oversized still the meter is calibrated during zero and put on line and made few trial the mass flow rate is found OK after weight test. | Site visit made and reports is pasted in below equipement details | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFKK24069242697 | 2024-08-10 | 2024-08-15 | Cust 1120 | CALCUTTA | Analytical - CHAKAN | CLOSED | AMC | TOC ANALYSER | FM 101 TOC Analyser(FM 101 TOC) | AMc | Site visited, MOM done. | Site visited, MOM done. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFKK24070462698 | 2024-08-13 | 2024-08-16 | Cust 1074 | CALCUTTA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Turbidity Meter | Optex turbidity sensor(1) | EUS service-- ETP1 data issue | Visited. Issue resolved. MOM is attached | Visited. Job done. MOM made. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFKK24070892699 | 2024-08-13 | 2024-08-20 | Cust 1124 | CALCUTTA | Analytical - CHAKAN | CLOSED | AMC | Ph Tranmitter Smartpro | Smartpro Electronics Assembly(Smartpro-ele) | AMC JOB | AMC JOB HAS BEEN DONE | AMC VISIT HAS BEEN MADE | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFKK24070892699 | 2024-08-13 | 2024-08-20 | Cust 1124 | CALCUTTA | Analytical - CHAKAN | CLOSED | AMC | Ph Tranmitter Smartpro | Smartpro Electronics Assembly(Smartpro-ele) | AMC JOB | AMC JOB HAS BEEN DONE | AMC VISIT HAS BEEN MADE | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFKK24070892699 | 2024-08-13 | 2024-08-20 | Cust 1124 | CALCUTTA | Analytical - CHAKAN | CLOSED | AMC | Ph Tranmitter Smartpro | Smartpro Electronics Assembly(Smartpro-ele) | AMC JOB | AMC JOB HAS BEEN DONE | AMC VISIT HAS BEEN MADE | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFKK24070892699 | 2024-08-13 | 2024-08-20 | Cust 1124 | CALCUTTA | Analytical - CHAKAN | CLOSED | AMC | Ph Tranmitter Smartpro | Smartpro Electronics Assembly(Smartpro-ele) | AMC JOB | AMC JOB HAS BEEN DONE | AMC VISIT HAS BEEN MADE | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFKK24070892699 | 2024-08-13 | 2024-08-20 | Cust 1124 | CALCUTTA | Analytical - CHAKAN | CLOSED | AMC | Ph Tranmitter Smartpro | Smartpro Electronics Assembly(Smartpro-ele) | AMC JOB | AMC JOB HAS BEEN DONE | AMC VISIT HAS BEEN MADE | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFKK24072322700 | 2024-08-18 | 2024-08-22 | Cust 915 | CALCUTTA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | Flow cell for Multiparameter(AWA-FLOWCELL) | Reading is always at zero | Site visit made. Report is attached | Site visit made. Report is attached | NO | Paid | Core | N |  | Analytical - CHAKAN
AFKK24072772701 | 2024-08-19 | 2024-08-20 | Cust 974 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to APPCB and CPCB central portal server. | Data is now live on the CPCB and APPCB servers. | Data is now live on the CPCB and APPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFKK24073452702 | 2024-08-20 | 2024-08-22 | Cust 974 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | APPCB data uploading issue | Data is now live on the APPCB server. | Data is now live on the APPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFKK24073482703 | 2024-08-20 | 2024-08-22 | Cust 1125 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software re-installation | We have re - installed the software and completed the CPCB configuration. Now data is live on the CPCB server. | We have re - installed the software and completed the CPCB configuration. Now data is live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFKK24073752704 | 2024-08-20 | 2024-08-22 | Cust 1126 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue. | We have checked and currently, the stack Data is showing live on CPCB servers. | We have checked and currently, the stack Data is showing live on CPCB servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFKK24075092705 | 2024-08-22 | 2024-08-23 | Cust 1113 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Average time and date is not uploading on CPCB server. | We have taken up this issue with AIPL team. now issue has been resolved and Average time & date is uploading on CPCB server for stack_1_Boiler_120TPH. | We have taken up this issue with AIPL team. now issue has been resolved and Average time &amp;amp; date is uploading on CPCB server for stack_1_Boiler_120TPH. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFKK24076872706 | 2024-08-26 | 2024-08-29 | Cust 974 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not updating to APPCB server. | now the data is continuously reflecting to the APPCB server. Data availability is almost 100%. | now the data is continuously reflecting to the APPCB server. Data availability is almost 100%. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFKK24076882707 | 2024-08-26 | 2024-09-02 | Cust 1107 | CALCUTTA | Analytical - CHAKAN | CLOSED | COMMISSIONING | pH Transmitter Aquamon | Ph Transmitter Aquamon 2000(PH-Trans Aqu-2000) | ORP meters-- 4 nos | Visited. Job done. MOM is attached | Visited. Job done. MOM made. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK24077202708 | 2024-08-26 | 2024-08-29 | Cust 623 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | EnviroConnect software installation in new PC. | Data is live on CPCB server. | Data is live on CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFKK24077472709 | 2024-08-27 | 2024-08-29 | Cust 1099 | CALCUTTA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | Flowcell disc of Multiparameter(1) | AMC | Site visited, MOM done. | Site visited, MOM done. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFKK24078372710 | 2024-08-28 | 2024-08-29 | Cust 1099 | CALCUTTA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | pH input card-FM of Multiparameter(1) | AMC | Site visited, MOM done. | Site visited, MOM done. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFKK24078872711 | 2024-08-29 | 2024-08-30 | Cust 844 | CALCUTTA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | Flow cell for Multiparameter(AWA-FLOWCELL) | AMC | Visit made. | Site visited, MOM done. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFKK24079312712 | 2024-08-29 | 2024-09-01 | Cust 1014 | CALCUTTA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | 24 VDC SMPS (1) | AMC | Site visit made. MOM is attached | Site visit made. MOM is attached | NO | Paid | Core | N |  | Analytical - CHAKAN
AFKK24079332713 | 2024-08-29 | 2024-09-01 | Cust 1100 | CALCUTTA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | CPU Display card for Multiparameter(AWA-CPU1) | AMC activity | Visit made, MOM done. | Site visited, MOM done. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFKK24084002714 | 2024-09-05 | 2024-09-06 | Cust 1112 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data on updating on CPCB Server. | We took access to the system and found Enviroconnect services were stop now data is live on CPCB Server. | We took access to the system and found Enviroconnect services were stop now data is live on CPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFKK24088912715 | 2024-09-16 | 2024-09-16 | Cust 1113 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera configuration needs to be done. | Camera connectivity completed now camera is streaming live on CPCB server. | Camera connectivity completed now camera is streaming live on CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFKK24090172716 | 2024-09-18 | 2024-09-19 | Cust 1113 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB Camera streaming issue. | we took access to your system and now the issue has been resolved and the camera is streaming live on CPCB server. | we took access to your system and now the issue has been resolved and the camera is streaming live on CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFKK24092572717 | 2024-09-23 | 2024-09-25 | Cust 1102 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | ETP and Boiler data not updating on CPCB Server. | We took access to the system and now data is live on CPCB Server. | We took access to the system and now data is live on CPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFKK24092712718 | 2024-09-23 | 2024-09-27 | Cust 1103 | CALCUTTA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Electrode | Chemotrode 120(HT238 760) | ETP EQMs | Visited. Job done. Report is attached. | Visited. Job done. Report made. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK24093402719 | 2024-09-25 | 2024-09-27 | Cust 1089 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB configuration needs to be done. | We have completed CPCB + Display Board configuration and now data is live on CPCB portal also showing on Display Board. | We have completed CPCB + Display Board configuration and now data is live on CPCB portal also showing on Display Board. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFKK24093442720 | 2024-09-25 | 2024-09-30 | Cust 1127 | CALCUTTA | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Smartpro Electronics Assembly(Smartpro-ele) | PH CONDUCTIVITY TURBIDITY & HARDNEDD | Commissioned. MOM is attached. | Commissioned. MOM is attached. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK24098092721 | 2024-10-02 | 2024-10-05 | Cust 719 | CALCUTTA | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | Flow cell for Multiparameter(AWA-FLOWCELL) | Spares Replacement and Commissioning of Product. | Spares replaced and calibrated the equipment. | Spares replaced and calibrated the equipment. | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFKK24099532722 | 2024-10-07 | 2024-10-14 | Cust 1128 | CALCUTTA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Aqua4Trans | Aqua4Trans(1) | pH reading Issue | Spoke to the client, and they will arrange a phone call from the Site for detail analysis m one we get a call, further complain will be logged. Presently the complain has been closed. | Spoke to the client, and they will arrange a phone call from the Site for detail analysis m one we get a call, further complain will be logged. Presently the complain has been closed. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFKK24099652723 | 2024-10-07 | 2024-10-10 | Cust 1129 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Re- installation of Eversense software . | We accessed the system and completed the configuration for CPCB. However, there was a hardware issue with the USB isolator converter.it was not communicating with local PC. | We accessed the system and completed the configuration for CPCB. However, there was a hardware issue with the USB isolator converter.it was not communicating with local PC. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFKK24100372724 | 2024-10-08 | 2024-10-13 | Cust 603 | CALCUTTA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | CPU with LCD of Multiparameter Display(1) | CPCB connectivity is down, NO MODSCAN data also | Site visit made and MOM is attached | Site visit made and MOM is attached | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFKK24101832725 | 2024-10-11 | 2024-10-17 | Cust 798 | CALCUTTA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Conductivity Electrode | CONDUCELL 2DC-PG-120(HT237 610) | pH and TDS Meter at our Kolkata site. Site :Provence trading & Industries pvt ltd Sankrail Indil park Dhulagarh, Bhagabhatipur Howrah 711313 | Site visit made. MOM is attached | Site visit made. MOM is attached | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK24104262726 | 2024-10-16 | 2024-10-18 | Cust 1074 | CALCUTTA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | EQMS of ETS#2 | ISSUE RESOLVED.EMAIL SENT. | Issue resolved. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFKK24104272727 | 2024-10-16 | 2024-10-22 | Cust 1117 | CALCUTTA | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | MULTIPARAMETER ANALYSER | mA Output card for Multiparameter(AWA-4-20mA) | EQMs for KMDA STP | Visited. Job done. Report send. | Visited. Job done. MOM made. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK24107772728 | 2024-10-23 | 2024-10-30 | Cust 823 | CALCUTTA | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | MULTIPARAMETER ANALYSER | CPU with LCD of Multiparameter Display(1) | Putting all newly received spares in CPU-1 analyser | Site visit made. MOM is attached | Site visit made. MOM is attached | NO | Paid | Core | N |  | Analytical - CHAKAN
AFKK24108492729 | 2024-10-25 | 2024-10-28 | Cust 823 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | cpcb data mismatch issue. | We accessed the system and checked the data using Modscan. The data received by Modscan is being uploaded to the CPCB server. | We accessed the system and checked the data using Modscan. The data received by Modscan is being uploaded to the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFKK24108802730 | 2024-10-25 | 2024-10-31 | Cust 1094 | CALCUTTA | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Showing Zero value in Sample. Need to check and calibrate. | Sample filter chocked. system working properly after cleaning. | Sample filter chocked. system working properly after cleaning. Report attached. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFKK24114782731 | 2024-11-09 | 2024-11-14 | Cust 1120 | CALCUTTA | Analytical - CHAKAN | CLOSED | AMC | TOC ANALYSER | Reaction Vessel Body(001?016A) | AMC | Site visited, MOM done. | Site visited, MOM done. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFKK24117452732 | 2024-11-13 | 2024-11-19 | Cust 1130 | CALCUTTA | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | MULTIPARAMETER ANALYSER | Mother board-PCX2 of Multiparameter Analyzer(1) | E & C activity | Site visit made. MOM is attached | Site visit made. MOM is attached | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK24117452732 | 2024-11-13 | 2024-11-19 | Cust 1130 | CALCUTTA | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | Aqua4Trans | Aqua4Trans(1) | E & C activity of Dissolve Oxygen | Site visit made. MOM is attached | Site visit made. MOM is attached | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK24117452732 | 2024-11-13 | 2024-11-19 | Cust 1130 | CALCUTTA | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | Aqua4Trans | Aqua4Trans(1) | E & C activity of Dissolve Oxygen | Site visit made. MOM is attached | Site visit made. MOM is attached | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK24117452732 | 2024-11-13 | 2024-11-19 | Cust 1130 | CALCUTTA | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | Ph Tranmitter Smartpro | Smartpro Electronics Assembly(Smartpro-ele) | E & C activity | Site visit made. MOM is attached | Site visit made. MOM is attached | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK24117452732 | 2024-11-13 | 2024-11-19 | Cust 1130 | CALCUTTA | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | MULTIPARAMETER ANALYSER | Mother board-PCX2 of Multiparameter Analyzer(1) | E & C activity | Site visit made. MOM is attached | Site visit made. MOM is attached | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK24117882733 | 2024-11-14 | 2024-11-20 | Cust 1084 | CALCUTTA | Analytical - CHAKAN | CLOSED | AMC | SULPHIDE ANALYSER | Vessel(35911001300) | CAMC | CAMC visit done. | CAMC visit made. | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFKK24120802734 | 2024-11-20 | 2024-11-24 | Cust 797 | CALCUTTA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | Cleaning Pump(1) | AMC | Visit made. | Visit made. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFKK24120812735 | 2024-11-20 | 2025-02-28 | Cust 1084 | CALCUTTA | Analytical - CHAKAN | CLOSED | SALES PROMOTION | SULPHIDE ANALYSER | Vessel(35911001300) | HCL PUMP FAILED FOR ALL 03 ANALYZERS | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210132449 LR Dt - 12-FEB-2025 | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210132449 LR Dt - 12-FEB-2025 | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFKK24121632736 | 2024-11-22 | 2024-11-29 | Cust 1117 | CALCUTTA | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | CX2000+CPCB+EUS | Visited. Commissioned. MOM is attached. | Visited. Commissioned. MOM made. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK24121692737 | 2024-11-22 | 2024-12-24 | Cust 1117 | CALCUTTA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | BLACK BOX commissioning | Replacemnt MARK EXPRESS PRIVATE LIMITED vide LR No - 210125034 LR Dt - 16-DEC-2024 | Replacemnt MARK EXPRESS PRIVATE LIMITED vide LR No - 210125034 LR Dt - 16-DEC-2024 | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFKK24125022738 | 2024-11-28 | 2024-11-29 | Cust 603 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation + CPCB connectivity needs to be done | we took access to your system and completed software reinstallation + CPCB connectivity, now data is live on CPCB server and the same is verified by you. | we took access to your system and completed software reinstallation + CPCB connectivity, now data is live on CPCB server and the same is verified by you. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFKK24125872739 | 2024-11-29 | 2024-12-02 | Cust 1084 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data uploaded issue. | As of now we are closing this complaint and will reopen it again once we get remote access of the system from coustomer end. | As of now we are closing this complaint and will reopen it again once we get remote access of the system from coustomer end. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFKK24128682740 | 2024-12-03 | 2024-12-09 | Cust 1131 | CALCUTTA | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | MULTIPARAMETER ANALYSER | CPU Display card for Multiparameter(AWA-CPU1) | E & C activity | Site visited, MOM done. | Site visited, MOM done. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK24128682740 | 2024-12-03 | 2024-12-09 | Cust 1131 | CALCUTTA | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | MULTIPARAMETER ANALYSER | CPU Display card for Multiparameter(AWA-CPU1) | E & C activity | Site visited, MOM done. | Site visited, MOM done. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK24131752741 | 2024-12-09 | 2024-12-12 | Cust 954 | CALCUTTA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | pH input card-FM of Multiparameter(1) | AMC | AMC done. | Site visited, MOM done. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFKK24135862742 | 2024-12-18 | 2024-12-22 | Cust 1080 | CALCUTTA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | DO Transmitters | VISI pro DO(VISI pro DO) | Visiwater DO P ARC FC10-- issue in sensor-- KMDA Kanchrapara | Visited. Checked. Report made. | Visited. Job done. MOM made. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK24135872743 | 2024-12-18 | 2024-12-22 | Cust 1117 | CALCUTTA | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | RS485 CARD for Multiparameter(AWA-RS485) | EUS+CPCB connectivity | Visited. EUS+CPCB connectivity completed. MOM is attached. | Visited. Commissioned. MOM made. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK24136572744 | 2024-12-20 | 2024-12-23 | Cust 823 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data mismatch issue. | As discussed, today we took access to your system and checked the data on MODSCAN and as per calculations (mA counts), same data uploading on the CPCB server for COD parameter. | As discussed, today we took access to your system and checked the data on MODSCAN and as per calculations (mA counts), same data uploading on the CPCB server for COD parameter. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFKK24138802745 | 2024-12-23 | 2024-12-30 | Cust 1086 | CALCUTTA | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | pH input card-FM of Multiparameter(1) | OCEMS at KMDA Chinchura | Visited. Job done. Report made. | Visited. Job done. Report made. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK24139122746 | 2024-12-24 | 2024-12-26 | Cust 797 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | RTDMs data uploaded issue. | We have made changes in the DeviceID id from the server side so now the Data is live on the Central srever. | We have made changes in the DeviceID id from the server side so now the Data is live on the Central srever. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFKK24139212747 | 2024-12-24 | 2024-12-28 | Cust 1086 | CALCUTTA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | DO Transmitters | VISI pro DO(VISI pro DO) | DO meter at Basin#2 at Chinsurah STP | Visited. Job done. Report made. | Visited. Job done. Report made. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK24139312748 | 2024-12-24 | 2024-12-30 | Cust 1132 | CALCUTTA | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | MULTIPARAMETER ANALYSER | Flowcell disc of Multiparameter(1) | E & C activity | Commissioned and MOM is attached | Commissioned and MOM is attached | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK24140552749 | 2024-12-27 | 2024-12-30 | Cust 1094 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | BLACK BOX CONFIGURATION FOR CPCB DATA TRANSFER. | Connectivity completed . Mail communication attached. | Connectivity completed . Mail communication attached. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFKK24146572750 | 2025-01-05 | 2025-01-10 | Cust 604 | CALCUTTA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Turbidity Meter | Optex turbidity sensor(1) | Turbidity analyser (0-3000)NTU is not working properly & showing the in correct value of the turbidity of the raw water. | Visited. Job done. Report is made. | Visited. Job done. Report is made. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK24147072751 | 2025-01-06 | 2025-01-10 | Cust 604 | CALCUTTA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | pH Transmitter Aquamon | Ph Transmitter Aquamon 4000(PH-Trans-4000) | pH meter for clear water application | Visited. Job done. Report is made. | Visited. Job done. Report is made. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK24147472752 | 2025-01-07 | 2025-01-10 | Cust 1117 | CALCUTTA | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | CPU Display card for Multiparameter(AWA-CPU1) | Re-commissioning of MPA & FMBB against supplied RF materials. | Visited. Job done. MOM made. | Visited. Job done. MOM made. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK24147592753 | 2025-01-07 | 2025-01-11 | Cust 1115 | CALCUTTA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc visit | amc visited | AMC VISIT HAS BEEN MADE | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFKK24148062754 | 2025-01-08 | 2025-02-02 | Cust 1117 | CALCUTTA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Cloud server data uploaded issue. | Replaced FMBB, mother board of MPA | Replaced FMBB, mother board of MPA | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFKK24148172755 | 2025-01-08 | 2025-01-10 | Cust 1132 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Need to change the scanner address from 5 to 6 for Stack 5_300 TPD_DRI-5 | We changed the address from the software at variable modify and now correct data is updating on CPCB server. | We changed the address from the software at variable modify and now correct data is updating on CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFKK24148322756 | 2025-01-08 | 2025-01-13 | Cust 1133 | CALCUTTA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Electrode | Polilyte Plus 120(HT242 431) | Issue in PH ELECTRODE-- POLILYTE PLUS PHI S8 120 (OLD PART NO. HT238450) Sr no--5437,5444,5455,5458 | Visited. Job done. Report made. | Visited. Job done. Report made. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK24150412757 | 2025-01-14 | 2025-01-17 | Cust 844 | CALCUTTA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | Flow cell for Multiparameter(AWA-FLOWCELL) | AMC-4TH VISIT | Site visited, MOM done. | Site visited, MOM done. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFKK24150512758 | 2025-01-14 | 2025-01-17 | Cust 1099 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data uploaded issue. | We took access of the BB and found a time sync issue and now we have synchronised the BB time with thr real-time so now the Data is live on the server. | We took access of the BB and found a time sync issue and now we have synchronised the BB time with thr real-time so now the Data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFKK24150752759 | 2025-01-14 | 2025-01-19 | Cust 1081 | CALCUTTA | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning | Job done | job completed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK24150782760 | 2025-01-14 | 2025-01-17 | Cust 1100 | CALCUTTA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | Flow cell for Multiparameter(AWA-FLOWCELL) | AMC-FINAL VISIT | site visited, mom done. | Site visited, MOM done. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFKK24151162761 | 2025-01-15 | 2025-02-19 | Cust 650 | CALCUTTA | Analytical - CHAKAN | CLOSED | DIGITAL CONEECTIVITY | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Cloud server data upload issue. | Data is now live on the Cloud srever. | Data is now live on the Cloud srever. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFKK24151722762 | 2025-01-16 | 2025-01-20 | Cust 650 | CALCUTTA | Analytical - CHAKAN | CLOSED | DIGITAL CONEECTIVITY | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | FM Cloud connectivity of CX-2000 through Teltonika Router. | Issue resolved. Cloud connectivity completed. Refer mail communication. | Issue resolved. Cloud connectivity completed. Refer mail communication. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFKK24151732763 | 2025-01-16 | 2025-01-17 | Cust 1123 | CALCUTTA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | PH2 Filter(1) | TSS/COD READING HANGED | Site visited, MOM done. | Site visited, MOM done. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK24152172764 | 2025-01-17 | 2025-01-20 | Cust 1134 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | DATCON | DATCON(DATCON) | CPCB data upload issue. | Now data is live on the CPCB server Stack 1 & 2. | Now data is live on the CPCB server Stack 1 &amp;amp; 2. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFKK24153992765 | 2025-01-21 | 2025-01-25 | Cust 1081 | CALCUTTA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC, Calibration | Maintenance activity & calibration done | job completed | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFKK24155082766 | 2025-01-23 | 2025-02-28 | Cust 534 | CALCUTTA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Electrode | Polilyte PRO 120(HT238 411) | Customer is experiencing technical difficulties at our ITC, Uluberia site. The specific issues are as follows: pH meter of Equalization tank-1 pH meter of Equalization tank-2 DO meter of both Aeration tank Online Monitoring System Kindly arrange for site visit at the earliest. Contact Details : Mr. Manik ( ITC, Uluberia) - 7908673463 Site Address: ITC, Limited PCPB East ? Uluberia. Uluberia, Howrah, 711303. India. | Apurba Das visited the site guidelines provided | Apurba Das visited the site guidelines provided | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFKK24155782767 | 2025-01-24 | 2025-01-30 | Cust 1132 | CALCUTTA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | Peristaltic sampling pump motor for Multiparameter(AWA-PSPM) | Failure of mother board and sample pump | Site visit made and identified the spares for making it operational and shared with division and it is despatched | Site visit made and identified the spares for making it operational and shared with division and it is despatched | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFKK24155842768 | 2025-01-25 | 2025-01-27 | Cust 1117 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | New BB Commissioning and CPCB + Cloud server connectivity needs to be done. | we have completed the new FM BB configuration + CPCB and cloud server connectivity for 9.5 MLD Sewage Treatment Plant and the data was live | we have completed the new FM BB configuration + CPCB and cloud server connectivity for 9.5 MLD Sewage Treatment Plant and the data was live | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFKK24156202769 | 2025-01-25 | 2025-02-02 | Cust 1117 | CALCUTTA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MIL card of Multiparameter(1) | MPA not working | Visited. Job done. MOM made. | Visited. Job done. MOM made. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK24164562770 | 2025-02-07 | 2025-03-12 | Cust 1094 | CALCUTTA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | Peristaltic sampling pump motor for Multiparameter(AWA-PSPM) | Pump is not rotating although Motor is rotating. Motor Rack and pinion damaged. Need to send one Motor urgently. Photos shared via Wapp. | Replacement material sent. Refer attached documents. | Replacement material sent. Refer attached documents. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFKK24164752771 | 2025-02-08 | 2025-02-14 | Cust 1135 | CALCUTTA | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Smartpro-con) | Commissioning of ORP, Turbidity, pH, TDS analysers, mag meters, PT's | installation guidelines provided. one each pH, ORP,TDS, Turbidity and Mag flowmeter commissioned | installation guidelines provided. one each pH, ORP,TDS, Turbidity and Mag flowmeter commissioned | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFKK24166282772 | 2025-02-11 | 2025-02-13 | Cust 1132 | CALCUTTA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | 12 V POWER SUPPLY CARD for Multiparameter(AWA-PS1) | DISPLAY HANGED | The site visited by Rajib, resolved the issue, job report sent. | The site visited by Rajib, resolved the issue, job report sent. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFKK24166492773 | 2025-02-12 | 2025-02-19 | Cust 1102 | CALCUTTA | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Value zero | Resolved. | Resolved. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFKK24166512774 | 2025-02-12 | 2025-02-19 | Cust 526 | CALCUTTA | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | Conductivity Transmitter Aquacon | Conductivity Transmitter Aquamon 2000(CON-TR-2000) | Currently at JSR site near Malda West Bengal , our team found one Ph transmitter display is not getting on. Power supply is coming up to PH transmitter, but display is not getting powered on. | Email the client to send the meter to our Pune facility for repairs. | Email the client to send the meter to our Pune facility for repairs. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFKK24166572775 | 2025-02-12 | 2025-02-14 | Cust 1132 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | FMBB configuration needs to be done. | we took access to the FMBB, and completed the configuration. now data is coming till the FMBB. | we took access to the FMBB, and completed the configuration. now data is coming till the FMBB. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFKK24168972776 | 2025-02-18 | 2025-02-21 | Cust 1136 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to CPCB server. | We have re-installed the software with an updated version and data is now live on the Local and CPCB servers. Please verify the readings and revert in case of any query. | We have re-installed the software with an updated version and data is now live on the Local and CPCB servers. Please verify the readings and revert in case of any query. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFKK24169882777 | 2025-02-20 | 2025-02-23 | Cust 712 | CALCUTTA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | pH/TSS reading error. Need calibration and procedure explanation. | VISIT DONE. REPORT ATTACHED FOR REFERENCE. | VISIT DONE. REPORT ATTACHED FOR REFERENCE. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFKK24172792778 | 2025-02-26 | 2025-03-03 | Cust 604 | CALCUTTA | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4Trans(1) | pH meter for treated water application of PHE- WTP Model-- A4TW01000 | Visited. Commissioned. MOM made. | Visited. Commissioned. MOM made. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK24173362779 | 2025-02-26 | 2025-03-03 | Cust 1103 | CALCUTTA | Analytical - CHAKAN | CLOSED | DIGITAL CONEECTIVITY | Ph Electrode | Polyclave VP 225(HT238 456) | EUS + dashboard for 2CEMs+1 OCEMs | Visited. Job done. MOM made. | Visited. Job done. MOM made. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFKK24173372780 | 2025-02-26 | 2025-03-02 | Cust 1127 | CALCUTTA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Turbidity Meter | Optex(FMSC-P3) | Calibration | Site visit is made and found earlier suggested erection modification is not done. Agreed by customer to do the same and inform to schedule for next visit | Site visit is made and found earlier suggested erection modification is not done. Agreed by customer to do the same and inform to schedule for next visit | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFKK24173592781 | 2025-02-26 | 2025-02-28 | Cust 1135 | CALCUTTA | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | Aqua4Trans | 4-20 mA input card of Aqua4trans(1) | Erection guidelines needed | Installation parameter provided and discussed | Installation parameter provided and discussed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFKK24174062782 | 2025-02-27 | 2025-03-03 | Cust 604 | CALCUTTA | Analytical - CHAKAN | CLOSED | COMMISSIONING | Turbidity Meter | Thermo series(Thermo series) | Turbidity Analyzer-- Model---FMTC-mi100 | Visited. Commissioned. MOM made. | Visited. Commissioned. MOM made. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK24174622783 | 2025-02-28 | 2025-03-04 | Cust 823 | CALCUTTA | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | MULTIPARAMETER ANALYSER | LCD of Multiparameter Display(1) | Display is flickering | Site visit is made and MOM is attached | Site visit is made and MOM is attached | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK24176042784 | 2025-03-01 | 2025-03-07 | Cust 1084 | CALCUTTA | Analytical - CHAKAN | CLOSED | AMC | SULPHIDE ANALYSER | Power supply board(359110,20XXX) | CAMC | Visit made, MOM done. | Visit made, MOM done. | EXTENDED | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK24176152785 | 2025-03-01 | 2025-03-04 | Cust 1058 | CALCUTTA | Analytical - CHAKAN | CLOSED | AMC | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Calibration & maintenance | job done | job completed | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFKK24176962786 | 2025-03-04 | 2025-03-10 | Cust 1086 | CALCUTTA | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | MPA--COD, BOD, TSS, pH, CH4, NO3 | Visited. Commissioned. MOM made. | Visited. Commissioned. MOM made. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK24177872787 | 2025-03-05 | 2025-03-07 | Cust 1086 | CALCUTTA | Analytical - CHAKAN | CLOSED | DIGITAL CONEECTIVITY | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Digital connectivity needs to be done. | Completed the cloud server connectivity for ETP and Data is live on the server. | Completed the cloud server connectivity for ETP and Data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFKK24178772788 | 2025-03-07 | 2025-03-21 | Cust 1137 | CALCUTTA | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Commissioning of PH sensor at RO water outlet | Commissioning of ph sensor completed successfully. Report attached. | Commissioning of ph sensor completed successfully. Report attached. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFKK24178782789 | 2025-03-07 | 2025-03-07 | Cust 1132 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB connectivity needs to be done, | We have completed the CPCB connectivity for the ETP station and now the Data is live on the CPCB server. | We have completed the CPCB connectivity for the ETP station and now the Data is live on the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFKK24178852790 | 2025-03-07 | 2025-03-10 | Cust 1074 | CALCUTTA | Analytical - CHAKAN | CLOSED | DIGITAL CONEECTIVITY | MULTIPARAMETER ANALYSER | FM make pH card(1) | ETP#2 EUS data connectivity issue | Visited. Job done. MOM made. | Visited. Job done. MOM made. | EXTENDED | Paid | Core | N |  | Analytical - CHAKAN
AFKK24179082791 | 2025-03-08 | 2025-06-16 | Cust 1084 | CALCUTTA | Analytical - CHAKAN | CLOSED | SALES PROMOTION | SULPHIDE ANALYSER | Vessel(35911001300) | NEED A FEW ELECTRONIC CARDS FOR REPAIRING OF SULPHIDE ANALYZER | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210145237 LR Dt - 12-JUN-2025 | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210145237 LR Dt - 12-JUN-2025. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFKK24179122792 | 2025-03-09 | 2025-03-12 | Cust 1099 | CALCUTTA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | FM make pH card(1) | AMC | Site visited,MOM done. | Site visited, MOM done. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFKK24179452793 | 2025-03-10 | 2025-03-13 | Cust 1112 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB server. | We took access to the system and observed the services were stopped after restarting the services data was live. | We took access to the system and observed the services were stopped after restarting the services data was live. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFKK24179862794 | 2025-03-10 | 2025-03-17 | Cust 604 | CALCUTTA | Analytical - CHAKAN | CLOSED | COMMISSIONING | Turbidity Meter | Optex turbidity sensor(1) | Turbidity meter FMSC T3 | Visited. Commissioned. MOM made. | Visited. Commissioned. MOM made. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK24179942795 | 2025-03-10 | 2025-03-12 | Cust 1099 | CALCUTTA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | CPU Display card for Multiparameter(AWA-CPU1) | AMC-Final visit | Site visited,MOM done. | Site visited, MOM done. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFKK24181302796 | 2025-03-12 | 2025-04-02 | Cust 534 | CALCUTTA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | Xenon Lamp Power supply for Multiparameter(AWA-XLAMPPS) | XENON LAMP 1 KV POWER SUPPLY BOARD FAULTY. OUTPUT 1000 VOLT NOT GENERATING. NEED URGENT REPLACEMENT. END USER IS ITC (CUSTOMER CODE 9379, PERSONAL CARE PRODUCTS FACTORY, BIRSHIBPUR, HOWRAH, WEST BENGAL). SPARES NEED TO SENT AT SITE. | Spares changed and Analyzer recommissioned. | Spares changed and Analyzer recommissioned. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFKK24186242797 | 2025-03-24 | 2025-03-30 | Cust 1127 | CALCUTTA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | THERMO SERIES | AV38 TRANSMITTER(1) | Calcium Hardness analyser | Site visit made, report sent to the customer | Site visit made, report sent to the customer | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFKK24186362798 | 2025-03-25 | 2025-03-31 | Cust 1120 | CALCUTTA | Analytical - CHAKAN | CLOSED | AMC | TOC ANALYSER | U.V. Lamp Power Supply 110V(001?205A) | TOC SERVICING | Visit made, MOM done. | Site visited, MOM done. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFKK24187872799 | 2025-03-27 | 2025-04-08 | Cust 954 | CALCUTTA | Analytical - CHAKAN | CLOSED | SALES PROMOTION | Aqua4Trans | Display card of Aqua4Trans(1) | Display LED need to be replaced | The RF material has been sent to to Client. | The RF material has been sent to to Client. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFKK24189902800 | 2025-03-31 | 2025-04-03 | Cust 1134 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software re-installation | We have reinstalled the software with an upgraded version and completed Local and CPCB configuration. Also we have updated the Enviroconnect DAS version. Now the data is live on both servers | We have reinstalled the software with an upgraded version and completed Local and CPCB configuration. Also we have updated the Enviroconnect DAS version. Now the data is live on both servers | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFKK25001282801 | 2025-04-04 | 2025-04-11 | Cust 910 | CALCUTTA | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Electrode Cable(Cond-Elec-cbl) | Contact person Mr. Dwivedi 8299446431 Mr. Pawankumar 970957262 | Currently the instrument engineer i not at site. Once he will join there, we will support for commissioning in remote support mode | Currently the instrument engineer i not at site. Once he will join there, we will support for commissioning in remote support mode | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFKK25002022802 | 2025-04-07 | 2025-04-14 | Cust 564 | CALCUTTA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Conductivity Tranmitter Polymetron | Conductivity Tranmitter Polymetron 9125(9125) | 3 no's turbidity analyser after given the 24V power supply ,the instrument is not turning on.And chlorine analyser in service condition continuous basis 0.8 ppm value is showing ,without dose the chlorine . | Support given remotely, visit rescheduled | remote support given, visit rescheduled | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK25002292803 | 2025-04-07 | 2025-04-09 | Cust 1094 | CALCUTTA | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | Peristaltic sampling pump motor for Multiparameter(AWA-PSPM) | NEW SAMPLE PUMP CHANGE AND CALIBRATION OF THE ANALYZER. | Spares changed and issue resolved. report attached. | Spares changed and issue resolved. report attached. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFKK25002652804 | 2025-04-08 | 2025-04-12 | Cust 1107 | CALCUTTA | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | FM make current output card(1) | Commissioning activity | Commissioned. | Site visited, MOM done. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK25002962805 | 2025-04-08 | 2025-04-15 | Cust 1138 | CALCUTTA | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | KMDA Jangipur+Raghunathganj+Berhanpur LS& STP projects-- 03 nos of OCEMs | Job done. Report made. | Job done. Report made. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK25003462806 | 2025-04-09 | 2025-04-15 | Cust 636 | CALCUTTA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | CX2000--2 Nos in Infosys Kolkata Sr no-- 231024/231023 | Visited. Job done. MOM made. | Visited. Job done. MOM made. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK25003532807 | 2025-04-09 | 2025-04-12 | Cust 1113 | CALCUTTA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | Flow cell for Multiparameter(AWA-FLOWCELL) | modbus faliure | The customer agreed to send the failty FMBB &amp; modbus isolator to the Pune for repairing under warranty with FOC basis. | The customer agreed to send the failty FMBB &amp;amp;amp; modbus isolator to the Pune for repairing under warranty with FOC basis. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFKK25003982808 | 2025-04-10 | 2025-04-10 | Cust 1058 | CALCUTTA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC VISIT FOR CR6 ANALYZER | AMVC VISIT MADE | AMC VISIT HAS BEEN MADE | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFKK25003982808 | 2025-04-10 | 2025-04-10 | Cust 1058 | CALCUTTA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC VISIT FOR CR6 ANALYZER | AMC VISIT MADE | AMC VISIT HAS BEEN MADE | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFKK25004052809 | 2025-04-10 | 2025-04-12 | Cust 1066 | CALCUTTA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Chlorine Analyzer-MAC 100 | MAC 100(1) | READING WAS NOT SHOWING | visited and need spare cards | visit has been done | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK25014602810 | 2025-05-02 | 2025-05-08 | Cust 1108 | CALCUTTA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Electrode | Polilyte Plus VP 360(HT242 442) | RTD faliure within warranty period | Material sent to Client. | Material sent to Client. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFKK25014672811 | 2025-05-02 | 2025-05-02 | Cust 1107 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation and CPCB connectivity | We have re-installed the software with an upgraded version and completed CPCB connectivity. Now the data is live on the same server. We accessed the DATCon and resolved the time sync issue for U#5. | We have re-installed the software with an upgraded version and completed CPCB connectivity. Now the data is live on the same server. We accessed the DATCon and resolved the time sync issue for U#5. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFKK25016662812 | 2025-05-06 | 2025-05-12 | Cust 797 | CALCUTTA | Analytical - CHAKAN | CLOSED | AMC | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | ARC VISIT | AMC VISITED AND FURTHER VISIT NEED FOR DHDT | AMC VISIT HAS BEEN MADE | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFKK25016792813 | 2025-05-07 | 2025-05-10 | Cust 1014 | CALCUTTA | DCS - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | MICROCON+ER SERIES | CONTROLLER MODULE (ER-201)(ER-201) | Data is not uploading to CPCB server. | We have accessed the Black Box and observed that the data from Stack 2 _ Unit-II 64 MW CPP Gas was not coming till the Black Box. the issue has been resolved from the hardware side as the data is now showing live on the CPCB server. | the issue has been resolved from the hardware side as the data is now showing live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFKK25017222814 | 2025-05-08 | 2025-05-09 | Cust 1107 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB connectivity needs to be done. | We have completed the CPCB connectivity and now the Data is live on the CPCB server. | We have completed the CPCB connectivity and now the Data is live on the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFKK25017452815 | 2025-05-08 | 2025-05-14 | Cust 1081 | CALCUTTA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Display touch is slow | System is working healthy in line from last 10yrs above | job completed | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFKK25017552816 | 2025-05-08 | 2025-05-14 | Cust 1058 | CALCUTTA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Calibration Error | Remote support given, problem solved | job completed | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK25019982817 | 2025-05-15 | 2025-05-16 | Cust 823 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue. | we are not getting proper data till the PC, Multiparameter Analyzer has recently been replaced, So it needs to do some settings in MPL so for now we are closing this complaint. | we are not getting proper data till the PC, Multiparameter Analyzer has recently been replaced, So it needs to do some settings in MPL so for now we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFKK25020052818 | 2025-05-15 | 2025-05-16 | Cust 1132 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data upload issue. | we accessed both black boxes and now internet side issue. been resolved and data is live on CPCB server for both devices. | we accessed both black boxes and now internet side issue. been resolved and data is live on CPCB server for both devices. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFKK25020192819 | 2025-05-15 | 2025-05-18 | Cust 1083 | CALCUTTA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | AQUADO Analyzer | DO2 Arc Sensors(DO2 Arc Sensors) | DO at STP-Halisahar | Visited. Job done. Report made. | Visited. Job done. Report made. | NO | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK25020502820 | 2025-05-16 | 2025-05-19 | Cust 1129 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | camera streming issue. | We took access and found that the EnviroconnectDAS server was not in running so after restarting the same, the Data started updating, and for the Camera, the Live video streaming window (run.sh) was not in running so after so after restarting the same the camera streaming started live. | Now the issue has been resolved, and the Camera streaming is live on the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFKK25020532821 | 2025-05-16 | 2025-05-17 | Cust 797 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | ETP-1, ETP-2 & ETP-3_RO Outlet COD, BOD, TSS & pH analysers are showing Offline in Forbes Marshall online portal | Not found any issue as data starts updating automatically. | Not found any issue as data starts updating automatically. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFKK25021192822 | 2025-05-18 | 2025-05-21 | Cust 604 | CALCUTTA | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Electrode | Polilyte Plus VP 120(HT242 428) | pH at WTP o/L | Visited. Commissioned. MOM made. | Visited. Commissioned. MOM made. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK25021202823 | 2025-05-18 | 2025-05-23 | Cust 915 | CALCUTTA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | Mother board-PCX2 of Multiparameter Analyzer(1) | Calibration | Site visit made and MOM is attached | Site visit made and MOM is attached | NO | Paid | Core | N |  | Analytical - CHAKAN
AFKK25021292824 | 2025-05-19 | 2025-05-23 | Cust 910 | CALCUTTA | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | E & C activity | Site visit made and MOM is attached | Site visit made and MOM is attached | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK25022952825 | 2025-05-21 | 2025-05-26 | Cust 797 | CALCUTTA | Analytical - CHAKAN | CLOSED | AMC | TOC ANALYSER | FM 101 TOC Analyser(FM 101 TOC) | AMC JOB | AMC VISISTED | AMC VISIT HAS BEEN MADE | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFKK25022952825 | 2025-05-21 | 2025-05-26 | Cust 797 | CALCUTTA | Analytical - CHAKAN | CLOSED | AMC | TOC ANALYSER | FM 101 TOC Analyser(FM 101 TOC) | AMC JOB | AMC VISISTED | AMC VISIT HAS BEEN MADE | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFKK25022952825 | 2025-05-21 | 2025-05-26 | Cust 797 | CALCUTTA | Analytical - CHAKAN | CLOSED | AMC | TOC ANALYSER | FM 101 TOC Analyser(FM 101 TOC) | AMC JOB | AMC VISISTED | AMC VISIT HAS BEEN MADE | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFKK25023182826 | 2025-05-22 | 2025-05-25 | Cust 1132 | CALCUTTA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | pH input card-FM of Multiparameter(1) | TSS READING- FLUCTUATION | Site visited, MOM done. | Site visited, MOM done. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFKK25023662827 | 2025-05-23 | 2025-05-27 | Cust 847 | CALCUTTA | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Smartpro Electronics Assembly(Smartpro-ele) | pH, Cond, Flow at Boxer Pepsi | Visited. Commissioned. MOM made. | Visited. Commissioned. MOM made. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK25023982828 | 2025-05-23 | 2025-05-28 | Cust 1120 | CALCUTTA | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Proactive visit to discuss regarding the TOC analyser | Visit made, discussion done. | Visit made BY iNDRANIL Sarkhel, discussion done. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFKK25025812829 | 2025-05-27 | 2025-06-02 | Cust 1086 | CALCUTTA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | Xenon Lamp Power supply for Multiparameter(AWA-XLAMPPS) | Issue in MPA | SITE VISITED,MOM DONE. | Site visited, MOM done. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFKK25026592830 | 2025-05-28 | 2025-05-30 | Cust 823 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data mismatch issue. | we took access of the system and made changes in configuration (Analyzer ranges), now the issue has been resolved and data is live on CPCB server and matching with the actual readings. | we took access of the system and made changes in configuration (Analyzer ranges), now the issue has been resolved and data is live on CPCB server and matching with the actual readings. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFKK25026732831 | 2025-05-28 | 2025-06-02 | Cust 823 | CALCUTTA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | CPU Display card for Multiparameter(AWA-CPU1) | New CPU card installation in old ETP analyser | Site visited, MOM done. | Site visited, MOM done. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFKK25027252832 | 2025-05-29 | 2025-05-30 | Cust 1139 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Blackbox configuration for PM data. | We took access to the system and now correct stack data is updating on CPCB server. | We took access to the system and now correct stack data is updating on CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFKK25027652833 | 2025-05-30 | 2025-06-02 | Cust 1113 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB Connectivity needs to be done. | We completed the CPCB connectivity, and the Data is live on the CPCB server. | We completed the CPCB connectivity, and the Data is live on the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFKK25027672834 | 2025-05-30 | 2025-06-05 | Cust 1132 | CALCUTTA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | 12 V POWER SUPPLY CARD for Multiparameter(AWA-PS1) | MPA ANALYZER' TSS READING GET HIGH | The analyser is under observation. Job report sent. | The analyser is under observation. Job report sent. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFKK25028032835 | 2025-05-30 | 2025-05-31 | Cust 692 | CALCUTTA | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | Chlorine sensor | Chlorine sensor-Krohne(1) | SENSOR GIVING ERRATIC READING . NEED TO CHECK FURTHER AND GIVE FINAL CONCLUSION. | Checked the sensor and found faulty. Offer given of new Sensor. | Checked the sensor and found faulty. Offer given of new Sensor. | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFKK25030362836 | 2025-06-03 | 2025-06-09 | Cust 1078 | CALCUTTA | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | MULTIPARAMETER ANALYSER | Flowcell disc of Multiparameter(1) | Issues in Multiparameter analyser | Site visit made and MOM is attached | Site visit made and MOM is attached | NO | Paid | Core | N |  | Analytical - CHAKAN
AFKK25031722837 | 2025-06-05 | 2025-06-08 | Cust 797 | CALCUTTA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc visit | amc visit has been made | AMC VISIT HAS BEEN MADE | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFKK25032362838 | 2025-06-08 | 2025-06-08 | Cust 797 | CALCUTTA | Analytical - CHAKAN | CLOSED | AMC | TOC ANALYSER | Ultra Violet Lamp(001?015) | AMC VISIT | AMC VISIT | AMC VISIT HAS BEEN MADE | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFKK25034872839 | 2025-06-13 | 2025-06-16 | Cust 1099 | CALCUTTA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | DATCON | DATCON(DATCON) | CPCB data upload issue. | Commercially not clear | Commercially not clear | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFLN23035672840 | 2023-06-06 | 2023-06-07 | Cust 1140 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Need to update UPPCB domain URL. | UPPCB Url is updated and data is uploading on the server | UPPCB Url is updated and data is uploading on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFLN23037442841 | 2023-06-11 | 2023-06-11 | Cust 498 | LUCKNOW | Analytical - CHAKAN | CLOSED | COMMISSIONING | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | commissioning of cx 2000 | TRAINING PROVIDED TO THE CUSTOMER AND SYSTEM IS WORKING FINE | COMMISSIONING OF THE CX 2000 DONE AND TRAIUNING PROVIDED TO THE CUSTOMER | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFLN23041052842 | 2023-06-16 | 2023-06-16 | Cust 1141 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB + UPPCB data upload issue. | Data is live on CPCB + UPPCB server, hence we are closing this complaint. | Data is live on CPCB + UPPCB server, hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFLN23045512843 | 2023-06-25 | 2023-07-18 | Cust 1142 | LUCKNOW | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Transmitter Aquacon | Conductivity Transmitter Aquamon 2000(CON-TR-2000) | VALUES OF NOT MATCHING WITH THE LAB REPORTS | there are minor mismatch with the lab reports | readings are ok there are minor mismatch with the lab reports | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFLN23045512843 | 2023-06-25 | 2023-07-18 | Cust 1142 | LUCKNOW | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Transmitter Aquacon | Conductivity Transmitter Aquamon 2000(CON-TR-2000) | VALUES OF NOT MATCHING WITH THE LAB REPORTS | there are minor mismatch with the lab reports | readings are ok there are minor mismatch with the lab reports | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFLN23054182844 | 2023-07-11 | 2023-07-12 | Cust 600 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading on CPCB server. | Data is live on CPCB + UPPCB server, hence we are closing this complaint. | Data is live on CPCB + UPPCB server, hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFLN23067832845 | 2023-08-03 | 2023-08-04 | Cust 1092 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB + UPPCB configuration needs to be done. | We took the access to the black box and now data is coming till the local system, hence we are closing this complaint. | We took the access to the black box and now data is coming till the local system, hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFLN23074012846 | 2023-08-17 | 2023-08-18 | Cust 1092 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB + UPPCB data upload issue. | The issue is from hardware side, after repair analyzer data is live on CPCB server, and we also change the UPPCB server URL and now data is live on both the server. | The issue is from hardware side, after repair analyzer data is live on CPCB server, and we also change the UPPCB server URL and now data is live on both the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFLN23087972847 | 2023-09-13 | 2023-09-20 | Cust 1143 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | System crashed, software reinstallation needs to be done. | We check the data on MODSCAN and same data is reflected on CPCB server. as of now we are closing this complaint. | We check the data on MODSCAN and same data is reflected on CPCB server. as of now we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFLN23088002848 | 2023-09-13 | 2023-09-15 | Cust 862 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Enviroconnect and hardware data mismatch | Took access of the system and checked the data as per calculation correct data displaying on software. | Took access of the system and checked the data as per calculation correct data displaying on software. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFLN23091112849 | 2023-09-20 | 2023-09-22 | Cust 827 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation needs to be done. | We took the access to the system and successfully did the software reinstallation local + CPCB + UPPCB configuration and now data is liv eon all the servers. | We took the access to the system and successfully did the software reinstallation local + CPCB + UPPCB configuration and now data is liv eon all the servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFLN23092232850 | 2023-09-22 | 2023-11-28 | Cust 1143 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data mismatch issue. | Now data is matching and also live on UPPCB server. | Now data is matching and also live on UPPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFLN23098212851 | 2023-10-01 | 2023-10-01 | Cust 1092 | LUCKNOW | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | Cleaning Pump(1) | paramaters are not matching | no problem observed | parametrs are ok now | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFLN23099702852 | 2023-10-04 | 2023-10-06 | Cust 827 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | UPPCB data upload issue. | Now data is liv eon UPPCB server, hence we are closing this complaint. | Now data is liv eon UPPCB server, hence we are closing this complaint. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFLN23105512853 | 2023-10-17 | 2023-10-18 | Cust 1141 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation | reinstalled the software and completed the Local + CPCB + UPPCB connectivity. Now data is continuously reflecting on the same. | reinstalled the software and completed the Local + CPCB + UPPCB connectivity. Now data is continuously reflecting on the same. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFLN23107282854 | 2023-10-20 | 2023-10-20 | Cust 835 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB | We took access to the system and data from the ETP outlet is now live on CPCB server. | We took access to the system and data from the ETP outlet is now live on CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFLN23108172855 | 2023-10-23 | 2023-10-25 | Cust 1141 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | issues in online Camera connectivity. | We took access to the system and Camera streaming is now live on CPCB server. | We took access to the system and Camera streaming is now live on CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFLN23112222856 | 2023-10-30 | 2023-11-01 | Cust 1144 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Software reinstallation needs to be done. | Successfully did the CPCB configuration and now data is live on CPCB server. | Successfully did the CPCB configuration and now data is live on CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFLN23114462857 | 2023-11-01 | 2023-11-03 | Cust 1141 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not streaming on CPCB Server. | Camera not playing in VLC as of now we are closing this complaint. | Camera not playing in VLC as of now we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFLN23117592858 | 2023-11-08 | 2023-11-08 | Cust 835 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB connectivity issue | Data from ETP_Outlet_Flow is now live on the CPCB server. | Data from ETP_Outlet_Flow is now live on the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFLN23126932859 | 2023-11-28 | 2023-12-01 | Cust 936 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | stack_1_(50TPH) slope fired boiler data not showing on CPCB/UPCB server | As of now we are closing this complaint, we will reopen this once the issue is resolved from hardware side. (analyzer data is not coming till the Black Box) | As of now we are closing this complaint, we will reopen this once the issue is resolved from hardware side. (analyzer data is not coming till the Black Box) | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFLN23134772860 | 2023-12-12 | 2024-04-29 | Cust 498 | LUCKNOW | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | Xenon Lamp Power supply for Multiparameter(AWA-XLAMPPS) | CX2000 Analyser Xenon lamp not flashing | Visit Done By Krishnanand. Upadhyay last week. New Power supply Installed & Units is working FIne after that | Visit Done By Krishnanand Upadhyay last week. New Power supply Installed &amp;amp;amp; Units is working FIne after that | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFLN23150512861 | 2024-01-08 | 2024-01-11 | Cust 1092 | LUCKNOW | DCS - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | MICROCON+ER SERIES | CONTROLLER MODULE (ER-201)(ER-201) | CPCB + UPPCB data upload issue. | Data is not coming till the BB , hence as of now we are closing this complaint and will reopen it after hardware side issue is resoled. | Data is not coming till the BB , hence as of now we are closing this complaint and will reopen it after hardware side issue is resoled. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFLN23151452862 | 2024-01-10 | 2024-04-09 | Cust 861 | LUCKNOW | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | PH Sensor for Multiparameter(AWA-PH-SENS) | card is replaced | card dispatched | card is replaced | YES | Paid | Process | N |  | Analytical - CHAKAN
AFLN23151452862 | 2024-01-10 | 2024-04-09 | Cust 861 | LUCKNOW | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | RS485 CARD for Multiparameter(AWA-RS485) | RS 485 card is found failed and display brightness issue so card and display needs to be replaced. | card replaced | card is replaced | YES | Paid | Process | N |  | Analytical - CHAKAN
AFLN23156322863 | 2024-01-19 | 2024-01-22 | Cust 1145 | LUCKNOW | DCS - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB and GPCB server. | Customer side issue data stared uploading after restarting the PC. | Customer side issue data stared uploading after restarting the PC. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFLN23161122864 | 2024-01-30 | 2024-02-02 | Cust 1141 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on CPCB Server. | We took access to the system and now camera is live on CPCB Server. | We took access to the system and now camera is live on CPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFLN23167662865 | 2024-02-08 | 2024-04-29 | Cust 874 | LUCKNOW | Analytical - CHAKAN | CLOSED | ENGINEERING ISSUE | THERMO SERIES | AV38 TRANSMITTER(1) | Thermo silica analyzers are not working, kindly arrange a visit to the Thermo Engineer to conclude the required spares to make both silica analyzers functional. | Thermo suggested New Model to be procured | Thermo suggested To go for new Model | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFLN23180312866 | 2024-02-29 | 2024-02-29 | Cust 1092 | LUCKNOW | Analytical - CHAKAN | CLOSED | APPLICATION ISSUE | Conductivity Electrode | Forbes Marshall Cond. Cell Model 8310, K=0.01(FM -8310) | calibration | calibration done and its under observation | calibration of conductivity meter done | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFLN23180352867 | 2024-02-29 | 2024-02-29 | Cust 476 | LUCKNOW | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(Ph-Trans-Smartpro) | calibration of ph transmitter | calibration done | calibration done | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFLN23180402868 | 2024-02-29 | 2024-02-29 | Cust 1144 | LUCKNOW | Analytical - CHAKAN | CLOSED | APPLICATION ISSUE | MULTIPARAMETER ANALYSER | mA Output card for Multiparameter(AWA-4-20mA) | r to i convertor not working | convertor changed and its working fine | r to i convertor changed | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFLN23181662869 | 2024-03-04 | 2024-03-07 | Cust 1141 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | DATA and CAMERA not updating on CPCB server. | Camera was not live on CPCB Server. We had taken up this issue with development team now camera is live on CPCB Server. | Camera was not live on CPCB Server. We had taken up this issue with development team now camera is live on CPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFLN23192732870 | 2024-03-25 | 2024-03-25 | Cust 600 | LUCKNOW | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | commissioning | commissioning done | commissioning of multiparameter analyser | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFLN23192742871 | 2024-03-25 | 2024-03-25 | Cust 600 | LUCKNOW | Analytical - CHAKAN | CLOSED | DIGITAL CONEECTIVITY | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | digital connectivity of analyzer | digital connectivity done | digital connectivity done | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFLN23196882872 | 2024-03-30 | 2024-03-30 | Cust 821 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB+UPPCB Portal. | We took access of the system and data is live on LOCAL + CPCB+ UPPCB. | We took access of the system and data is live on LOCAL + CPCB+ UPPCB. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFLN24002662873 | 2024-04-05 | 2024-04-05 | Cust 1141 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation | We have reinstalled the software with the updated version and data is now live on the Local + CPCB + UPPCB servers. | We have reinstalled the software with the updated version and data is now live on the Local + CPCB + UPPCB servers. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFLN24012432874 | 2024-04-26 | 2024-04-29 | Cust 1140 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | CPCB data upload issue. | we took access to the system and changed the configuration (Scanner to RS485). Currently, Multiparameter data is live on the CPCB + UPPCB server. | we took access to the system and changed the configuration (Scanner to RS485). Currently, Multiparameter data is live on the CPCB + UPPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFLN24014372875 | 2024-04-30 | 2024-05-03 | Cust 1146 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | CPCB data upload issue. | As of now we are closing this complaint and will reopen it once the hardware side issue is resolved. | As of now we are closing this complaint and will reopen it once the hardware side issue is resolved. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFLN24015342876 | 2024-04-30 | 2024-04-30 | Cust 1092 | LUCKNOW | Analytical - CHAKAN | CLOSED | APPLICATION ISSUE | MULTIPARAMETER ANALYSER | LCD of Multiparameter Display(1) | ocms not working | ocms has to send pune | ocms has to send to pune | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFLN24016142877 | 2024-05-02 | 2024-05-08 | Cust 1147 | LUCKNOW | Steam System Division | CLOSED | PRODUCT FAILURE | CRS | CONDENSATE RECOVERY METER - MODEL CRM-485(CRM-485R) | display is blank | Please send the CRM to Chakan | Please send the CRM to Chakan | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFLN24018142878 | 2024-05-07 | 2024-05-14 | Cust 862 | LUCKNOW | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Customer reported for troubleshooting of multiparameter analyzer COD & TSS value are in error-1 & PH1-PH-2 values are in offset condition. | display send for repair | display send for repair | NO | Paid | Core | N |  | Analytical - CHAKAN
AFLN24024752879 | 2024-05-22 | 2024-07-24 | Cust 937 | LUCKNOW | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data upload issue. | today we took access of the black box and completed CPCB connectivity, and now data is live on CPCB server | today we took access of the black box and completed CPCB connectivity, and now data is live on CPCB server | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFLN24028062880 | 2024-05-28 | 2024-06-04 | Cust 915 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB + UPPCB data upload issue. | Now the internet side issue is resolve from customers end and data is live on CPCB + UPPCB server. | Now the internet side issue is resolve from customers end and data is live on CPCB + UPPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFLN24030232881 | 2024-05-31 | 2024-06-01 | Cust 822 | LUCKNOW | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Customer reported for calibration of 2 multiparameter analyzer | Calibration done | Calibration done | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFLN24030232881 | 2024-05-31 | 2024-06-01 | Cust 822 | LUCKNOW | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Customer reported for calibration of 2 multiparameter analyzer | Calibration done | Calibration done | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFLN24031522882 | 2024-06-01 | 2024-06-03 | Cust 1148 | LUCKNOW | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | Customer reported for calibration of conductivity transmitter's. | Transmitter send for checking. | Calibration done. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFLN24031522882 | 2024-06-01 | 2024-06-03 | Cust 1148 | LUCKNOW | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | Customer reported for calibration of conductivity transmitter's. | Calibration done | Calibration done. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFLN24031522882 | 2024-06-01 | 2024-06-03 | Cust 1148 | LUCKNOW | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | Customer reported for calibration of conductivity transmitter's. | Calibration Done | Calibration done. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFLN24031522882 | 2024-06-01 | 2024-06-03 | Cust 1148 | LUCKNOW | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | Customer reported for calibration of conductivity transmitter's. | Calibration Done | Calibration done. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFLN24031802883 | 2024-06-03 | 2024-06-10 | Cust 1149 | LUCKNOW | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | DO Transmitters | optex series(optex series) | Customer reported for calibration of DO sensor installed at there site. | checked the sensor calibration found ok | checked the sensor calibration found ok | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFLN24035602884 | 2024-06-11 | 2024-06-14 | Cust 915 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB + UPPCB data upload issue for STP. | Need to arrange site visit, we will reopen it once engineer arrange site visit. | Need to arrange site visit, we will reopen it once engineer arrange site visit. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFLN24040102885 | 2024-06-20 | 2024-06-21 | Cust 600 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB connectivity needs to be done. | We have completed the Black Box configuration and now the ETP data is live on the CPCB server. | We have completed the Black Box configuration and now the ETP data is live on the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFLN24040572886 | 2024-06-21 | 2024-06-21 | Cust 1141 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data and Camera configuration for CPCB server. | We took access to the system and now data and Camera is live on CPCB Server. | We took access to the system and now data and Camera is live on CPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFLN24041722887 | 2024-06-24 | 2024-06-26 | Cust 1141 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on CPCB Server. | We took access to the system and found camera streaming window was closed and now camera is live on CPCB Server. | We took access to the system and found camera streaming window was closed and now camera is live on CPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFLN24043362888 | 2024-06-27 | 2024-06-29 | Cust 1141 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on CPCB Server. | We took access to the system and now Camera is live on CPCB server and found that camera setting were changed due to which camera was lnactive on the server. | We took access to the system and now Camera is live on CPCB server and found that camera setting were changed due to which camera was lnactive on the server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFLN24044252889 | 2024-06-28 | 2024-06-29 | Cust 1092 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data mismatch issue. | we took the access to your system and checked the data on MODSCAN, and as per calculations the same data uploading onto the CPCB server for COD & TSS parameters. | we took the access to your system and checked the data on MODSCAN, and as per calculations the same data uploading onto the CPCB server for COD &amp;amp; TSS parameters. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFLN24052342890 | 2024-07-11 | 2024-07-12 | Cust 937 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB connectivity needs to be done (BB) | today we took access of the black box and completed CPCB connectivity, and now data is live on CPCB server | today we took access of the black box and completed CPCB connectivity, and now data is live on CPCB server | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFLN24062302891 | 2024-07-29 | 2024-07-31 | Cust 1146 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | CPCB data uploading issue | corrected all tags and slave IDs according to given details and data is now live on the CPCB server. | I have corrected all tags and slave IDs according to given details and data is now live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFLN24066892892 | 2024-08-05 | 2024-08-12 | Cust 1150 | LUCKNOW | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Smartpro-con) | We are facing a problem that Online Conductivity meter (Boiler) is not working properly from last some days . While we had replaced the Sensor & Cable but problem is same. So please rectify problem as soon as possible . | Analyzer send to pune for repair | Analyzer send for repair to pune. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFLN24068912893 | 2024-08-09 | 2024-08-12 | Cust 936 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CAMERA STREAMING ISSUE. | Now the Camera streaming is live continuously on the CPCB server. | Now the Camera streaming is live continuously on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFLN24072942894 | 2024-08-19 | 2024-08-21 | Cust 1147 | LUCKNOW | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(Ph-Trans-Smartpro) | ph sensor calibration work | ph sensor and control valve calibration work | ph meter & control valve calibration work | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFLN24075052895 | 2024-08-22 | 2024-08-23 | Cust 1141 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on CPCB Server. | We took access to the system and now camera is live on CPCB Server. | We took access to the system and now camera is live on CPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFLN24077852896 | 2024-08-27 | 2024-08-30 | Cust 946 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera configuration needs to be done. | As of now we are closing this complaint and will reopen it once the camera issue is resolved. | As of now we are closing this complaint and will reopen it once the camera issue is resolved. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFLN24077952897 | 2024-08-27 | 2024-08-30 | Cust 1141 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera issue | As of now we are closing this complaint. We will reopen it once the camera issue is resolved from customers end. | As of now we are closing this complaint. We will reopen it once the camera issue is resolved from your end. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFLN24079072898 | 2024-08-29 | 2024-08-30 | Cust 1151 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Incorrect Date and Time displaying on LOCAL Server. | We took access to the system and sync hardware clock after which correct date updated on LOCAL Server. | We took access to the system and sync hardware clock after which correct date updated on LOCAL Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFLN24083582899 | 2024-09-04 | 2024-11-11 | Cust 1141 | LUCKNOW | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | While commissioning CX2000 analyser noticed continuous Air bubbles in flow cell. After dismantling flow cell one disc found missing. Even after replacement the leakage still there seems some damage internally to flow cell | Changes have been done in AIS. We started 3 days logging of analyser data on Different value standards & distilled water. So any such issues will be caught in pre dispatch testing only. | Changes have been done in AIS. We started 3 days logging of analyser data on Different value standards &amp;amp;amp; distilled water. So any such issues will be caught in pre dispatch testing only. | YES | No Trip | Process | Y | Product / equipment failure impacting safety of equipment, process or people. | Analytical - CHAKAN
AFLN24083772900 | 2024-09-04 | 2024-09-06 | Cust 1141 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Blackbox configuration and CPCB + Cloud server connectivity needs to be done. | The new Analyser CX2000 has been commissioned at this site and for now we have completed the CPCB and UPPCB connectivity and currently, the Data is live for COD, BOD, TSS and pH parameters. | The new Analyser CX2000 has been commissioned at this site and for now we have completed the CPCB and UPPCB connectivity and currently, the Data is live for COD, BOD, TSS and pH parameters. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFLN24086882901 | 2024-09-11 | 2024-09-13 | Cust 915 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data upload issue. | Now data is live on CPCB + UPPCB server. | Now data is live on CPCB + UPPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFLN24090232902 | 2024-09-18 | 2024-09-18 | Cust 936 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera streaming is not live on CPCB server. | We took access to the system and now camera streaming is live on CPCB server. | We took access to the system and now camera streaming is live on CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFLN24113112903 | 2024-11-05 | 2025-02-10 | Cust 874 | LUCKNOW | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | CX2000 Flowcell body crack noticed during commissioning. Causing water ingress in Flow cell | Replacment material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210128111 LR Dt - 10-JAN-2025 | Replacment material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210128111 LR Dt - 10-JAN-2025 | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFLN24114532904 | 2024-11-08 | 2024-11-08 | Cust 497 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | FMBB configuration needs to be done. | As discussed, today we took access of the black box and completed configuration, now data is coming till the Black Box. | As discussed, today we took access of the black box and completed configuration, now data is coming till the Black Box. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFLN24116092905 | 2024-11-12 | 2024-11-15 | Cust 1152 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to CPCB server. | Data is live on the CPCB server. | Data is live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFLN24116202906 | 2024-11-12 | 2024-11-15 | Cust 936 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera streaming issue on CPCB server. | Camera streaming is live on the CPCB server. | Camera streaming is live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFLN24116522907 | 2024-11-12 | 2024-11-15 | Cust 1153 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB + GPCB data upload issue. | issue is resolved from the customer end and data is live on CPCB + GPCB server also completed Camera configuration for SPENT WASH OUTLET , now camera is streaming live on CPCB server. | issue is resolved from the customer end and data is live on CPCB + GPCB server also completed Camera configuration for SPENT WASH OUTLET , now camera is streaming live on CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFLN24117902908 | 2024-11-14 | 2025-02-10 | Cust 1092 | LUCKNOW | Analytical - CHAKAN | CLOSED | SALES PROMOTION | MULTIPARAMETER ANALYSER | Cleaning Pump(1) | Cx 1000 is not working due to failure of motherboard and sample pump needs to be replaced. | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210128128 LR Dt - 10-JAN-2025 | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210128128 LR Dt - 10-JAN-2025 | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFLN24118172909 | 2024-11-15 | 2024-11-15 | Cust 1154 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Localhost license issue. | PC name was changed due to which localhost was not accessible. We changed the PC name and now localhost is accessible and data is live on Server. | PC name was changed due to which localhost was not accessible. We changed the PC name and now localhost is accessible and data is live on Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFLN24119822910 | 2024-11-19 | 2024-11-22 | Cust 1154 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB Server. | As of now we are closing this complaint as CPCB URL is Blocked in customers network. | As of now we are closing this complaint as CPCB URL is Blocked in customers network. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFLN24119942911 | 2024-11-19 | 2024-11-26 | Cust 1152 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera streaming issue | Camera streaming continuously live on the CPCB server. | Camera streaming continuously live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFLN24120182912 | 2024-11-19 | 2024-11-24 | Cust 798 | LUCKNOW | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | We are facing a issues with pH Sensor at PAPL ,Varanasi Client - Parle Agro Pvt. Ltd. Address- Varanasi Contact no- Mr Vishnu- 9369835171-/ Rajneesh- 9792200599 Problem- pH sensor not working. | now ph sensor working fine | now ph sensor working fine | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFLN24125032913 | 2024-11-28 | 2024-11-29 | Cust 1155 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB connectivity needs to be done. | we took access of the black box and completed CPCB connectivity for ETP1 _SPENT WASH, and also the internet side issue is resolved for STACK_20TPH_Boiler. and now data is live on CPCB server for both parameters. | we took access of the black box and completed CPCB connectivity for ETP1 _SPENT WASH, and also the internet side issue is resolved now data is live on CPCB server for both parameters. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFLN24128662914 | 2024-12-03 | 2024-12-06 | Cust 1146 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Eversense Login page not opening. | We have re-configured the system and data from all devices are now live on the Local and CPCB server and the same is verified by customer. | We have re-configured the system and data from all devices are now live on the Local and CPCB server and the same is verified by customer. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFLN24135332915 | 2024-12-17 | 2024-12-18 | Cust 936 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera streaming issue on CPCB server. | We have accessed the system, and the camera is now live on the CPCB server. | We have accessed the system, and the camera is now live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFLN24139092916 | 2024-12-24 | 2024-12-31 | Cust 866 | LUCKNOW | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Customer reported for calibration for MPA | Calibration done. | Calibration done. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFLN24140022917 | 2024-12-26 | 2024-12-28 | Cust 937 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | cpcb data mismatch issue. | we took access of the black box and made some changes in FMBB configuration and also updated the new patch of FMBB, now data is live on CPCB server and the same is verified by you also. | we took access of the black box and made some changes in FMBB configuration and also updated the new patch of FMBB, now data is live on CPCB server and the same is verified by you also. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFLN24142042918 | 2024-12-30 | 2024-12-31 | Cust 937 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB Camera streaming issue. | we took access to your system, and now the issue has been resolved and Camera is streaming live on CPCB server. | we took access to your system, and now the issue has been resolved and Camera is streaming live on CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFLN24145262919 | 2025-01-02 | 2025-01-03 | Cust 1141 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to CPCB | Now the camera is live on the CPCB server. Multiparameter data is reflecting and currently, flowmeter is not running hence readings will not be available on the CPCB server. | Now the camera is live on the CPCB server. Multiparameter data is reflecting and currently, flowmeter is not running hence readings will not be available on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFLN24145292920 | 2025-01-02 | 2025-01-03 | Cust 871 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB+LOCAL server. | We have observed that data is live on CPCB so we are closing this complaint from our end. | We have observed that data is live on CPCB so we are closing this complaint from our end. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFLN24148572921 | 2025-01-09 | 2025-01-12 | Cust 1156 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | software re -installation | Software re-installation and cpcb configuration completed. Currently, data is not coming till modscan. As of now we are closing this complaint. We will reopen this complaint once the issue is resolved from the hardware side. | As of now we are closing this complaint. We will reopen this complaint once the issue is resolved from the hardware side. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFLN24153942922 | 2025-01-21 | 2025-01-24 | Cust 1156 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | CPCB connectivity needs to be done. | We have completed local and CPCB configuration and data from the multiparameter is now live on the CPCB server. | We have completed local and CPCB configuration and data from the multiparameter is now live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFLN24153952923 | 2025-01-21 | 2025-01-22 | Cust 1157 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on CPCB server. | Now the data is live on the CPCB server. | Now the data is live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFLN24153982924 | 2025-01-21 | 2025-01-24 | Cust 819 | LUCKNOW | DCS - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | MICROCON+ER SERIES | CONTROLLER MODULE (ER-201)(ER-201) | Data upload issue on CPCB server. | We have accessed the black box, and the data is now live on the CPCB server. | We have accessed the black box, and the data is now live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFLN24158942925 | 2025-01-29 | 2025-02-28 | Cust 946 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | online Effluent monitoring data is not uploading | We have formatted the FMBlack Box and now the Data is live continously. | We have formatted the FMBlack Box and now the Data is live continously. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFLN24164142926 | 2025-02-06 | 2025-02-07 | Cust 915 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB + UPPCB configuration needs to be done. | today we have took access of the STP black box, and completed black box configuration + CPCB + UPPCB configuration, now data is live on both servers. | today we have took access of the STP black box, and completed black box configuration + CPCB + UPPCB configuration, now data is live on both servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFLN24166652927 | 2025-02-12 | 2025-02-15 | Cust 946 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB camere streaming issue. | We have updated the Camera utility form.ogg to webm and now the Camera streaming is live on the CPCB server. | We have updated the Camera utility form.ogg to webm and now the Camera streaming is live on the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFLN24167412928 | 2025-02-13 | 2025-03-14 | Cust 1152 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue. | We have updated the DAS and now the Data is live. | We have updated the DAS and now the Data is live. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFLN24167652929 | 2025-02-14 | 2025-02-18 | Cust 835 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue. | we took access to your system and made some changes in configuration, now data is live on CPCB server. | we took access to your system and made some changes in configuration, now data is live on CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFLN24167672930 | 2025-02-14 | 2025-02-18 | Cust 497 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Local configuration needs to be done. | As discussed, we took access of the black box and completed Black Box + FTP file configuration, now data is live on the Enviroconnect local server through Black Box. | As discussed, we took access of the black box and completed Black Box + FTP file configuration, now data is live on the Enviroconnect local server through Black Box. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFLN24169652931 | 2025-02-19 | 2025-02-21 | Cust 936 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera streaming issue on CPCB server. | Now the camera is streaming live on the CPCB server. | Now the camera is streaming live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFLN24177812932 | 2025-03-05 | 2025-03-07 | Cust 936 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera streaming issue on CPCB server. | Camera streaming is now live on the CPCB server. | Camera streaming is now live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFLN24181312933 | 2025-03-12 | 2025-03-14 | Cust 936 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to CPCB server. | Data is now live on the CPCB server. | Data is now live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFLN24182752934 | 2025-03-17 | 2025-03-20 | Cust 835 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | data upload issue on CPCB server | Now the data is showing live on the CPCB server. Also we have updated the Enviroconnect DAS version to the latest. | Now the data is showing live on the CPCB server. Also we have updated the Enviroconnect DAS version to the latest. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFLN25001912935 | 2025-04-07 | 2025-04-11 | Cust 1146 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Data from RO PERMEATE (1ST PHASE) is not uploading to CPCB server. | Now the issue has been resolved and data is live on the CPCB server. | Now the issue has been resolved and data is live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFLN25006822936 | 2025-04-18 | 2025-04-18 | Cust 936 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera streaming issue on CPCB server. | Now the camera is streaming live on the CPCB server. | Now the camera is streaming live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFLN25016272937 | 2025-05-06 | 2025-05-09 | Cust 811 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data uploaded issue. | Commercially not clear. | Commercially not clear. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFLN25019152938 | 2025-05-13 | 2025-05-15 | Cust 1150 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB connectivity needs to be done for mass flow meter. | we took access of the black box and now the internet side issue has been resolved also we have completed UPPCB configuration, now data is live on CPCB + UPPCB server. | we took access of the black box and now the internet side issue has been resolved also we have completed UPPCB configuration, now data is live on CPCB + UPPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFLN25026272939 | 2025-05-28 | 2025-05-30 | Cust 498 | LUCKNOW | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | Sample Pump(1) | Sample detection switch not working | Replacement material Sent on manual DC on 29-05-2025. Through DTDC AIR Courier vide LR NO- P76757849. | Replacement material Sent on manual DC on 29-05-2025. Through DTDC AIR Courier vide LR NO- P76757849. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFLN25030072940 | 2025-06-02 | 2025-06-04 | Cust 1150 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data upload issue. | now the internet issue is resolved from your side and data is live on CPCB server. | now the internet issue is resolved from your side and data is live on CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFLN25030552941 | 2025-06-03 | 2025-06-04 | Cust 1146 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Data uploading issue on CPCB server | Now the issue has been resolved and data is live on the CPCB server. | Now the issue has been resolved and data is live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFLN25034302942 | 2025-06-12 | 2025-06-13 | Cust 937 | LUCKNOW | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | data upload issue at CPCB server. | NO AMC | NO AMC | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFMM23074312943 | 2023-08-17 | 2023-08-18 | Cust 1158 | MUMBAI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | GPCB Connectivity needs to be done. | The GPCB connectivity has been completed and now the data is live on the server | The GPCB connectivity has been completed and now the data is live on the server | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFMM23101562944 | 2023-10-09 | 2023-10-12 | Cust 1159 | MUMBAI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | GPCB data uploded issue. | We took access of the Black Box for the GSPBL1 unit and found that the Flow data was not communicating with the Black Box and after that, the same hardware issue has been resolved and now the data is live on the server. | Issue has been resolved hence we are closing this complaint.. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFMM23140602945 | 2023-12-23 | 2023-12-25 | Cust 1160 | MUMBAI | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | DO Transmitters | VISI pro DO(VISI pro DO) | DO sensor commissioning require | Recommended to procure the new meter as it is obsolete | Recommended to procure the new meter as it is obsolete | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFMM23140612946 | 2023-12-23 | 2023-12-25 | Cust 1160 | MUMBAI | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | DO Transmitters | VISI pro DO(VISI pro DO) | DO sensor is not showing reading | recommended to send the indicator to Pune for the upgradation. | recommended to send the indicator to Pune for the upgradation. | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMM23154462947 | 2024-01-16 | 2024-01-16 | Cust 1161 | MUMBAI | Steam System Division | CLOSED | QUALITY PROBLEM | PRS | PRS - SIZE 100X125 - WITH ARCA-POSITIONER(PRS100X125ARCA-POSITIONER) | MEASURETROL - FMPRS 1) Cable Quality and Cable lengths to be re-defined as the current cable length is not sufficient. 2) Reliable Conduits for cable and its routing to be decided. | Jitesh Patil, Sachin Lawange, Ved Valsangkar, Umesh Kene visited the site and have arranged for the JB locally and additional cable - Inputs have been taken up with R&D, MFG and QC. | Jitesh Patil, Sachin Lawange, Ved Valsangkar, Umesh Kene visited the site and have arranged for the JB locally and additional cable - Inputs have been taken up with R&amp;amp;D, MFG and QC. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFMM23154462947 | 2024-01-16 | 2024-01-16 | Cust 1161 | MUMBAI | Steam System Division | CLOSED | QUALITY PROBLEM | PRS | PRS - SIZE 100X125 - WITH ARCA-POSITIONER(PRS100X125ARCA-POSITIONER) | MEASURETROL - FMPRS 1) Cable Quality and Cable lengths to be re-defined as the current cable length is not sufficient. 2) Reliable Conduits for cable and its routing to be decided. | Jitesh Patil, Sachin Lawange, Ved Valsangkar, Umesh Kene visited the site and have arranged for the JB locally and additional cable - Inputs have been taken up with R&D, MFG and QC. | Jitesh Patil, Sachin Lawange, Ved Valsangkar, Umesh Kene visited the site and have arranged for the JB locally and additional cable - Inputs have been taken up with R&amp;amp;D, MFG and QC. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFMM23154462947 | 2024-01-16 | 2024-01-16 | Cust 1161 | MUMBAI | Steam System Division | CLOSED | QUALITY PROBLEM | PRS | PRS - SIZE 100X125 - WITH ARCA-POSITIONER(PRS100X125ARCA-POSITIONER) | MEASURETROL - FMPRS 1) Cable Quality and Cable lengths to be re-defined as the current cable length is not sufficient. 2) Reliable Conduits for cable and its routing to be decided. | Jitesh Patil, Sachin Lawange, Ved Valsangkar, Umesh Kene visited the site and have arranged for the JB locally and additional cable - Inputs have been taken up with R&D, MFG and QC. | Jitesh Patil, Sachin Lawange, Ved Valsangkar, Umesh Kene visited the site and have arranged for the JB locally and additional cable - Inputs have been taken up with R&amp;amp;D, MFG and QC. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFMM23154462947 | 2024-01-16 | 2024-01-16 | Cust 1161 | MUMBAI | Steam System Division | CLOSED | QUALITY PROBLEM | PRS | PRS - SIZE 100X125 - WITH ARCA-POSITIONER(PRS100X125ARCA-POSITIONER) | MEASURETROL - FMPRS 1) Cable Quality and Cable lengths to be re-defined as the current cable length is not sufficient. 2) Reliable Conduits for cable and its routing to be decided. | Jitesh Patil, Sachin Lawange, Ved Valsangkar, Umesh Kene visited the site and have arranged for the JB locally and additional cable - Inputs have been taken up with R&D, MFG and QC. | Jitesh Patil, Sachin Lawange, Ved Valsangkar, Umesh Kene visited the site and have arranged for the JB locally and additional cable - Inputs have been taken up with R&amp;amp;D, MFG and QC. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFMM23154472948 | 2024-01-16 | 2024-01-16 | Cust 1161 | MUMBAI | Steam System Division | CLOSED | QUALITY PROBLEM | PRS | PRS - SIZE 100X150 - WITH ARCA-POSITIONER(PRS100X150ARCA-POSITIONER) | MEASURETROL FMPRS - Add-on accessories / complete package (panels, RTU) loose wiring / incomplete connections / no mention in UM | Jitesh Patil, Abhijeet Deshpande, Jay Kumar, Sachin Lawange, Ved Valsangkar, Umesh Kene visited the site and address the issue. The site inputs have been taken up with RND, NPD, QC and MFG | Jitesh Patil, Abhijeet Deshpande, Jay Kumar, Sachin Lawange, Ved Valsangkar, Umesh Kene visited the site and address the issue. The site inputs have been taken up with RND, NPD, QC and MFG | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFMM23154472948 | 2024-01-16 | 2024-01-16 | Cust 1161 | MUMBAI | Steam System Division | CLOSED | QUALITY PROBLEM | PRS | PRS - SIZE 100X150 - WITH ARCA-POSITIONER(PRS100X150ARCA-POSITIONER) | MEASURETROL FMPRS - Add-on accessories / complete package (panels, RTU) loose wiring / incomplete connections / no mention in UM | Jitesh Patil, Abhijeet Deshpande, Jay Kumar, Sachin Lawange, Ved Valsangkar, Umesh Kene visited the site and address the issue. The site inputs have been taken up with RND, NPD, QC and MFG | Jitesh Patil, Abhijeet Deshpande, Jay Kumar, Sachin Lawange, Ved Valsangkar, Umesh Kene visited the site and address the issue. The site inputs have been taken up with RND, NPD, QC and MFG | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFMM23154472948 | 2024-01-16 | 2024-01-16 | Cust 1161 | MUMBAI | Steam System Division | CLOSED | QUALITY PROBLEM | PRS | PRS - SIZE 100X150 - WITH ARCA-POSITIONER(PRS100X150ARCA-POSITIONER) | MEASURETROL FMPRS - Add-on accessories / complete package (panels, RTU) loose wiring / incomplete connections / no mention in UM | Jitesh Patil, Abhijeet Deshpande, Jay Kumar, Sachin Lawange, Ved Valsangkar, Umesh Kene visited the site and address the issue. The site inputs have been taken up with RND, NPD, QC and MFG | Jitesh Patil, Abhijeet Deshpande, Jay Kumar, Sachin Lawange, Ved Valsangkar, Umesh Kene visited the site and address the issue. The site inputs have been taken up with RND, NPD, QC and MFG | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFMM23154472948 | 2024-01-16 | 2024-01-16 | Cust 1161 | MUMBAI | Steam System Division | CLOSED | QUALITY PROBLEM | PRS | PRS - SIZE 100X150 - WITH ARCA-POSITIONER(PRS100X150ARCA-POSITIONER) | MEASURETROL FMPRS - Add-on accessories / complete package (panels, RTU) loose wiring / incomplete connections / no mention in UM | Jitesh Patil, Abhijeet Deshpande, Jay Kumar, Sachin Lawange, Ved Valsangkar, Umesh Kene visited the site and address the issue. The site inputs have been taken up with RND, NPD, QC and MFG | Jitesh Patil, Abhijeet Deshpande, Jay Kumar, Sachin Lawange, Ved Valsangkar, Umesh Kene visited the site and address the issue. The site inputs have been taken up with RND, NPD, QC and MFG | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFMM23154482949 | 2024-01-16 | 2024-01-16 | Cust 1161 | MUMBAI | Steam System Division | CLOSED | QUALITY PROBLEM | PRS | PRS - SIZE 100X125 - WITH ARCA-POSITIONER(PRS100X125ARCA-POSITIONER) | MEASURETROL-FMPRS - Bluetooth connectivity issue | Jitesh Patil, Jay Kumar, Yogesh Latkar, Sachin Lawange, Abhijeet Deshpande, Ved Valsangkar, Umesh Kene visited the site visit new spares and address site. Inputs taken up with RND, NPD, MFG and QC Teams. | Jitesh Patil, Jay Kumar, Yogesh Latkar, Sachin, Abhijeet Deshpande, Ved Valsangkar, Umesh Kene visited the site visit new spares and address site. Inputs taken up with RND, NPD, MFG and QC Teams. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFMM23154482949 | 2024-01-16 | 2024-01-16 | Cust 1161 | MUMBAI | Steam System Division | CLOSED | QUALITY PROBLEM | PRS | PRS - SIZE 100X125 - WITH ARCA-POSITIONER(PRS100X125ARCA-POSITIONER) | MEASURETROL-FMPRS - Bluetooth connectivity issue | Jitesh Patil, Jay Kumar, Yogesh Latkar, Sachin Lawange, Abhijeet Deshpande, Ved Valsangkar, Umesh Kene visited the site visit new spares and address site. Inputs taken up with RND, NPD, MFG and QC Teams. | Jitesh Patil, Jay Kumar, Yogesh Latkar, Sachin, Abhijeet Deshpande, Ved Valsangkar, Umesh Kene visited the site visit new spares and address site. Inputs taken up with RND, NPD, MFG and QC Teams. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFMM23154482949 | 2024-01-16 | 2024-01-16 | Cust 1161 | MUMBAI | Steam System Division | CLOSED | QUALITY PROBLEM | PRS | PRS - SIZE 100X125 - WITH ARCA-POSITIONER(PRS100X125ARCA-POSITIONER) | MEASURETROL-FMPRS - Bluetooth connectivity issue | Jitesh Patil, Jay Kumar, Yogesh Latkar, Sachin Lawange, Abhijeet Deshpande, Ved Valsangkar, Umesh Kene visited the site visit new spares and address site. Inputs taken up with RND, NPD, MFG and QC Teams. | Jitesh Patil, Jay Kumar, Yogesh Latkar, Sachin, Abhijeet Deshpande, Ved Valsangkar, Umesh Kene visited the site visit new spares and address site. Inputs taken up with RND, NPD, MFG and QC Teams. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFMM23154482949 | 2024-01-16 | 2024-01-16 | Cust 1161 | MUMBAI | Steam System Division | CLOSED | QUALITY PROBLEM | PRS | PRS - SIZE 100X125 - WITH ARCA-POSITIONER(PRS100X125ARCA-POSITIONER) | MEASURETROL-FMPRS - Bluetooth connectivity issue | Jitesh Patil, Jay Kumar, Yogesh Latkar, Sachin Lawange, Abhijeet Deshpande, Ved Valsangkar, Umesh Kene visited the site visit new spares and address site. Inputs taken up with RND, NPD, MFG and QC Teams. | Jitesh Patil, Jay Kumar, Yogesh Latkar, Sachin, Abhijeet Deshpande, Ved Valsangkar, Umesh Kene visited the site visit new spares and address site. Inputs taken up with RND, NPD, MFG and QC Teams. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFMM23154492950 | 2024-01-16 | 2024-01-18 | Cust 1161 | MUMBAI | Steam System Division | CLOSED | QUALITY PROBLEM | PRS | PRS - SIZE 100X125 - WITH ARCA-POSITIONER(PRS100X125ARCA-POSITIONER) | MEASURETROL-FMPRS - Leakage in the Piezo Block - Completed the commissioning of the Block-3 PRS. But for Block -4 PRS, could not be able to complete the commissioning, as valve is malfunctioning & not able to perform the position calibration. As per Jaykumar there is internal leakage into the valve & due to that not able to perform position calibration as well as pressure tuning | Jitesh, Jay Kumar, Abhijeet, Sachin, Ved, Umesh visited the site with new spares and site address. Inputs taken up with RND, NPD, QC and MFG | Jitesh, Jay Kumar, Abhijeet, Sachin, Ved, Umesh visited the site with new spares and site address. Inputs taken up with RND, NPD, QC and MFG | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFMM23183602951 | 2024-03-10 | 2024-03-12 | Cust 1162 | MUMBAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Electrode | Chemotrode 120(HT238 760) | Need to install new pH sensor and commission | commission the pH sesnor. | commission the pH sesnor. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFMM23183832952 | 2024-03-11 | 2024-03-13 | Cust 1160 | MUMBAI | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | DO sensor is showing low value as compare to lab value | done the setting and recommended to change interchange the 2 DO sensor to find the best result. | done the setting and recommended to change interchange the 2 DO sensor to find the best result. | EXTENDED | No Trip | Process | N |  | Analytical - CHAKAN
AFMM23194202953 | 2024-03-27 | 2024-04-03 | Cust 1161 | MUMBAI | Steam System Division | CLOSED | MANUFACTURING ISSUES | PRV - MEASURETROL | FM MEASURETROL(SIZE 080NB) | BLUETOOTH CARD NEEDS TO REPLACE. | VISITED AT SITE ANFD REPLACED THE BLUTOOTH CARD. | VISITED AT SITE ANFD REPLACED THE BLUTOOTH CARD. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFMM23194202953 | 2024-03-27 | 2024-04-03 | Cust 1161 | MUMBAI | Steam System Division | CLOSED | MANUFACTURING ISSUES | PRV - MEASURETROL | FM MEASURETROL(SIZE 080NB) | BLUETOOTH CARD NEEDS TO REPLACE. | VISITED AT SITE ANFD REPLACED THE BLUTOOTH CARD. | VISITED AT SITE ANFD REPLACED THE BLUTOOTH CARD. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFMM23194202953 | 2024-03-27 | 2024-04-03 | Cust 1161 | MUMBAI | Steam System Division | CLOSED | MANUFACTURING ISSUES | PRV - MEASURETROL | FM MEASURETROL(SIZE 080NB) | BLUETOOTH CARD NEEDS TO REPLACE. | VISITED AT SITE ANFD REPLACED THE BLUTOOTH CARD. | VISITED AT SITE ANFD REPLACED THE BLUTOOTH CARD. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFMM23194202953 | 2024-03-27 | 2024-04-03 | Cust 1161 | MUMBAI | Steam System Division | CLOSED | MANUFACTURING ISSUES | PRV - MEASURETROL | FM MEASURETROL(SIZE 080NB) | BLUETOOTH CARD NEEDS TO REPLACE. | VISITED AT SITE ANFD REPLACED THE BLUTOOTH CARD. | VISITED AT SITE ANFD REPLACED THE BLUTOOTH CARD. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFMM24007812954 | 2024-04-18 | 2024-04-22 | Cust 1163 | MUMBAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | GPCB connectivity require | visited and done the GPCB connectivity with the help of our Connectivity team. | visited and done the GPCB connectivity with the help of our Connectivity team. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFMM24007812954 | 2024-04-18 | 2024-04-22 | Cust 1163 | MUMBAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | GPCB connectivity require | visited and done the GPCB connectivity with the help of our Connectivity team. | visited and done the GPCB connectivity with the help of our Connectivity team. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFMM24007812954 | 2024-04-18 | 2024-04-22 | Cust 1163 | MUMBAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | GPCB connectivity require | visited and done the GPCB connectivity with the help of our Connectivity team. | visited and done the GPCB connectivity with the help of our Connectivity team. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFMM24007812954 | 2024-04-18 | 2024-04-22 | Cust 1163 | MUMBAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | GPCB connectivity require | visited and done the GPCB connectivity with the help of our Connectivity team. | visited and done the GPCB connectivity with the help of our Connectivity team. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFMM24033052955 | 2024-06-05 | 2024-06-09 | Cust 1164 | MUMBAI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | CPU Display card for Multiparameter(AWA-CPU1) | ANALYZER SPARES REPLACEMENT & CALIBRATION REQUIRED AT SITE | Done the replacement of the spares. Now ok. | Done the replacement of the spares. Now ok. | YES | No Trip | Process | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFMM24037292956 | 2024-06-14 | 2024-06-14 | Cust 1159 | MUMBAI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | STATION NAME NEEDS TO BE CHANGED ON GPCB SREVER, | Please check now the Station name has been changed by the GPCB team for both units . | Please check now the Station name has been changed by the GPCB team for both units . | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFMM24037942957 | 2024-06-17 | 2024-06-24 | Cust 1164 | MUMBAI | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | MULTIPARAMETER ANALYSER | 12 V POWER SUPPLY CARD for Multiparameter(AWA-PS1) | Require calibration for the analyser | Done the calibration work and submitted the report | Done the calibration work and submitted the report | EXTENDED | No Trip | Process | N |  | Analytical - CHAKAN
AFMM24050902958 | 2024-07-09 | 2024-07-15 | Cust 1161 | MUMBAI | Steam System Division | CLOSED | QUALITY PROBLEM | PRV - MEASURETROL | FM MEASURETROL(SIZE 080NB) | Block-4 PRV (80 NB Measurtrol) is also not working since the last two days, PRV is not getting open & due to that not getting required downstream pressure. Customer has bypassed the PRV & running the plant on bypass. But due to that safety valve is getting pop-up continuously. Kindly arrange a visit or permanent solution as the customer is now losing patience. | Reference RF OA Number 5525103010 has been booked for M/s. CENTAUR PHARMACEUTICALS PVT. LTD. Items: 1) 080EVLV-Q1EKS412MI-L1HWNSNNX = Qty01No 2) 040EVLV-Q1EFS412MI-M1HENSNNX = Qty01No 3) PID-RFWSM202101 = Qty02Nos | Reference RF OA Number 5525103010 has been booked for M/s. CENTAUR PHARMACEUTICALS PVT. LTD. 1) 080EVLV-Q1EKS412MI-L1HWNSNNX 2) 040EVLV-Q1EFS412MI-M1HENSNNX 3) PID-RFWSM202101 | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFMM24060652959 | 2024-07-27 | 2024-07-29 | Cust 1159 | MUMBAI | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | DO sensor is not showing correct reading | DO sensor cap need to change due to DO reading abnormal. | DO sensor cap need to change due to DO reading abnormal. | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMM24087552960 | 2024-09-12 | 2024-09-16 | Cust 1165 | MUMBAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Electrode | Chemotrode 150(HT238 762) | ph sensor need to commissioning. | done the testing with proper wiring. now ok. | done the testing with proper wiring. now ok. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFMM24087552960 | 2024-09-12 | 2024-09-16 | Cust 1165 | MUMBAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Electrode | Chemotrode 150(HT238 762) | ph sensor need to commissioning. | done the testing with proper wiring. now ok. | done the testing with proper wiring. now ok. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFMM24089532961 | 2024-09-17 | 2024-09-18 | Cust 1166 | MUMBAI | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | General Panel | VP6 CABLE(1) | DPI202 commissioning required | done the setting in the DPI 202. now ok. | done the setting in the DPI 202. now ok. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFMM24097922962 | 2024-10-02 | 2024-10-04 | Cust 1163 | MUMBAI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data upload issue. | Now data is live on cpcb server | Now data is live on cpcb server | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFMM24103842963 | 2024-10-15 | 2024-10-19 | Cust 1167 | MUMBAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Transmitter Aquacon | Aquacon 2000(AQCN 2000 p) | commissioning required | Done the recalibration and taken trial. Now ok. | Done the recalibration and taken trial. Now ok. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFMM24103842963 | 2024-10-15 | 2024-10-19 | Cust 1167 | MUMBAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Transmitter Aquacon | Aquacon 2000(AQCN 2000 p) | commissioning required | Done the recalibration and taken trial. Now ok. | Done the recalibration and taken trial. Now ok. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFMM24103842963 | 2024-10-15 | 2024-10-19 | Cust 1167 | MUMBAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Transmitter Aquacon | Aquacon 2000(AQCN 2000 p) | commissioning required | Done the recalibration and taken trial. Now ok. | Done the recalibration and taken trial. Now ok. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFMM24113992964 | 2024-11-07 | 2024-11-08 | Cust 1168 | MUMBAI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB connectivity needs to be done. | we took access of the black box and completed CPCB connectivity and now data is live on CPCB server and the same is verified by @Maurya, Sarvesh . | we took access of the black box and completed CPCB connectivity and now data is live on CPCB server and the same is verified by @Maurya, Sarvesh . | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFMM24115512965 | 2024-11-11 | 2024-11-11 | Cust 1168 | MUMBAI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | GPCB connectivity needs to be done. | we took access of the black box and completed GPCB connectivity for RO_Permeate_from_GSTPL, now data is live on the GPCB server. | we took access of the black box and completed GPCB connectivity for RO_Permeate_from_GSTPL, now data is live on the GPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFMM24116732966 | 2024-11-12 | 2024-11-13 | Cust 1169 | MUMBAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | commissioning require of DO | done the commissioning of the DO sensor. | done the commissioning of the DO sensor. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFMM24116732966 | 2024-11-12 | 2024-11-13 | Cust 1169 | MUMBAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | commissioning require of DO | done the commissioning of the pH sensor. | done the commissioning of the DO sensor. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFMM24116732966 | 2024-11-12 | 2024-11-13 | Cust 1169 | MUMBAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | commissioning require of DO | done the commissioning of the pH sensor. | done the commissioning of the DO sensor. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFMM24127952967 | 2024-12-02 | 2024-12-04 | Cust 1169 | MUMBAI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB connectivity needs to be done. | we took access of the black box and completed CPCB connectivity, now data is live on CPCB server. | we took access of the black box and completed CPCB connectivity, now data is live on CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFMM24129512968 | 2024-12-04 | 2024-12-05 | Cust 1169 | MUMBAI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | cpcb connectivity needs to be done. | we took access of the black box and completed CPCB connectivity, now data is live on CPCB server. | we took access of the black box and completed CPCB connectivity, now data is live on CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFMM24131942969 | 2024-12-09 | 2024-12-12 | Cust 1169 | MUMBAI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | data mismatch issue. | As of now we are closing this complaint and will reopen it once the hardware side issue is resolved. | As of now we are closing this complaint and will reopen it once the hardware side issue is resolved. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFMM24137392970 | 2024-12-21 | 2024-12-23 | Cust 1170 | MUMBAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Smartpro Electronics Assembly(Smartpro-Elc-Assembly) | TDS system commissioning require | Done the commissioning of the TDS sensor. Given drawing for the same | Done the commissioning of the TDS sensor. Given drawing for the same | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFMM24137432971 | 2024-12-21 | 2024-12-23 | Cust 1169 | MUMBAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | SCADA Software | Envirosense Software for CPCB(ESENSE1) | GPCB and CPCB connectivity require | Done the connectivity through CPCB team. Now it is ok. | Done the connectivity through CPCB team. Now it is ok. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFMM24137772972 | 2024-12-22 | 2024-12-22 | Cust 1171 | MUMBAI | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | MULTIPARAMETER ANALYSER | Current output card-FM(1) | ONLINE EFFLUENT MONITORING SYSTEM OUTPUT FEEDBACK NOT SHOWING CORRECT | Old COD card was removed and new COD card is installed | 1) Old PH sensor was removed and new PH sensor is installed due to error occurring in previous SENSOR 2)Old COD card was removed and new COD card is installed | NO | No Trip | Process | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFMM24137772972 | 2024-12-22 | 2024-12-22 | Cust 1171 | MUMBAI | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | MULTIPARAMETER ANALYSER | Current output card-FM(1) | ONLINE EFFLUENT MONITORING SYSTEM OUTPUT FEEDBACK NOT SHOWING CORRECT | PH sensor is installed due to error occurring in previous sensor | 1) Old PH sensor was removed and new PH sensor is installed due to error occurring in previous SENSOR 2)Old COD card was removed and new COD card is installed | NO | No Trip | Process | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFMM24139252973 | 2024-12-24 | 2024-12-26 | Cust 1169 | MUMBAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | DO sensor commissioning require | Done the commissioning of the sensor. now ok. | DOne the commissioning of the sensor. now ok. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFMM24140702974 | 2024-12-28 | 2024-12-29 | Cust 1172 | MUMBAI | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | MULTIPARAMETER ANALYSER | Cleaning Pump(1) | Multiparameter analyses cleaning pump not working properly. | We found that the pump was damaged. | We found that the pump was damaged. | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMM24141662975 | 2024-12-29 | 2024-12-29 | Cust 1171 | MUMBAI | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | MULTIPARAMETER ANALYSER | mA Output card for Multiparameter(AWA-4-20mA) | ONLINE EFFLUENT MONITORING SYSTEM OUTPUT FEEDBACK NOT SHOWING CORRECT | ) I changed the COD range in the plc panel ,then on the PLC I added a range of 0mg/l for 4 mA and 800 mg/l for 20 mA 2) I changed the BOD ranged 0 mg/l for 4 mA and 400 mg/l for 20 mA | ) I changed the COD range in the plc panel ,then on the PLC I added a range of 0mg/l for 4 mA and 800 mg/l for 20 mA 2) I changed the BOD ranged 0 mg/l for 4 mA and 400 mg/l for 20 mA | NO | No Trip | Process | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFMM24142612976 | 2024-12-30 | 2025-01-03 | Cust 1162 | MUMBAI | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | Aqua4Trans | PH input card(1) | pH sensor is not showing correct reading | found issue in the pH sensor. submitted the offer. | found issue in the pH sensor. submitted the offer. | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMM24149502977 | 2025-01-12 | 2025-01-12 | Cust 1173 | MUMBAI | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Multi parameter analyser is not working fine . | spare parts replace and complete servicing done . | spare parts replace and complete servicing done . | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMM24156122978 | 2025-01-25 | 2025-01-30 | Cust 1174 | MUMBAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | pH commissioning require | Done the pH sensor commissioning. Control valve offer is submitted. | Done the pH sensor commissioning. Control valve offer is submitted. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFMM24156122978 | 2025-01-25 | 2025-01-30 | Cust 1174 | MUMBAI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | pH commissioning require | Done the pH sensor commissioning. Control valve offer is submitted. | Done the pH sensor commissioning. Control valve offer is submitted. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFMM24157632979 | 2025-01-27 | 2025-01-27 | Cust 1173 | MUMBAI | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | MULTIPARAMETER ANALYSER | Water Detector for Multiparameter(AWA-WD) | multiparameter analyser reading not showing on online . | the signal of PH 2 continue to show high and an error is showing . | the signal of PH 2 continue to show high and an error is showing . | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMM24171182980 | 2025-02-24 | 2025-02-25 | Cust 1173 | MUMBAI | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | MULTIPARAMETER ANALYSER | LCD of Multiparameter Display(1) | Multiparameter display not working . | i replaced the PH 2 detector and the display, as the latter was experiencing data initialization problems. | i replaced the PH 2 detector and the display, as the latter was experiencing data initialization problems. | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMM24171182980 | 2025-02-24 | 2025-02-25 | Cust 1173 | MUMBAI | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | MULTIPARAMETER ANALYSER | PH2 detector(1) | photo detector is not working. Due to this, the signal of PH 2 continues to show high. | i replaced the PH 2 detector and the display, as the latter was experiencing data initialization problems. | i replaced the PH 2 detector and the display, as the latter was experiencing data initialization problems. | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMM24180322981 | 2025-03-10 | 2025-03-13 | Cust 1171 | MUMBAI | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | online effluent monitoring system out put feedback Ph not showing correct. | Ph output cord not working | Ph output cord not working | NO | No Trip | Process | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFMM25024602982 | 2025-05-25 | 2025-05-25 | Cust 1175 | MUMBAI | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | Ph Tranmitter Smartpro | Smartpro Electronics Assembly(Smartpro-Elc-cbl) | installation of single line plate mounted PH measurement system & boiler observation . | we successfully installed the plate mounted ph measurement system . but PH not show accurate .so i wiil check next visit . | we successfully installed the plate mounted ph measurement system . but PH not show accurate .so, i will check next visit. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFMM25026802983 | 2025-05-28 | 2025-06-13 | Cust 1175 | MUMBAI | Boiler Division | CLOSED | PRODUCT FAILURE | Ph Electrode | Polilyte PRO 120(HT238 411) | The readings consistently showed between 1.98 and 2.2. When we tested it in pH 7 water, the reading displayed as 0 pH, sr no .8094 | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210145209 LR Dt - 12-JUN-2025 | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210145209 LR Dt - 12-JUN-2025 | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFMM25031122984 | 2025-06-04 | 2025-06-07 | Cust 1176 | MUMBAI | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | pH Transmitter Aquamon | Aquamon Electronics Assembly(ph ele-set-2000) | ph analyzer installation and commissioning. | checked the packing material for the pH sampling and dosing system, and all materials were found to be in order. Regarding the next steps, we kindly request that you install the panel with a 230V power supply at our designated location as soon as possible. Additionally, please provide the pH sampling point as per our discussion. | all materials were found to be in order. we kindly request that you install the panel with a 230V power supply at our designated location as soon as possible. and pH sampling point. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFMY23043542985 | 2023-06-21 | 2023-06-21 | Cust 1177 | MYANMAR | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | PH1 filter(1) | PH1 filter replacement. | Finished, system working well and reading ok. | monitor the result for 1 week on sampling waste water. | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY23067382986 | 2023-08-02 | 2023-08-02 | Cust 1178 | MYANMAR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Monthly service | Finished | finished | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY23067392987 | 2023-08-02 | 2023-08-02 | Cust 1179 | MYANMAR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Monthly service for July | Finished, | Need to change the Xenon power supply. | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY23067402988 | 2023-08-02 | 2023-08-02 | Cust 1180 | MYANMAR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | monthly service | Finished | Finished | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY23068002989 | 2023-08-03 | 2023-08-03 | Cust 1180 | MYANMAR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Enviroconnect software hanging, and how to download the report. | The PC working well and get back connection with analyzer, | - Don`t turn off the PC and enviroconnect software. - replace the cleaning pump head. | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY23078662990 | 2023-08-25 | 2023-08-25 | Cust 1178 | MYANMAR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC for on-line waste water monitoring system for August | Finished. | Finished | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY23078742991 | 2023-08-25 | 2023-08-25 | Cust 1179 | MYANMAR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC for On-line waste water monitoring system, Quarterly calibration | - xenon power supply board does not work, so need to change urgently. - calibration will be done after changing the spare parts. | Finished for August, | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY23078762992 | 2023-08-25 | 2023-08-25 | Cust 1180 | MYANMAR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC of on-line waste water monitoring system, quarterly calibration. | Finished, need to check the TDS sensor. | Need to check TDS sensor, | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY23092202993 | 2023-09-22 | 2023-09-22 | Cust 1181 | MYANMAR | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | system commissioning | Finished commissioning, system working well and reading OK. | Finished | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY23098162994 | 2023-10-01 | 2023-10-01 | Cust 1178 | MYANMAR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Monthly service | Done , Finished | finsihed. | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY23098172995 | 2023-10-01 | 2023-10-01 | Cust 1180 | MYANMAR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Monthly service | Done, Finished | finsihed | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY23099032996 | 2023-10-03 | 2023-10-03 | Cust 1179 | MYANMAR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Monthly service | Done , Finished. waiting the spare parts for xenon power supply board. | Finished | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY23105162997 | 2023-10-16 | 2023-10-23 | Cust 1180 | MYANMAR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | monthly service and spare parts replacements | Finished, | Done | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY23108472998 | 2023-10-24 | 2023-10-24 | Cust 1178 | MYANMAR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC, Monthly service for October, Multi Parameter Analyzer (online waste water monitoring system) | Done, Finished | Done. on 24 Oct. | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY23119802999 | 2023-11-15 | 2023-11-15 | Cust 1180 | MYANMAR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Monthly service | Done Finished, calibrated the PH, TSS, COD . Reading good. need to check the TDS sensor and transmitter, | Finished. Calibration and spare parts change. | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY23122603000 | 2023-11-21 | 2023-11-21 | Cust 1178 | MYANMAR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC service for the Multi analyzer , November, Quarterly calibration. | Finished. | calibration result will be sent via email | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY23124043001 | 2023-11-24 | 2023-11-24 | Cust 1180 | MYANMAR | Analytical - CHAKAN | CLOSED | AMC | Aqua4Trans | Aqua4Trans(1) | troubleshooting the TDS sensor and transmitter of Multi Analyzer, online waste water monitoring. | - cannot check the sensor, postponed to next week - PC online issue solved. | postponed to next week. | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY23126303002 | 2023-11-27 | 2023-12-02 | Cust 1179 | MYANMAR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC monthly service of online waste water monitoring for Nov, Quarterly calibration and spare parts replacement. | finished. Done the calibration. Done the spare parts changing. | Finished and working well . | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY23126993003 | 2023-11-28 | 2023-11-28 | Cust 1180 | MYANMAR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Aqua4Trans | Aqua4 trans-A4TW033000(1) | checking the TDS sensor and transmitter for data showing wrongly. | transmitter showing - when short sensor , the data goes down to 20ppm and sensor in the air , show around 45 to 60ppm. also reading not correct with new TDS sensor. Need to changed the transmitter. | need to change the transmitter, AquaTDS (SMARTPro 8969P) | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY23135153004 | 2023-12-13 | 2023-12-13 | Cust 1181 | MYANMAR | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | installation guidelines for the WW online monitoring | Give the proper guidelines for installation | Done. | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY23136223005 | 2023-12-14 | 2023-12-18 | Cust 1180 | MYANMAR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Enviroconnect PC is hanging and very slow response. | the PC is working Ok, took software backup. | Done and Finished. | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY23138493006 | 2023-12-19 | 2023-12-22 | Cust 1180 | MYANMAR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Enviroconnect PC is hanging and very slow response. | we took access of the system and observer that the PC is working Ok | we took access of the system and observer that the PC is working Ok | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFMY23138793007 | 2023-12-19 | 2023-12-19 | Cust 1180 | MYANMAR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC Monthly service | - TDS transmitter have error , need to replace - PC hanging, need to upgrade the system . | Finished | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY23138803008 | 2023-12-19 | 2023-12-19 | Cust 1178 | MYANMAR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Monthly service for month December | - Cleaned sensor and flow cell chamber. - repaired the level float . | Done Finished. | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY23152833009 | 2024-01-12 | 2024-01-12 | Cust 1179 | MYANMAR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC service , for Jan . On-Line waste water monitoring system | Reinstalled the Xenon power supply card , Done the monthly service for jan. test run and analyzer doing working well. | remining to check the EnviroConnect software on PC for TSS data. | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY23152843010 | 2024-01-12 | 2024-01-12 | Cust 1178 | MYANMAR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC service , for Jan . On-Line waste water monitoring system | Done | finished | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY23152853011 | 2024-01-12 | 2024-01-12 | Cust 1180 | MYANMAR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC service, for Jan. On-Line waste water monitoring. | Done | finished | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY23152863012 | 2024-01-12 | 2024-01-18 | Cust 1181 | MYANMAR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | no output voltage and failed electric circuit after 20days used. Xenon power supply card of Multi Parameter Analyzer | Material dispatch against RF OA 544240032 from Singapore | Material dispatch against RF OA 544240032 from Singapore | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFMY23154873013 | 2024-01-16 | 2024-01-19 | Cust 1179 | MYANMAR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | TSS data mismatch issue. | after changing the TSS mA range to 750 from the analyser side the data starts uploading correctly. | after changing the TSS mA range to 750 from the analyser side the data starts uploading correctly. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFMY23156893014 | 2024-01-22 | 2024-01-22 | Cust 1179 | MYANMAR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | TSS data showing on Enviroconnect and analyzer are difference. To check the software and Analyzer setting. | Solved the issue. - changed the mA output of TSS on the analyzer - changed the Range of TSS on the analyzer. | Finished. test ok. | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY23158733015 | 2024-01-24 | 2024-01-29 | Cust 1180 | MYANMAR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | auto power restart and then only show the pH and DO. no signal PH2 , error 8. need to check the Xenon power supply card and MUVX card. | Repaired the Xenon power supply card and output voltage around 500 V DC. the signal reading ok. adjustment for mA output and range. | signal reading ok. but need to keep the spare for the Xenon power supply card. | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY23176943016 | 2024-02-26 | 2024-02-26 | Cust 1179 | MYANMAR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Monthly service (quarterly calibration) | Done and Finished. | Done. | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY23179463017 | 2024-02-29 | 2024-02-29 | Cust 1180 | MYANMAR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Monthly service for February (quarterly calibration) | Calibrated COD, TSS, pH. reading OK, Finished | Done | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY23179473018 | 2024-02-29 | 2024-02-29 | Cust 1178 | MYANMAR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Monthly service and Quarterly calibration | Calibrated COD, TSS, pH. Reading ok. Finished. | Done | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY23183303019 | 2024-03-08 | 2024-08-09 | Cust 1182 | MYANMAR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Aqua4Trans | Aqua4Trans(1) | lcd display appears dark black line | Replacement material shipped on 26th June | Replacement material shipped on 26th June | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFMY23184583020 | 2024-03-12 | 2024-03-20 | Cust 1182 | MYANMAR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Aqua4Trans | Aqua4Trans(1) | Boiler Feed water pH monitoring system . Sample water line blocked and no drain water from system. | Done | Finidhed. | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY23190183021 | 2024-03-20 | 2024-03-20 | Cust 1179 | MYANMAR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Monthly service for March | Finished | Done | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY23190963022 | 2024-03-21 | 2024-03-21 | Cust 1178 | MYANMAR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Monthly service | Finished | Done | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY23190973023 | 2024-03-21 | 2024-03-21 | Cust 1180 | MYANMAR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Monthly service | Finished | Done | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY23193473024 | 2024-03-26 | 2024-03-26 | Cust 1178 | MYANMAR | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | MULTIPARAMETER ANALYSER | Cleaning Pump(1) | replacement of cleaning pump set with new. | replaced with new, Test run ok. | replaced with new, test run ok. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFMY23195073025 | 2024-03-28 | 2024-04-04 | Cust 1182 | MYANMAR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | pH Transmitter Aquamon | Ph Transmitter Aquamon 4000(PH-Trans-4000) | 8tph boiler pH monitoring (Sampling system with analyzer ) no water flow, | The high temperature safety shut off valve is closing and back pressure regulator valve is closing. so the sample water cannot flow throw the system. now open these two valves and checked the system working condition by operating the boiler feed water pump. Test run ok. | checked the system working condition by operating the boiler feed water pump. Test run ok. | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY24004273026 | 2024-04-09 | 2024-04-09 | Cust 1180 | MYANMAR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Monthly service , Multi parameter analyzer | Done. checked and cleaned the system.. | Done. | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY24004903027 | 2024-04-10 | 2024-04-10 | Cust 1178 | MYANMAR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Monthly service , | Done and Finished, | Finished. | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY24013523028 | 2024-04-29 | 2024-04-29 | Cust 1179 | MYANMAR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | monthly service | Finished on 26 April. | finished. | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY24035513029 | 2024-06-10 | 2024-06-10 | Cust 1178 | MYANMAR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | TSS reading negative value | cleaning flow cell chamber and zeroing. after that reading ok. | solved. reading ok. | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY24049813030 | 2024-07-07 | 2024-07-07 | Cust 1180 | MYANMAR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Monthly service for June | Finished , | need to replace the Xenon power supply card and LCD. | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY24049823031 | 2024-07-07 | 2024-07-07 | Cust 1178 | MYANMAR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | monthly service for June | finished | finished | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY24058803032 | 2024-07-24 | 2024-07-24 | Cust 1178 | MYANMAR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | July Monthly service for WW online monitoring system | Done. | need to keep the spare for the all tubing. | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY24058813033 | 2024-07-24 | 2024-07-31 | Cust 1180 | MYANMAR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | monthly service , replacing the LCD, xenon power supply card and CMOS battery. | site visit done | site visit done | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY24058823034 | 2024-07-24 | 2024-07-24 | Cust 1179 | MYANMAR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | breakdown service call, - COD reading too high and some error showing. | changed the PH 1 optical filter and PH1 photo detector. | Finished , done | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY24067073035 | 2024-08-05 | 2024-08-05 | Cust 1180 | MYANMAR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | monthly service, signal not stable. | Need to check the motherboard and CPU by replacing with new. | Need to check the motherboard and CPU by replacing with new. | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY24069853036 | 2024-08-12 | 2024-08-16 | Cust 1180 | MYANMAR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Analyzer troubleshooting for PH1 PH2 reading too high. | changed the card slot for MUVX card. after that signal PH1 & PH2 stable. | finished . | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY24072913037 | 2024-08-19 | 2024-08-22 | Cust 1180 | MYANMAR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Xenon power supply card power failure | solved, xenon card is good. after changed the card slot of MUVX card , the signal was stable and sample reading ok. | after service ,, the signal and sample reading ok. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFMY24073663038 | 2024-08-20 | 2024-08-23 | Cust 1180 | MYANMAR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | The customer, Emerald Brewery EBML, setup the new PC for Enviroconnect for waste water online monitoring (multi parameter analyzer). So need to install the environment software., | we took access of the system and completed the software installation and now the data is live on the Local + Myanmar server. | we took access of the system and completed the software installation and now the data is live on the Local + Myanmar server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFMY24089863039 | 2024-09-17 | 2024-09-17 | Cust 1182 | MYANMAR | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | Aqua4Trans | Aqua4Trans(1) | remove the old damage one and replace the new Aqua4Trans installation and commissioning. WHRB boiler ph and TDS monitoring system | the old transmitter is damaged the display. Removed the old one and installed the new Aqua4Trans and configured for pH and TDS. | Finished. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFMY24094263040 | 2024-09-26 | 2024-09-26 | Cust 1180 | MYANMAR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Monthly service | sample pump discharge pipe crack and waste water entering inside the panel. UV lamp, flow cell chamber and Optical filter 1&2, and cleaning pump are sink in the waste water. Cleaned it all. | Power off and wait for dry. | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY24094273041 | 2024-09-26 | 2024-09-26 | Cust 1180 | MYANMAR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | UV lamp , flow cell , Optical filter 1&2 are sink in the waste water, so cleaned it all and checking these parts after dry. | the system was dry after 1 day, checked the sample pump and cleaning pump motor which are working well. Failed the Xenon Lamp , Xenon power supply card, optical filter 1 & 2, PhotoX 100pF and 220pF, | Need to replace the following parts Xenon Lamp , Xenon power supply card, optical filter 1 &amp;amp; 2, PhotoX 100pF and 220pF, | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY24094583042 | 2024-09-27 | 2024-10-07 | Cust 1182 | MYANMAR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Electrode | FM1104(1) | ph high value , calibration fails all time, | failed the pH probe. | change new one | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY24110143043 | 2024-10-28 | 2024-10-30 | Cust 1181 | MYANMAR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MANYMAR DATA UPLOAD ISSUE FOR STANDARD DISTILLERY LTD | Now the Data is live on the Local server + Myanmar server. | Now the Data is live on the Local server + Myanmar server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFMY24112443044 | 2024-11-04 | 2024-11-04 | Cust 1181 | MYANMAR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Standard Distillery, RS485 converter install and EnviroConnect server down issue. | Solved the issue. by repairing from HO. remote support. | Finished . Done | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY24123203045 | 2024-11-25 | 2025-01-08 | Cust 1182 | MYANMAR | Steam System Division | CLOSED | PRODUCT FAILURE | CRS | CONDENSATE RECOVERY METER - MODEL CRM-485(CRM-485R) | CRMR2210687, damage in commissioning, | RF 138244949 under draft, on hold until further details | RF 138244949 under draft, on hold until further details | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFMY24127113046 | 2024-11-30 | 2024-11-30 | Cust 1178 | MYANMAR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Monthly service for Analyzer (online waste water monitoring system) | tube leakage and cracking of sample pump, need to change with new. PH signal too low, need to replace the quartz disc. water entering into the level Float, need to replace, | finished for November service, In next week, will replace the all tube.. | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY24132413047 | 2024-12-10 | 2024-12-10 | Cust 1178 | MYANMAR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | all Tube replacement of Analyzer, | replaced the all tube, Test run ok. no leakage. | test run ok. No leakage | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY24139183048 | 2024-12-24 | 2024-12-24 | Cust 1178 | MYANMAR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Monthly service | Done and finished. Need to replace the quartz disc. | Done. | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY24151983049 | 2025-01-16 | 2025-01-16 | Cust 1178 | MYANMAR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Monthly service for January | - PH signal too low because of quarts disc failed. - LCD black line and sport. | need to change the Quarts disc and LCD. | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY24164353050 | 2025-02-06 | 2025-02-06 | Cust 1178 | MYANMAR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | monthly service | - quartz disc failed so PH1 reading too low. - black line on the LCD. | need to change the quartz disc. | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY24164803051 | 2025-02-08 | 2025-02-10 | Cust 1178 | MYANMAR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Local and Myanmar serevr data uploded issue. | Taken up this issue with the AIPL team and they have updated the DAS with 4.3 version and now the Data is live. | Taken up this issue with the AIPL team and they have updated the DAS with 4.3 version and now the Data is live. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFMY24168243052 | 2025-02-16 | 2025-02-16 | Cust 1178 | MYANMAR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | EnviroConnect PC data receiving delay. Server error. | Solved by remote support from head office. | solved. | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY24168253053 | 2025-02-16 | 2025-02-16 | Cust 1179 | MYANMAR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Monthly service for February. | Finished . | Need to change the LCD and PH1 set. | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY24170593054 | 2025-02-21 | 2025-02-21 | Cust 1182 | MYANMAR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Electrode | Polilyte PRO 120(HT238 411) | pH sampling system calibration with Aqua4Trans pH transmitter, 16.8 tph and 8 tph | calibration failed, the transmitter is good. pH probe failed. | need to change the pH probe. | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY24170593054 | 2025-02-21 | 2025-02-21 | Cust 1182 | MYANMAR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Electrode | Polilyte PRO 120(HT238 411) | pH sampling system calibration with Aqua4Trans pH transmitter, 16.8 tph and 8 tph | calibration failed, the transmitter is good. pH probe failed. | need to change the pH probe. | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY24186483055 | 2025-03-25 | 2025-03-25 | Cust 1179 | MYANMAR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Monthly service for March | Finished. Need to replace the 24V power supply for Flacon MiNi, TDS transmitter. Also Need to replace the LCD. | Finished. | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY24186493056 | 2025-03-25 | 2025-03-25 | Cust 1178 | MYANMAR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC for online waste water monitoring system, Monthly service for March, and spare parts (quartz discs) replacement service. | Replaced the spare parts and done the monthly service. | Done | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY25011013057 | 2025-04-28 | 2025-04-28 | Cust 1178 | MYANMAR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | monthly service | Done | Done, The LCD is not good and need to keep the spare. | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY25011043058 | 2025-04-28 | 2025-04-28 | Cust 1179 | MYANMAR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | monthly service for April | done | Done, PH1 signal too low need to change the PH1 set. | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFMY25011073059 | 2025-04-28 | 2025-04-28 | Cust 1180 | MYANMAR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | monthly service for April. | 12 power supply card failed. so the external sample pump cannot work. | Need to change the 12V dc power supply for external sampling pump. | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFNG23034983060 | 2023-06-05 | 2023-06-26 | Cust 1183 | NAGPUR | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | MULTIPARAMETER ANALYSER | RS485 CARD for Multiparameter(AWA-RS485) | Found RS 485 card faulty | CARD SENT | MATERIAL HAS SENT | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFNG23037293061 | 2023-06-09 | 2023-06-09 | Cust 1184 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Data not updating on CPCB server. | Reported issue has been checked & found Black Box RS 485 terminal not working. Requested to procure the spare to establish the connectivity back. | closed | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG23039253062 | 2023-06-13 | 2023-06-18 | Cust 1185 | NAGPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | Dissolved Oxygen Analyzer Panel mounted(AQUADO4000P) | commissioning of SWAN make Portable DO analyzer. | Commissioning of DO analyzer is completed. | closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG23039753063 | 2023-06-14 | 2023-06-20 | Cust 1081 | NAGPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | VISI pro DO(VISI pro DO) | Commissioning of Optical DO sensor. | Commissioning of optical DO is completed at Unit #3. | closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG23039753063 | 2023-06-14 | 2023-06-20 | Cust 1081 | NAGPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | VISI pro DO(VISI pro DO) | Commissioning of Optical DO sensor. | Commissioning of optical DO is completed at Unit #4 | closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG23042513064 | 2023-06-19 | 2023-06-21 | Cust 1186 | NAGPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4 trans-A4TW033000(1) | Replacement received at site. Call for the commissioning of same. | Received replacement checked & found in good physical installation. Electrode, Transmitter installation completed & power on Analyzer successfully. Commissioning completed & taken Analyzer into service. Online working performance observed & found ok. | commissioning completed at site. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG23043203065 | 2023-06-21 | 2023-06-23 | Cust 1187 | NAGPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Analyzer showing zero values always. | Online support provided. Reported issue of Analyzer not working has been checked & resolved successfully. | closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG23047113066 | 2023-06-27 | 2023-07-10 | Cust 1186 | NAGPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Expert Services to attend the Multiparameter Analyzer. | Maintenance activities carried out & taken Analyzer into service. Online working of Analyzer observed & found working ok. | closed | NO | Paid | Core | N |  | Analytical - CHAKAN
AFNG23047113066 | 2023-06-27 | 2023-07-10 | Cust 1186 | NAGPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Expert Services to attend the Multiparameter Analyzer. | Maintenance activities carried out & taken Analyzer into service. Online working of Analyzer observed & found working ok. | closed | NO | Paid | Core | N |  | Analytical - CHAKAN
AFNG23047113066 | 2023-06-27 | 2023-07-10 | Cust 1186 | NAGPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Expert Services to attend the Multiparameter Analyzer. | After analysis, observed that issue in Display card, Detector assembly. Spares required to put the Analyzer into operational. | closed | NO | Paid | Core | N |  | Analytical - CHAKAN
AFNG23047403067 | 2023-06-28 | 2023-06-29 | Cust 708 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data uploaded issue. | Here we took access of the and found that the internet is not pinging till the BB and now the same issue has been resolved by the customer's IT team and the data is live. | Issue has been resolved and data is live hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG23047693068 | 2023-06-28 | 2023-06-28 | Cust 1055 | NAGPUR | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | Connectivity of pH,Conductivity analyzer to eversense software. | Connected the pH & conductivity to eversense server. | Connected the pH & conductivity to eversense server. | NO | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFNG23050613069 | 2023-07-03 | 2023-07-07 | Cust 1188 | NAGPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | Commissioning of Conductivity transmitter and sensor at M/s JSW steels kalmeshwar. | Commissioning is completed by providing online support to customer. | closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG23051573070 | 2023-07-05 | 2023-07-09 | Cust 1187 | NAGPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Display received at site. Call for commissioning of same. | After Display replacement aslo issue in Keypad was observed. Door With Acrylic needed at site. The same was conveyed to division. | closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG23056403071 | 2023-07-14 | 2023-07-14 | Cust 1186 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data uploaded issue. | We have manually changed the time in BB to sync to real-time and now data started uploading on the CPCB server so we ask them to open ntp port from their network for automatic time synchronization. | Data is live hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG23059643072 | 2023-07-21 | 2023-07-24 | Cust 1189 | NAGPUR | Analytical - CHAKAN | CLOSED | AMC | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Installation and configuration of FM black box. | Installation and configuration of FM black box is completed. | closed | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG23060813073 | 2023-07-24 | 2023-07-24 | Cust 1187 | NAGPUR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | DO Sensor Oxysens DO is not working. | Checked and suspected the issue in DO sensor, suggested to send the sensor to our chakan plant for warranty repair/ replacement. | closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG23063303074 | 2023-07-28 | 2023-07-28 | Cust 1067 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data uploaded issue. | We did some changes in server configuration end and now data is live on all the server. | We did some changes in server configuration end and now data is live on all the server. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFNG23067813075 | 2023-08-03 | 2023-08-04 | Cust 466 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | RSPCB data upload issue. | Now data is live on RSPCB server, hence we are closing this complaint. | Now data is live on RSPCB server, hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG23067903076 | 2023-08-03 | 2023-08-04 | Cust 1190 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB camera configuration needs to be done. | Camera is live on CPCB server, hence we are closing this complaint. | Camera is live on CPCB server, hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG23070153077 | 2023-08-08 | 2023-08-14 | Cust 1187 | NAGPUR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Display card and door with acrylic received at site. Call for the commissioning of same. | Spare replaced at site & taken Analyzer into service. Online working performance observed & found ok. | Closed. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG23071123078 | 2023-08-09 | 2023-08-11 | Cust 1187 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data uploaded issue. | I took access of the Kara unit and updated the version of BB to the latest version and now the data is live on the server And for Nimora Plant we found that the data is not coming till the BB and we also checked the data by bypassing the Isolator but the data is not coming so after that we checked the data directly on modscan and the data is coming on modscan. Hence it seems that the RS485 Isolator is damaged so needs to be replaced the same. | Issue has been resolved hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG23071193079 | 2023-08-09 | 2023-08-14 | Cust 1187 | NAGPUR | Analytical - CHAKAN | CLOSED | DIGITAL CONEECTIVITY | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Pre-commissioning activities Multi Parameter Analyzer. | Pre-commissioning activities carried out at site & Analyzers made ready for CPCB connectivity. | Closed. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG23071203080 | 2023-08-09 | 2023-08-14 | Cust 1187 | NAGPUR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Data not uploading on Server. | While at site reported issue of connectivity has been checked, observed that not functioning & needs replacement. | closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG23074263081 | 2023-08-17 | 2023-08-18 | Cust 1191 | NAGPUR | Analytical - CHAKAN | CLOSED | DIGITAL CONEECTIVITY | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Auto reboot patch needs to be update. | We took access of all 3 Black Boxes and upgraded them with the Latest patch. | We took access of all 3 Black Boxes and upgraded them with the Latest patch. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG23076563082 | 2023-08-21 | 2023-08-28 | Cust 1192 | NAGPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC visit. | Preventive maintenance carried out & taken Multiparameter into service. Online working performance observed & found ok. | closed | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFNG23076693083 | 2023-08-21 | 2023-08-23 | Cust 497 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data uploaded issue. | We took access of the BB and made some changes in the configuration so now the data is live on the server. | We took access of the BB and made some changes in the configuration so now the data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG23080653084 | 2023-08-28 | 2023-09-18 | Cust 1193 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPB data uploaded issue. | ur development team updated RTDAS the version, and Now the data is uploading. | Our development team updated RTDAS the version, and Now the data is uploading. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG23081233085 | 2023-08-29 | 2024-01-18 | Cust 1184 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | RTDMS portal data upload issue. | Customers wants to remove device from RTDMS portal and, already sent a mail to CPCB regarding the same. (srchiefchemistmppgcl.sarni@gmail.com) | Customers wants to remove device from RTDMS portal and, already sent a mail to CPCB regarding the same. (srchiefchemistmppgcl.sarni@gmail.com) | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG23082973086 | 2023-09-01 | 2023-09-04 | Cust 1194 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data upload issue. | We took the access to the system and restart the services, after that the data is live on MPCB server. | We took the access to the system and restart the services, after that the data is live on MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG23084693087 | 2023-09-05 | 2023-09-06 | Cust 777 | NAGPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Electrode Cable(Ph-Cab) | Call for commissioning of 03 nos pH & 03 Nos Conductivity Tansmitters. | Supplied Analyzer are very old. Transmitters power on Successfully. Out of 6 ,01 Transmitter found not working & needs replacement. All sensor needs replacement. List to procure the material has been shared with customer. | closed | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG23085443088 | 2023-09-07 | 2023-09-11 | Cust 1186 | NAGPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | VISI pro DO(VISI pro DO) | Commissioning of Visitrace DO 120 Sensors at Unit 8,9 & 10 SWAS. | Commissioning of Visitrace DO 120 Sensors at Unit 8 is completed. | closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG23085443088 | 2023-09-07 | 2023-09-11 | Cust 1186 | NAGPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | VISI pro DO(VISI pro DO) | Commissioning of Visitrace DO 120 Sensors at Unit 8,9 & 10 SWAS. | Commissioning of Visitrace DO 120 Sensors at Unit 9 SWAS is completed. | closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG23085443088 | 2023-09-07 | 2023-09-11 | Cust 1186 | NAGPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | VISI pro DO(VISI pro DO) | Commissioning of Visitrace DO 120 Sensors at Unit 8,9 & 10 SWAS. | Installation, tubing, wiring and configuration is completed at Unit 10. final commissioning is pending because of samples unavailability, as the plant in shut down. | closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG23085503089 | 2023-09-08 | 2023-09-15 | Cust 1195 | NAGPUR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | M/S VISAG BIOFUELS PVT LTD. VILLAGE KHAPA, TESHIL -WARASEONI, DISTRICT -BALAGHAT, MADHYA PRADESH (481331) Site Contact Person :- Mr.Lakhan Kumar :- Mobile No. 9005283203. | Plant under shitdown till Monday. Visit planned on Monday i.e.18th Sept | Plant under shitdown till Monday. Visit planned on Monday i.e.18th Sept | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFNG23086893090 | 2023-09-11 | 2023-09-12 | Cust 528 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB and MPPCB. | Data was already updating on both the enviroconnect and MPPCB server. | Data was already updating on both the enviroconnect and MPPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG23088733091 | 2023-09-14 | 2023-09-15 | Cust 497 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data uploaded issue. | We did the software reinstallation + local + CPCB configuration, now data is live on CPCB server. for MPCB customers provides access on monday morning as of now we are closing this complaint. | We did the software reinstallation + local + CPCB configuration, now data is live on CPCB server. for MPCB customers provides access on monday morning as of now we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG23091123092 | 2023-09-20 | 2023-09-24 | Cust 497 | NAGPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC visit. | Maintenance activities carried out & taken Analyzer into service. Online working performance observed & found ok. | closed | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFNG23092753093 | 2023-09-25 | 2023-09-25 | Cust 1189 | NAGPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Maintenance and calibration of Multiparameter analyzer. | Maintenance and calibration of Multiparameter analyzer is carried out. | closed | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFNG23093263094 | 2023-09-25 | 2023-09-25 | Cust 1194 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data was not updating on CPCB Portal. | Took access of the system and done configuration changes. | Took access of the system and done configuration changes. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFNG23093663095 | 2023-09-26 | 2023-10-25 | Cust 555 | NAGPUR | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | Ph Electrode | FM1106(1) | Supplied PH electrode is found in leakage and dry condition | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210045615 LR Dt - 11-OCT-2023 | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210045615 LR Dt - 11-OCT-2023 | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFNG23094583096 | 2023-09-27 | 2023-09-30 | Cust 497 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data break issue | We took access of the system and updated the .jar file in the Local configuration so the data is now live on the serve | The issue has been resolved hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG23096263097 | 2023-09-29 | 2023-10-02 | Cust 555 | NAGPUR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Electrode | Easycontrol 225(HT238 542) | o.a. 5524103841 number . New PH electrode supplied at Site failed / Not working. Required replacement | Already logged . | Already logged | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFNG23099683098 | 2023-10-04 | 2023-10-05 | Cust 1194 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB Server for Effluent. | Took access of the system and restarted services and data is live on LOCAL+CPCB+MPCB. | Took access of the system and restarted services and data is live on LOCAL+CPCB+MPCB. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG23099693099 | 2023-10-04 | 2023-10-06 | Cust 1196 | NAGPUR | Analytical - CHAKAN | CLOSED | APPLICATION ISSUE | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | Zero reading issue in Oxysens DO sensor. | Checked and cleaned the sensor. observed the response in air, tap water and in process. sensor is found responding ok. | closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG23099733100 | 2023-10-04 | 2023-10-06 | Cust 1186 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | BB configuration. | we took access of all 3 Black Box and completed the BB configuration + CPCB connectivity so now data is live on the server. | we took access of all 3 Black Box and completed the BB configuration + CPCB connectivity so now data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG23099773101 | 2023-10-04 | 2023-10-07 | Cust 1186 | NAGPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Commissioning of FM BB at ETP, STP and AWR locations | Visited the site from 4 to 6 oct 23 commissioning of FM BB @ ETP, STP & AWR done . Patch Modification balance in ETP analyzer | FM BB commissioning done ETP,STP &AWR Plant | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG23099773101 | 2023-10-04 | 2023-10-07 | Cust 1186 | NAGPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Commissioning of FM BB at ETP, STP and AWR locations | Visited the site from 4 to 6 oct 23 commissioning of FM BB @ ETP, STP & AWR done . Patch Modification balance in ETP analyzer | FM BB commissioning done ETP,STP &AWR Plant | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG23099773101 | 2023-10-04 | 2023-10-07 | Cust 1186 | NAGPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Commissioning of FM BB at ETP, STP and AWR locations | Visited the site from 4 to 6 oct 23 commissioning of FM BB @ ETP, STP & AWR done . Patch Modification balance in ETP analyzer | FM BB commissioning done ETP,STP &AWR Plant | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG23099873102 | 2023-10-05 | 2023-10-06 | Cust 1186 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB and MPCB data uploaded issue. | Issue has been resolved hence we are closing this issue. | Issue has been resolved hence we are closing this issue. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG23100523103 | 2023-10-06 | 2023-10-09 | Cust 497 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data breaking on MPPCB server. | Took access of the system and and update patch 55 and data is live without breaking on MPCB Server. | Took access of the system and and update patch 55 and data is live without breaking on MPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFNG23100573104 | 2023-10-06 | 2023-10-09 | Cust 1197 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue. | RS485 card faulty, send to FM for repair, as of now we are closing this complaint. | RS485 card faulty, send to FM for repair, as of now we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG23100663105 | 2023-10-06 | 2023-10-08 | Cust 1197 | NAGPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 1st AMC visit. | Maintenance activities carried out & taken Analyzer into service. Online working performance observed & found ok. RS 485 Card found malfunctioning & needs replacement. | AMC visit done. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFNG23103463106 | 2023-10-12 | 2023-10-17 | Cust 1198 | NAGPUR | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | TOC ANALYSER | TOC 1950 Single StreamANALYSER ANATEL(1950 TOC) | Contractor finalization,Site inspection etc | Discussion with contractor & Customer done for erection & Commissioning activities | Visit done , Discussion done | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG23106843107 | 2023-10-19 | 2023-11-27 | Cust 1199 | NAGPUR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Aqua4Trans | Aqua4Trans(1) | 1. PO number - 3100204374 Material received on 22nd August 2023. Qty Reqd - 01 Nos. 2. Sr number - 22A4T0562-P013100 Qty Reqd - 01 No. Failed during warranty period. | Replacement material has been booked through TCI EXPRESS vide LR No - 070900222 LR Dt - 23-NOV-2023 | Replacement material has been booked through TCI EXPRESS vide LR No - 070900222 LR Dt - 23-NOV-2023 | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFNG23106843107 | 2023-10-19 | 2023-11-27 | Cust 1199 | NAGPUR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Aqua4Trans | Aqua4Trans(1) | 1. PO number - 3100204374 Material received on 22nd August 2023. Qty Reqd - 01 Nos. 2. Sr number - 22A4T0562-P013100 Qty Reqd - 01 No. Failed during warranty period. | Replacement material has been booked through TCI EXPRESS vide LR No - 070900222 LR Dt - 23-NOV-2023 | Replacement material has been booked through TCI EXPRESS vide LR No - 070900222 LR Dt - 23-NOV-2023 | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFNG23106943108 | 2023-10-19 | 2023-10-25 | Cust 1186 | NAGPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Commissioning and CPCB connectivity of FMBB at ETP1. | Visited the site on 20.10,23 (1) Spares replacement done STP analyzer. (2) System given on line for Connectivity of ETP analyzer for Configuration of FM BB | Multiparameter spares replacement done at STP analyzer | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG23106953109 | 2023-10-19 | 2023-10-26 | Cust 1186 | NAGPUR | Analytical - CHAKAN | CLOSED | APPLICATION ISSUE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | zero reading. Spares installation. | ETP 1 analyzer tubing changed. zero calibration carried out. analyzer is now working ok. | closed | NO | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG23107213110 | 2023-10-20 | 2023-10-23 | Cust 708 | NAGPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | pH measurement showing erratic reading. | After inspection, Problem suspected in pH Electrode. M/s FMPL will take the corrective action to resolve the same. | closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG23108233111 | 2023-10-24 | 2023-10-30 | Cust 1187 | NAGPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | troubleshooting and calibration of inlet analyzer. | replaced the detectors and xenon lamp power supply. analyzer is now working ok but the current output cards are faulty. | closed | NO | Paid | Core | N |  | Analytical - CHAKAN
AFNG23108243112 | 2023-10-24 | 2023-10-30 | Cust 1187 | NAGPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | calibration of disposal/outlet multiparameter analyzer. | calibration of disposal/outlet multiparameter analyzer is carried out successfully. | closed | NO | Paid | Core | N |  | Analytical - CHAKAN
AFNG23108833113 | 2023-10-25 | 2023-10-27 | Cust 1186 | NAGPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning of ETP-4 Analyzer. | Installation checked ,Power supply checked & power ON Multiparameter Analyzer successfully. Necessary parameter setting done & successfully completed the commissioning as per attached report. | commissioning completed at site. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG23108843114 | 2023-10-25 | 2023-10-27 | Cust 1186 | NAGPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Expert visit to check the Multiparameter Analyzer installed at ETP- 1, ETP - 2, ETP-3 & ETP-4. | All 04 ETP Analyzer checked & necessary maintenance activities carried out at site as per attached minutes. | AMC visit done. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFNG23109323115 | 2023-10-26 | 2023-10-28 | Cust 1200 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data uploaded issue. | We took access of the Stack Analyser system and updated the DAS locally so now the data is live on the server. | We took access of the Stack Analyser system and updated the DAS locally so now the data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG23112063116 | 2023-10-30 | 2023-11-01 | Cust 1201 | NAGPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Smartpro-con) | Commissioning of conductivity Transmitter. | Installation checked & commissioned successfully conductivity transmitter. | closed | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG23112063116 | 2023-10-30 | 2023-11-01 | Cust 1201 | NAGPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Smartpro-con) | Commissioning of conductivity Transmitter. | Installation checked & commissioned successfully conductivity transmitter. | closed | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG23112413117 | 2023-10-30 | 2023-10-30 | Cust 1183 | NAGPUR | Analytical - CHAKAN | CLOSED | APPLICATION ISSUE | MULTIPARAMETER ANALYSER | FM make RS485 card(1) | breakdown | closed | closed | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFNG23113133118 | 2023-10-31 | 2023-11-03 | Cust 1186 | NAGPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Commissioning of FM BB @ ETP,STP and AWR unit | Commissioning of FM BB done | Commissioning of FM BB done | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG23113133118 | 2023-10-31 | 2023-11-03 | Cust 1186 | NAGPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Commissioning of FM BB @ ETP,STP and AWR unit | Commissioning of FM BB done | Commissioning of FM BB done | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG23113133118 | 2023-10-31 | 2023-11-03 | Cust 1186 | NAGPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Commissioning of FM BB @ ETP,STP and AWR unit | Commissioning of FM BB done | Commissioning of FM BB done | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG23113393119 | 2023-10-31 | 2023-10-31 | Cust 1202 | NAGPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | Calibration of DO analyzer on telesupport. | Supported customer on telecall and performed the calibration job. | Supported customer on telecall and performed the calibration job. | NO | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFNG23115153120 | 2023-11-02 | 2023-11-02 | Cust 1194 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | data unavailable on MPCB site | Took access to the system and data from ETP Outlet is now live on the Local + MPCB server. | Took access to the system and data from ETP Outlet is now live on the Local + MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG23115183121 | 2023-11-02 | 2023-11-03 | Cust 1203 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | CPCB data uploaded issue. | now the data is live on Both servers. | now the data is live on Both servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG23116653122 | 2023-11-06 | 2023-11-10 | Cust 1184 | NAGPUR | Analytical - CHAKAN | CLOSED | AMC | Aqua4Trans | Aqua4Trans(1) | Maintenance and calibration of ETP analyzers. | Maintenance and calibration of ETP analyzers is completed. | closed | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFNG23116653122 | 2023-11-06 | 2023-11-10 | Cust 1184 | NAGPUR | Analytical - CHAKAN | CLOSED | AMC | Aqua4Trans | Aqua4Trans(1) | Maintenance and calibration of ETP analyzers. | Maintenance and calibration of ETP analyzers is completed. | closed | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFNG23118753123 | 2023-11-10 | 2023-11-11 | Cust 1203 | NAGPUR | Analytical - CHAKAN | CLOSED | APPLICATION ISSUE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | troubleshooting and Calibration of multiparameter analyzer. | Analyzer calibration completed by providing telephonic support. | closed by telephonic support. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG23118903124 | 2023-11-11 | 2023-11-11 | Cust 1187 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB Connectivity. | We have completed the CPCB connectivity for 75 MLD unit and data is live on the server. | We have completed the CPCB connectivity for 75 MLD unit and data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG23120543125 | 2023-11-16 | 2023-11-18 | Cust 1200 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data uploded issue. | We took access to check this issue and observed that the USB serial com port got changed in the system and now we have made the same changes in the configuration and the data upload is live on the server. | Issue has been resolved and data is live. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG23120813126 | 2023-11-17 | 2023-11-18 | Cust 1204 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation + CPCB + Display board configuration needs to be done. | We did software reinstallation + CPCB + Display board configuration and now data is live on CPCB server. | We did software reinstallation + CPCB + Display board configuration and now data is live on CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG23122263127 | 2023-11-20 | 2023-11-23 | Cust 1007 | NAGPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Maintenance & calibration job. | Maintenance activities carried out for Multiparameter Analyzer & Opacity Meter successfully. Taken Analyzer into service. Data not uploaded to the server. | Calibration done at site. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFNG23122783128 | 2023-11-21 | 2023-11-23 | Cust 1007 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB Connectivity issue. | Now the issue has been resolved and data is live on the CPCB server. | Now the issue has been resolved and data is live on the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG23124253129 | 2023-11-24 | 2023-11-25 | Cust 528 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on MPPCB Server. | It was customer side issue resolved and data is live now on CPCB+MPPCB. | Data was not updating it was customer side issue now data is live on CPCB+MPPCB. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFNG23126373130 | 2023-11-27 | 2024-04-29 | Cust 1187 | NAGPUR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | DO sensor not working; visited site on 27 Nov and 26 Dec with new sensors. sensors not responding to process, abnormal behavior. | New DO sensor and transmitter installed. | closed | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG23127983131 | 2023-11-29 | 2023-11-29 | Cust 1202 | NAGPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | CALIBRATION | CLOSED | CLOSED | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFNG23131443132 | 2023-12-04 | 2023-12-07 | Cust 708 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data uploded issue. | The Data is now live on the server. | The Data is now live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG23135063133 | 2023-12-13 | 2023-12-15 | Cust 708 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue. | As of now we are closing this complaint, and will reopen it after NTP server port 123 is allowed from customers end. | As of now we are closing this complaint, and will reopen it after NTP server port 123 is allowed from customers end. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG23135333134 | 2023-12-13 | 2023-12-15 | Cust 1186 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Black Box configuation needs to be done. | we have completed the Black Box Configuration and data is live on the server for ETP_1 device. | we have completed the Black Box Configuration and data is live on the server for ETP_1 device. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG23137283135 | 2023-12-17 | 2023-12-21 | Cust 1184 | NAGPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4Trans(1) | Commissioning of Aqua 4 transe pH & Conductivity transmitters. | Commissioning of Aqua 4 transe pH & Conductivity transmitters is completed successfully. | closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG23137283135 | 2023-12-17 | 2023-12-21 | Cust 1184 | NAGPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4Trans(1) | Commissioning of Aqua 4 transe pH & Conductivity transmitters. | Commissioning of Aqua 4 transe pH & Conductivity transmitters is completed successfully. | closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG23137283135 | 2023-12-17 | 2023-12-21 | Cust 1184 | NAGPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4Trans(1) | Commissioning of Aqua 4 transe pH & Conductivity transmitters. | Commissioning of Aqua 4 transe pH & Conductivity transmitters is completed successfully. | closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG23137283135 | 2023-12-17 | 2023-12-21 | Cust 1184 | NAGPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4Trans(1) | Commissioning of Aqua 4 transe pH & Conductivity transmitters. | Commissioning of Aqua 4 transe pH & Conductivity transmitters is completed successfully. | closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG23137283135 | 2023-12-17 | 2023-12-21 | Cust 1184 | NAGPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4Trans(1) | Commissioning of Aqua 4 transe pH & Conductivity transmitters. | Commissioning of Aqua 4 transe pH & Conductivity transmitters is completed successfully. | closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG23137413136 | 2023-12-17 | 2023-12-18 | Cust 1186 | NAGPUR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | FM BB failed @ ETP analyzer | visited the site on 13.12.23. System given online configuration of FM BB done. ETP analyzer | System given online configuration of FM BB | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG23140663137 | 2023-12-23 | 2023-12-25 | Cust 1205 | NAGPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Spare received at site. Call for the commissioning of spares of multiparameter Analyzer. | Received spares used & taken ETP-3into service. pH measurement not become online as the pH input card not received at site. | closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG23140673138 | 2023-12-23 | 2023-12-25 | Cust 1187 | NAGPUR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Repaired Black Box failed at site. Call for the commissioning of same. | Received RS 485isolator checked & found not working. The same need to replace to update the data into server. | Closed. | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG23140783139 | 2023-12-23 | 2023-12-26 | Cust 1206 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | BB CONFIGURATION NEEDS TO BE DONE. | We have completed the CPCB connectivity through FMBB and now the data is live on the server | We have completed the CPCB connectivity through FMBB and now the data is live on the server | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG23141883140 | 2023-12-25 | 2023-12-25 | Cust 1187 | NAGPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | COD & TSS reading not showing. always showing --- on the display. | Necessary parameter settings done. Calibration done in DI water & taken Analyzer into service. Online working performance observed & found ok. | closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG23142243141 | 2023-12-26 | 2024-01-11 | Cust 1186 | NAGPUR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | SWAS & Analysers | Steam & Water Analysis System(SWAS & Analysers) | Following spares and consumables are required. PH sensor , Buffer amaplifier and PH input card Aqua4 trans 01 no. each | Material received from division. Visit is plan for replacement | Material received from division. Visit is plan for replacement | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFNG23143783142 | 2023-12-28 | 2024-01-29 | Cust 708 | NAGPUR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Electrode | FM1104(1) | mV output is very low | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210077512 LR Dt - 03-JAN-2024 | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210077512 LR Dt - 03-JAN-2024 | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFNG23147343143 | 2024-01-02 | 2024-01-30 | Cust 1202 | NAGPUR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | AQUADO Analyzer | OXYSENS(OXYSENS) | Oxygen sensor is not working in process. | RF done for required spares. | RF done for required spares. | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFNG23151803144 | 2024-01-10 | 2024-01-16 | Cust 1207 | NAGPUR | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | TOC ANALYSER | FM 101 TOC Analyser(FM 101 TOC) | Erection activity for TOC Project to be started. | Issue of location change resolved by discussion on location. Erection / Installation activities started | Lined up Erection activities | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG23154723145 | 2024-01-16 | 2024-01-18 | Cust 1203 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | stack and ETP multiparameter reading not showing on MPCB server. | Data is now live on the MPCB and CPCB server. | Data is now live on the MPCB and CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG23155743146 | 2024-01-18 | 2024-01-23 | Cust 1205 | NAGPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | pH reading on higher side. | Calibration checked through standard pH simulator & found Analyzer working ok. pH Electrode kept overnight in 3Mole KCL solution for regeneration. Customer taken it into service & found working ok. | Calibration done at site. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG23158553147 | 2024-01-24 | 2024-01-26 | Cust 1190 | NAGPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4Trans(1) | Commissioning call. | Buffer Amplifier replaced, Calibration carried out & taken Analyzer into service. All the commissioned Analyzers are under observation. | closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG23159083148 | 2024-01-25 | 2024-01-25 | Cust 1190 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | DATCON | DATCON(DATCON) | Issue in Decimal value | correct readings are updated on the server. | correct readings are updated on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG23159213149 | 2024-01-25 | 2024-01-26 | Cust 708 | NAGPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Erratic pH reading. New supplied pH electrode not working. | pH electrode replaced at site. Calibration performed & taken Analyzer into service. Online working performance for 4 - 5 Hours observed & found ok. | closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG23161723150 | 2024-01-30 | 2024-03-25 | Cust 1184 | NAGPUR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Electrode | FM1104(1) | 03 nos of pH electrodes were found faulty during commissioning. | Replacement material has been booked through TCI EXPRESS vide LR No - 110007585004 LR Dt - 19-MAR-2024 | Replacement material has been booked through TCI EXPRESS vide LR No - 110007585004 LR Dt - 19-MAR-2024 | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFNG23161723150 | 2024-01-30 | 2024-03-25 | Cust 1184 | NAGPUR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Electrode | FM1104(1) | 03 nos of pH electrodes were found faulty during commissioning. | Replacement material has been booked through TCI EXPRESS vide LR No - 110007585004 LR Dt - 19-MAR-2024 | Replacement material has been booked through TCI EXPRESS vide LR No - 110007585004 LR Dt - 19-MAR-2024 | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFNG23161723150 | 2024-01-30 | 2024-03-25 | Cust 1184 | NAGPUR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Electrode | FM1104(1) | 03 nos of pH electrodes were found faulty during commissioning. | Replacement material has been booked through TCI EXPRESS vide LR No - 110007585004 LR Dt - 19-MAR-2024 | Replacement material has been booked through TCI EXPRESS vide LR No - 110007585004 LR Dt - 19-MAR-2024 | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFNG23162533151 | 2024-01-31 | 2024-02-05 | Cust 1205 | NAGPUR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Spares purchased . call for the commissioning of same. | Spares replaced, calibration done & taken Multi Parameter Analyzer into service. | closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG23164703152 | 2024-02-02 | 2024-02-05 | Cust 1208 | NAGPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Maintenance and calibration of ETP analyzer | Maintenance and calibration of repaired ETP analyzer is carried out. | closed | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFNG23165093153 | 2024-02-03 | 2024-02-06 | Cust 1205 | NAGPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Spares commissioning & its Calibration. | Spares replaced, Calibration done & taken Multiparameter into service. Online working performance observed & found ok. | Closed. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG23165093153 | 2024-02-03 | 2024-02-06 | Cust 1205 | NAGPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Spares commissioning & its Calibration. | Spares replaced, Calibration done & taken Multiparameter into service. Online working performance observed & found ok. | Closed. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG23165223154 | 2024-02-03 | 2024-02-09 | Cust 1209 | NAGPUR | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | TOC ANALYSER | FM 101 TOC Analyser(FM 101 TOC) | Completion of erection of TOC Shelter | Completion of erection of TOC Shelter almost done. Only sample tubing is pending | Lined up Erection activities | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG23166863155 | 2024-02-07 | 2024-02-13 | Cust 1187 | NAGPUR | Analytical - CHAKAN | CLOSED | APPLICATION ISSUE | AQUADO Analyzer | Dissolved Oxygen Analyzer Wall mounted(AQUADO4000W) | Installation of Aquamax DO on areation tank 2. Replacement of Oxysense DO sensor. | Installation of Visiwater DO is completed and working ok. | closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG23167623156 | 2024-02-08 | 2024-02-09 | Cust 1187 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Data Mismatch issue on CPCB server. | The issue has been resolved and now the correct data is uploading to the server | The issue has been resolved and now the correct data is uploading to the server | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG23167723157 | 2024-02-08 | 2024-02-19 | Cust 1190 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera configuration needs to be done. | Now the camera is streaming live on the CPCB server for Camera at Guard Pond Outlet. | Now the camera is streaming live on the CPCB server for Camera at Guard Pond Outlet. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG23167983158 | 2024-02-09 | 2024-03-21 | Cust 1205 | NAGPUR | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | Ph Electrode | Ph Electrode(FM1104(4C)) | Premature failures of pH Electrodes | pH sensor despatched from division to site. Visit required to commission the same. | pH sensor despatched from division to site. Visit required to commission the same. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFNG23170263159 | 2024-02-14 | 2024-02-16 | Cust 497 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data uploading issue to MPCB server | We have updated a DAS and now data is live on the MPCB server. | We have updated a DAS and now data is live on the MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG23171593160 | 2024-02-17 | 2024-02-20 | Cust 1055 | NAGPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | DO2 Arc Sensors(DO2 Arc Sensors) | Commissioning of DO analyzer. | Commissioning of DO analyzer was completed | Commissioning of DO analyzer was completed. | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFNG23171603161 | 2024-02-17 | 2024-02-20 | Cust 1202 | NAGPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | DO2 Arc Sensors(DO2 Arc Sensors) | Commissioning of new DO analyzer | Analyzer found not working at site. Sent to division for analysis | Analyzer found not working at site. Same was sent to division for analysis. | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFNG23172993162 | 2024-02-20 | 2024-02-26 | Cust 1184 | NAGPUR | Analytical - CHAKAN | CLOSED | AMC | Aqua4Trans | Turbidity Input card of Aqua4Trans(1) | Maintenance and calibration of pH, turbidity and Dust analyzers | Maintenance and calibration of pH, turbidity and Dust analyzers is completed successfully. | closed | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFNG23172993162 | 2024-02-20 | 2024-02-26 | Cust 1184 | NAGPUR | Analytical - CHAKAN | CLOSED | AMC | Aqua4Trans | Turbidity Input card of Aqua4Trans(1) | Maintenance and calibration of pH, turbidity and Dust analyzers | Maintenance and calibration of pH, turbidity and Dust analyzers is completed successfully. | closed | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFNG23173923163 | 2024-02-21 | 2024-02-26 | Cust 1209 | NAGPUR | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | TOC ANALYSER | TOC 1950 Single StreamANALYSER ANATEL(1950 TOC) | completion of erection | Completed Erection | Completed Erection activities | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG23174613164 | 2024-02-22 | 2024-02-26 | Cust 1186 | NAGPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | No communication Through Modbus. | Site readiness completed at Site & data available on LAN to upload it on Server. | closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG23181023165 | 2024-03-01 | 2024-03-02 | Cust 1210 | NAGPUR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | visit required to check the reported problem in Ph meters installed at PCTP plant. | Reported issue checked and attended successfully. Online readings were cross checked with offline found ok. | visited the site reported problem attended | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG23181023165 | 2024-03-01 | 2024-03-02 | Cust 1210 | NAGPUR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | visit required to check the reported problem in Ph meters installed at PCTP plant. | problem in transmitter suspected. After replacing pH sensor/ cable problem remains same meter showing zero reading. Suggested to send the meter to pune for checking | visited the site reported problem attended | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG23181593166 | 2024-03-04 | 2024-03-08 | Cust 1209 | NAGPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | TOC ANALYSER | FM 101 TOC Analyser(FM 101 TOC) | Powering UP of the Analyser & Connectivity Hardware | Powering UP of the Analyser & Connectivity Hardware completed successfully | Erection completed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG23185553167 | 2024-03-14 | 2024-03-18 | Cust 1209 | NAGPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | TOC ANALYSER | Crossflow Filter Body(001?255) | visit required for commissioning of TOC ananlyzer installed at by product line | Visited the site from 13 to 15 March . following activites were carried out. 1) Analyzer Power On. reagents were prepared . 2) Analyzer taken into service with DM water. 3) Analog output hookup to DCS. 4) Data availability checked upto FMBB found ok. 5) Required connectivity document handed over to NMDC. | Analyzer made ready for final commissioning | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG23188983168 | 2024-03-19 | 2024-03-21 | Cust 1186 | NAGPUR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | PH Input Card for Multiparameter(AWA-PH) | pH reading is not correct | pH card found faulty. Analyser re-configuration done , all other measurements made healthy. | attended and closed | NO | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG23188993169 | 2024-03-19 | 2024-03-23 | Cust 1186 | NAGPUR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Data not available on site | Data was available at Analyser O/p. Problem identified in Optical fibre. Attended by customer by replacing optical fibre core.Now data is available on server | attended and closed | YES | Periodic Maintenance (AMC) | Core | N | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG23191303170 | 2024-03-22 | 2024-03-23 | Cust 1186 | NAGPUR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | High pH reading & connectivity issue. | High pH reading & Aqua4Trans issue successfully resolved at site with help of R&D Engineers. | Calibration done at site. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG23191373171 | 2024-03-22 | 2024-03-23 | Cust 1186 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Local Host connectivity. | Local host connectivity established successfully with help of Support Team. | closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG23191433172 | 2024-03-22 | 2024-03-22 | Cust 1186 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software installation and Local Connectivity needs to be done | We have completed the software installation and now the data is Live on the Local server. | We have completed the software installation and now the data is Live on the Local server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG23197973173 | 2024-03-31 | 2024-03-31 | Cust 1211 | NAGPUR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | calibration failed, Need to send it to pune | closed | closed | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFNG24003223174 | 2024-04-08 | 2024-04-30 | Cust 1055 | NAGPUR | Analytical - CHAKAN | CLOSED | SALES PROMOTION | Ph Electrode | FM1104(1) | pH electrode not working | REPLACED | REPLACED | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFNG24003503175 | 2024-04-08 | 2024-04-15 | Cust 1190 | NAGPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning call. | Multiparameter Analyzer installation done successfully. Power on & taken into successfully. There is high pH reading observed & the same was rectified successfully. Spftware upgradation part pending. the same istaken with R&D & will resolve after firmware upgradation. | attended and closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24005503176 | 2024-04-12 | 2024-04-15 | Cust 497 | NAGPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 1st AMC visit | Maintenance activities carried out at site & taken Analyzer into service. Online working performance observed & found ok. | Calibration done at site. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFNG24006473177 | 2024-04-15 | 2024-04-20 | Cust 1212 | NAGPUR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | DO Transmitters | VISI pro DO(VISI pro DO) | Sensors need to check as all are switched off condition. | Visited site , checked all 3 sensors & found not getting powered on. Need replacement. Customer will send all these sensor to our factory for repair is any possible on chargeable basis. | closed | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24007113178 | 2024-04-17 | 2024-04-19 | Cust 1194 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB and MPCB data uploaded issue. | we took access and now the data is live on both sites. | we took access and now the data is live on both sites. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG24007303179 | 2024-04-17 | 2024-04-22 | Cust 1189 | NAGPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Maintenance and calibration of Multiparameter analyzer installed at ETP area. | Maintenance and calibration of Multiparameter analyzer installed at ETP area is completed. | closed | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFNG24007473180 | 2024-04-17 | 2024-04-18 | Cust 1186 | NAGPUR | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | Turbidity Meter | Thermo series(Thermo series) | visit required at Koradi 660 MW to give guideline for erection and precommissioning for thermo make turbidity ananlyzer @ETP plant qty 03 nos. and Krohne make level transmitters qty 03 nos. | Visited the site installation and erection guidelines given. Location finalised for the mounting of turbidity analyzers. Visit jointly done with MSPGCL persons. | erection guideline given. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24007473180 | 2024-04-17 | 2024-04-18 | Cust 1186 | NAGPUR | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | Turbidity Meter | Thermo series(Thermo series) | visit required at Koradi 660 MW to give guideline for erection and precommissioning for thermo make turbidity ananlyzer @ETP plant qty 03 nos. and Krohne make level transmitters qty 03 nos. | Visited the site installation and erection guidelines given. Location finalised for the mounting of turbidity analyzers. Visit jointly done with MSPGCL persons. | erection guideline given. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24008533181 | 2024-04-20 | 2024-04-22 | Cust 1081 | NAGPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | VISI pro DO(VISI pro DO) | Commissioning call. | Analyzer taken into service but readings are on higher side. FM Engineer will visit again to conclude the site. | commissioning completed at site. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24012023182 | 2024-04-26 | 2024-04-29 | Cust 1213 | NAGPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4Trans(1) | Commissioning of pH and Chlorine analyzers | Commissioning of pH analyzer and Chlorine is pending as the pH cable not found at site and clorine analyzer rotameter damaged. seperate complaint is logged for the replacement. | closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24012023182 | 2024-04-26 | 2024-04-29 | Cust 1213 | NAGPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4Trans(1) | Commissioning of pH and Chlorine analyzers | Commissioning of pH analyzer is completed at intake well. | closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24013193183 | 2024-04-29 | 2024-05-03 | Cust 1209 | NAGPUR | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | TOC ANALYSER | FM 101 TOC Analyser(FM 101 TOC) | Installation of Sample Filters | For Installation of Sample Filters , needed connectors. Purchased these connectors at Nagpur and dispatched to site | Lined up Erection activities | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24013583184 | 2024-04-29 | 2024-05-03 | Cust 1213 | NAGPUR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Chlorine Analyzer-MAC 100 | MAC 100(1) | Chlorine analyser Rotameter damaged. pH sensor cables & flowthrough chambers IN & OUT connector were not found at site. | Separate Complaint is being logged under short supply / item lost at site to enable supply of connectors. | Separate Complaint is being logged under short supply / item lost at site to enable supply of connectors. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFNG24014193185 | 2024-04-30 | 2024-04-30 | Cust 1055 | NAGPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | PH input card(1) | PH | CLOWS | CLOSED | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFNG24014283186 | 2024-04-30 | 2024-05-03 | Cust 1077 | NAGPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | SWAN Make Portable DO Transmitter AMC. | Maintenance done & taken Analyzer into service. Online reading & actual readings are not matched. Spares suggested to make the instrument healthy. | AMC visit done. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFNG24015523187 | 2024-05-01 | 2024-05-03 | Cust 708 | NAGPUR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | pH reading showing on higher side. Replacement received at our Nagpur office. Visit required to commission the same. | New supplied pH input card not compatible to the existing PCX-1 . New RF to be raised for complete set. | Closed. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24016713188 | 2024-05-03 | 2024-05-06 | Cust 1194 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data upload issue. | today we took access to your system and now the issue has been resolved and data is live on the MPCB server. | today we took access to your system and now the issue has been resolved and data is live on the MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFNG24017443189 | 2024-05-06 | 2024-05-14 | Cust 1186 | NAGPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Scheduled Visit for the checking of Multiparameter Analyzer. | Calibration done at site & taken into service. Online working of Analyzer observed & found ok. | Calibration done at site. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFNG24017883190 | 2024-05-06 | 2024-05-09 | Cust 1207 | NAGPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | TOC ANALYSER | FM 101 TOC Analyser(FM 101 TOC) | Commissioning of TOC analyzer and Installation of Sample filter | Visited the site from 6 to 8 Precommissioning activites were carried out . Sample filter erection done. Waiting for EOT to start final commissioning activity | Precommissionig activites were carried out | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24018453191 | 2024-05-08 | 2024-05-10 | Cust 1209 | NAGPUR | Analytical - CHAKAN | CLOSED | DIGITAL CONEECTIVITY | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Digital connectivity needs to be done. | Today we have completed the cloud server connectivity and now the data is live on the server. | Today we have completed the cloud server connectivity and now the data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG24018753192 | 2024-05-08 | 2024-05-12 | Cust 1184 | NAGPUR | Analytical - CHAKAN | CLOSED | AMC | Aqua4Trans | Turbidity Input card of Aqua4Trans(1) | maintenance and calibration of ETP analyzers installed at ATPS chachai. | maintenance and calibration of ETP 1 analyzers installed at ATPS chachai is completed. | closed | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFNG24018753192 | 2024-05-08 | 2024-05-12 | Cust 1184 | NAGPUR | Analytical - CHAKAN | CLOSED | AMC | Aqua4Trans | Turbidity Input card of Aqua4Trans(1) | maintenance and calibration of ETP analyzers installed at ATPS chachai. | maintenance and calibration of ETP 2 analyzers installed at ATPS chachai is completed. | closed | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFNG24020163193 | 2024-05-12 | 2024-05-22 | Cust 1214 | NAGPUR | Steam System Division | CLOSED | QUALITY PROBLEM | PRV - MEASURETROL | FM MEASURETROL(SIZE 050NB) | Prv not showing outlet pressure in app. And unable to control. Qty -2nos | Mr Jitesh Patil and Mr. Sagar visited the site and has upgraded the 07Nos Electronic PCB's with Zener Diode as a proactive measure. | Mr Jitesh Patil and Mr. Sagar visited the site and has upgraded the 07Nos Electronic PCB's with Zener Diode as a proactive measure. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFNG24020163193 | 2024-05-12 | 2024-05-22 | Cust 1214 | NAGPUR | Steam System Division | CLOSED | QUALITY PROBLEM | PRV - MEASURETROL | FM MEASURETROL(SIZE 080NB) | Prv not showing outlet pressure in app. And unable to control. Qty -2nos | Mr Jitesh Patil and Mr. Sagar visited the site and has upgraded the 07Nos Electronic PCB's with Zener Diode as a proactive measure. | Mr Jitesh Patil and Mr. Sagar visited the site and has upgraded the 07Nos Electronic PCB's with Zener Diode as a proactive measure. | YES | No Trip | Process | N | Product / equipment failure due to which customers plant or part of plant is stopped which leaads to production loss. | Analytical - CHAKAN
AFNG24020783194 | 2024-05-14 | 2024-05-17 | Cust 1215 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB connectivity needs to be done. | we took access to your system and completed CPCB connectivity for Stack 2 CEM, and now data is live on CPCB server. | we took access to your system and completed CPCB connectivity for Stack 2 CEM, and now data is live on CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG24021343195 | 2024-05-15 | 2024-05-17 | Cust 1187 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data uploaded issue. | We took access to your system and updated the "EnviroconnectDAS" and now we observed that ETP OUTLET data is not coming to the local system ,and checked on modscan. | We took access to your system and updated the &amp;quot;EnviroconnectDAS&amp;quot; and now we observed that ETP OUTLET data is not coming to the local system ,and checked on modscan. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG24023393196 | 2024-05-20 | 2024-05-22 | Cust 708 | NAGPUR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | pH reading showing always higher. M/s FMPL supplied spares like DIsplay Board ( PCX-1), pH input card & pH electrode. | Spares replaced at site & taken Analyzer into service. Online working performance observed & found ok. | closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24023903197 | 2024-05-21 | 2024-05-22 | Cust 1187 | NAGPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Display keypad not working. Analyzer showing continuous zero reading. | Reported isuue of TSS reading in inlet & outlet Multiparameter Analyzer checked & resolves successfully. | Calibration done at site. | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24025473198 | 2024-05-24 | 2024-05-26 | Cust 1197 | NAGPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC Visit. | 2nd AMC visit done at site & taken Analyzer into service. Online connectivity checked & found ok. | AMC visit done. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFNG24026833199 | 2024-05-27 | 2024-05-27 | Cust 1197 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue. | As discussed, today we took access to the system and now the hardware side issue is resolved and data is live on CPCB server for ETP. | As discussed, today we took access to the system and now the hardware side issue is resolved and data is live on CPCB server for ETP. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG24027923200 | 2024-05-28 | 2024-05-29 | Cust 1186 | NAGPUR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | 24 VDC SMPS (1) | Call from user that MPA not power ON . Online support required for troubleshooting | Online with user problem in SMPS observed. New SMPS arrange by the Kamal Agencies . SMPS installed and Analyzer taken into service | Online with user problem in SMPS observed. New SMPS arrange by the Kamal Agencies . SMPS installed and Analyzer taken into service | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFNG24028583201 | 2024-05-28 | 2024-06-03 | Cust 1207 | NAGPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | TOC ANALYSER | FM 101 TOC Analyser(FM 101 TOC) | Commissioning of TOC ananlyzer @NMDC plant .Irration Basin | Visited the site from 28 to 31 May. Commissioing of TOC analyzer done as per PO. Analyzer online with actual sample. | TOC Analyzer taken into service | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24028693202 | 2024-05-29 | 2024-06-01 | Cust 1187 | NAGPUR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Display went blank. Lining issue. | Reported Display issue has been checked & resolved successfully. Presently Analyzer is working ok. | Calibration done at site. | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24032473203 | 2024-06-04 | 2024-06-07 | Cust 528 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue. | As of now we are closing this complaint and will reopen it once the hardware side issue is resolved. | As of now we are closing this complaint and will reopen it once the hardware side issue is resolved. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG24033293204 | 2024-06-05 | 2024-06-07 | Cust 1205 | NAGPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | PH calibration issue. | new pH input card installed, Still problem persist. Issue observed in the PCX 1 card which needs software upgradation. | closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24036573205 | 2024-06-13 | 2024-06-14 | Cust 1216 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPP CEMS and Effluent data are not going to CPCB RTDMS | Data is live on RTDMS CPCB portal. | Data is live on RTDMS CPCB portal. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFNG24036703206 | 2024-06-13 | 2024-08-08 | Cust 708 | NAGPUR | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Almost all the spare expect PCX - 2 replaced at site. Still customer reported ananlyzer not working. Now reported issue of display flickering. | New PCX-1 Board received at Nagpur office. | Closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24036853207 | 2024-06-13 | 2024-06-13 | Cust 1185 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB connectivity needs to be done. | we took access of ZLD_Outlet/Plant_Outlet and now the internet side issue has been resolved and data is live on CPCB server for all parameters (pH,TSS,Temperature,Flow). | we took access of ZLD_Outlet/Plant_Outlet and now the internet side issue has been resolved and data is live on CPCB server for all parameters (pH,TSS,Temperature,Flow). | YES | No Trip | Core | Y | Product / equipment failure impacting safety of equipment, process or people. | Analytical - CHAKAN
AFNG24036983208 | 2024-06-14 | 2024-06-15 | Cust 1185 | NAGPUR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | THERMO SERIES | ADAPTER(1) | Visit required to check the thermo scientific analyzers installed at DM plant & SWAS room | Visited the site on 12 &13 June analyzers installed at DM plant & SWAS room checked. Calibration of same carried out. | Analyzer installed at DM plant and SWAS room checked | NO | Paid | Core | N |  | Analytical - CHAKAN
AFNG24037203209 | 2024-06-14 | 2024-06-14 | Cust 1194 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on CPCB + Local + MPCB server. | Data is live on CPCB + Local + MPCB server. | Data is live on CPCB + Local + MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFNG24038183210 | 2024-06-17 | 2024-06-18 | Cust 1185 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data upload issue. | we took access to your system and updated the new Black Box patch in Plant_Outlet and now data is live on the CPCB server. | we took access to your system and updated the new Black Box patch in Plant_Outlet and now data is live on the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG24039363211 | 2024-06-18 | 2024-06-20 | Cust 1194 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Enviroconnect software latest version to update | Updated the latest version of the EnviroConnect software. And now data is live on CPCB + Local + MPCB server. | Updated the latest version of the EnviroConnect software. And now data is live on CPCB + Local + MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFNG24041603212 | 2024-06-24 | 2024-06-25 | Cust 528 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB + MPPCB data upload issue. | the issue has been resolved and flow data is uploading on the CPCB + MPPCB server, | the issue has been resolved and flow data is uploading on the CPCB + MPPCB server, | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG24041963213 | 2024-06-24 | 2024-06-27 | Cust 1186 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB connectivity needs to be done. | we have completed the MPCB connectivity and now the data is live on the server for ETP units. | we have completed the MPCB connectivity and now the data is live on the server for ETP units. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG24043053214 | 2024-06-26 | 2024-06-29 | Cust 497 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB ISSUE. | the issue has been resolved from the hardware side and now the data started updating on the server. | the issue has been resolved from the hardware side and now the data started updating on the server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFNG24043693215 | 2024-06-27 | 2024-06-29 | Cust 1217 | NAGPUR | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | pre commissioning visit for installation guidelines. | Panel installation, Sample tubing, power supply requirnments & internet connection requirnments details explained. M/s VEPL will inform for final commissioning after completion of erection activities. | closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24044963216 | 2024-06-29 | 2024-07-02 | Cust 1218 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | AFBC BOLIER 2 ( DEVICE ) SHOWING HIGH VALUE ON DIGITAL SCANNER. | we have checked AFBC-2 online PM data on modscan and data is matching with the CPCB server. | we have checked AFBC-2 online PM data on modscan and data is matching with the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFNG24047523217 | 2024-07-02 | 2024-07-09 | Cust 1186 | NAGPUR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Analyser getting resetted frequently | Analyser hard reseting done. | attended and closed | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24047963218 | 2024-07-03 | 2024-07-05 | Cust 1185 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data upload issue. | we took access to your system and updated the new Black Box patch in Plant_Outlet and now data is live on the CPCB server. | we took access to your system and updated the new Black Box patch in Plant_Outlet and now data is live on the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG24048193219 | 2024-07-03 | 2024-07-05 | Cust 1186 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Blackbox commissioning. | We have completed the FM Black Box Configuration + MPCB connectivity and now the data is live on the MPCB server. | We have completed the FM Black Box Configuration + MPCB connectivity and now the data is live on the MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFNG24048423220 | 2024-07-04 | 2024-07-06 | Cust 1186 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | MPCB connectivity of ETP-4 Analyzer. | MPCB connectivity established successfully. | Calibration done at site. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24048433221 | 2024-07-04 | 2024-07-06 | Cust 1205 | NAGPUR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | pH input card-FM of Multiparameter(1) | inaccurate pH measurement. | Software upgradation done, pH input card replaced & taken analyzer into service. | Calibration done at site. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24048593222 | 2024-07-04 | 2024-07-06 | Cust 1219 | NAGPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Erratic measurement of Multiparameter analyzer. | Analyzer reset, Necessary parameter settings done & taken Analyzer into service. Spares required to make the pH measurement online. Requested customer to upgrade the instrument with CX2000 series Analyzer. | closed | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24049133223 | 2024-07-05 | 2024-07-06 | Cust 1220 | NAGPUR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | pH measurement not correct. | Software upgradation done, pH input card replaced & taken anlyzer into service. Online working of Aanlyzer observed & found ok. | Calibration done at site. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24050443224 | 2024-07-08 | 2024-07-11 | Cust 1187 | NAGPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | VISI pro DO(VISI pro DO) | Visit required to install ne DO sensor visiwater with DPI 201 | visited the site on 8.7.24. New sensor supplied by division installed . Perfromnace checked found ok. | new visiwater sensor installed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24050983225 | 2024-07-09 | 2024-07-12 | Cust 1186 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data breaking issue. | We took access of the BB and updated with the new patch hence now the data is live. | We took access of the BB and updated with the new patch hence now the data is live. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG24051023226 | 2024-07-09 | 2024-07-12 | Cust 1187 | NAGPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | VISI pro DO(VISI pro DO) | New sensor received at office , Call for the commissioning of same. | Demo unit installation complete at Site & taken Analyzer into service. Online working observed & found ok. | commissioning completed at site. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24051043227 | 2024-07-09 | 2024-07-12 | Cust 1186 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Spares installation & its connectivity to the CPCB server. | Spares installation done & taken analyzer into service. Observed that Display hanged & needs replacement. Data connectivity break intermittently. | closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24052543228 | 2024-07-12 | 2024-07-15 | Cust 528 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on CPCB server. | At site we have installed Enviroconnect software. hence facing frequent issues in data upload to CPCB + MPPCB server .As suggested by our software team also, the MPPCB driver has windows compatibility issues and on almost all sites we have replaced the software with FM black box so For your site We have asked customer to upgrade the software with FM black box and as per the site engineers confirmation the black box is already supplied at site .once we get any revert from the customer we will reopen this case again. | We have asked customer to upgrade the software with FM black box and as per the site engineers confirmation the black box is already supplied at site from the customer we will reopen this case again. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFNG24052923229 | 2024-07-12 | 2024-07-15 | Cust 497 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | MPCB data breaking issue. | We took access of the system and made some modifications Now the data is uploading continuously and the data availability is almost 99% on the MPCB server. | We took access of the system and made some modifications Now the data is uploading continuously and the data availability is almost 99% on the MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG24053133230 | 2024-07-13 | 2024-07-15 | Cust 1217 | NAGPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning of Multiparameter analyzer CX2000 | Commissioning is completed successfully. | closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24053753231 | 2024-07-15 | 2024-07-26 | Cust 1217 | NAGPUR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | HMI hang intermittently. | Replacement received and installed. Analyzer is working ok after HMI replacement. | closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24054043232 | 2024-07-15 | 2024-08-12 | Cust 1220 | NAGPUR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Isolator found Not working. | The material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210107960 LR Dt - 09-AUG-2024. | Closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24054113233 | 2024-07-15 | 2024-07-18 | Cust 1183 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on LOCAL+CPCB server. | We took access to the system and found there is converter issue. | We took access to the system and found there is converter issue. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFNG24055073234 | 2024-07-17 | 2024-07-20 | Cust 1218 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Re-installation of Enviroconnect software . | We took access to the system and successfully completed the software re-installation. Now the data is currently live on the CPCB server for the following stacks: Stack_2_KILN-2_WHRB, Stack_3_combined_KILN-3&4_WHRB, Stack_6_Ferro_1&2, Stack_7_Ferro_3&4, and Stack_8_Ferro_1. However, we have noticed that the data is not communicating with the system for Stack_1_KILN-1_WHRB, Stack_4_AFBC_Boiler-1, and Stack_5_AFBC_Boiler-2 (checked on MODSCAN). So the issue is from the hardware side. so we are closing this complaint. | we have noticed that the data is not communicating with the system for Stack_1_KILN-1_WHRB, Stack_4_AFBC_Boiler-1, and Stack_5_AFBC_Boiler-2 (checked on MODSCAN) | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFNG24055823235 | 2024-07-18 | 2024-07-22 | Cust 497 | NAGPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC visit | Maintenance done & taken Analyzer into service. Online working performance observed & found ok. | Calibration done at site. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFNG24057993236 | 2024-07-23 | 2024-08-09 | Cust 1185 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data uploaded issue. | we have completed all configurations for both the Display Board and files are generating from our end on your both PC's and the same is discussed with you also. | we have completed all configurations for both the Display Board and files are generating from our end on your both PC's and the same is discussed with you also. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG24058073237 | 2024-07-23 | 2024-07-23 | Cust 1218 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | local server data upload issue. | We took access to the system and successfully completed the software re-installation. Now the data is currently live on the CPCB + Local servers. | We took access to the system and successfully completed the software re-installation. Now the data is currently live on the CPCB + Local servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFNG24058153238 | 2024-07-23 | 2024-07-26 | Cust 1221 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading in CPCB / SPCB Server . | now the data is live on CPCB + Local servers. | now the data is live on CPCB + Local servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFNG24058173239 | 2024-07-23 | 2024-07-26 | Cust 1183 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Stack data not updating on CPCB+LOCAL Server. | We took access to the system and data is live on CPCB server. There was a converter side issue from customer end. | We took access to the system and data is live on CPCB server. There was a converter side issue from customer end. | YES | No Trip | Process | Y |  | Analytical - CHAKAN
AFNG24058353240 | 2024-07-23 | 2024-09-27 | Cust 1187 | NAGPUR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | AQUADO Analyzer | Aquamax system(1) | DO sensor membrane and DPI found faulty. | Replacement received and installed at site on 25th Sept 24. system is under observation | closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24060153241 | 2024-07-26 | 2024-07-31 | Cust 1214 | NAGPUR | Steam System Division | CLOSED | QUALITY PROBLEM | PRV - MEASURETROL | FM MEASURETROL(SIZE 040NB) | PRS not working due to multiple error. | Replacement has bene arranged against RF OA Number 5525103430 | Replacement has bene arranged against RF OA Number 5525103430 | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFNG24060153241 | 2024-07-26 | 2024-07-31 | Cust 1214 | NAGPUR | Steam System Division | CLOSED | QUALITY PROBLEM | PRV - MEASURETROL | FM MEASURETROL(SIZE 040NB) | PRS not working due to multiple error. | Replacement has bene arranged against RF OA Number 5525103430 | Replacement has bene arranged against RF OA Number 5525103430 | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFNG24063133242 | 2024-07-30 | 2024-08-09 | Cust 1221 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | EnviroConnect Software re-installation needs to be done. | we have completed all configurations for both the Display Board and files are generating from our end on your both PC's and the same is discussed with you also. | we have completed all configurations for both the Display Board and files are generating from our end on your both PC's and the same is discussed with you also. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFNG24064483243 | 2024-07-31 | 2024-08-02 | Cust 497 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB server. | We took access to the system and now data is updating on CPCB Server. | We took access to the system and now data is updating on CPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFNG24065153244 | 2024-08-01 | 2024-08-02 | Cust 1183 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB Server. | There was converter side issue from customer side. Converter has been replaced at customer end and we have completed the software side changes and now data is live on LOCAL+CPCB Server. | There was converter side issue from customer side. Converter has been replaced at customer end and we have completed the software side changes and now data is live on LOCAL+CPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFNG24068313245 | 2024-08-08 | 2024-08-09 | Cust 1186 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB AND MPCB DATA UPLODED ISSUE. | Now the Data is live on the CPCB and MPCB server. | Now the Data is live on the CPCB and MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG24068843246 | 2024-08-09 | 2024-09-23 | Cust 1222 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Data upload issue on CPCB + MPCB server | We accessed the router, & updated it to the latest version of Nexapp Router, and re-configured it. And now the data is live on the CPCB and MPCB servers. | We accessed the router, &amp;amp; updated it to the latest version of Nexapp Router, and re-configured it. And now the data is live on the CPCB and MPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFNG24071803247 | 2024-08-16 | 2024-08-19 | Cust 1187 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data uploded issue. | As of now we are closing this complaint and we will reopen this once the Hardware side issue is resolved for OUTLET from your end. | As of now we are closing this complaint and we will reopen this once the Hardware side issue is resolved for OUTLET from your end. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG24071843248 | 2024-08-16 | 2024-08-17 | Cust 528 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB + MPPCB data upload issue. | we again took access of the black box and now data is live on CPCB + MPPCB server. | we again took access of the black box and now data is live on CPCB + MPPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG24071923249 | 2024-08-16 | 2024-08-16 | Cust 1186 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB and MPCB data uploded issue. | Now the Data is live on the server. | Now the Data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG24073493250 | 2024-08-20 | 2024-08-22 | Cust 1192 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to Central portal | data is showing live on the Central portal. | data is showing live on the Central portal. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFNG24073743251 | 2024-08-20 | 2024-08-23 | Cust 1190 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | MPCB connectivity needs to be done. | today we took access of the MPL Black Box and completed the MPCB configuration, and now data is live on the MPCB server. | today we took access of the MPL Black Box and completed the MPCB configuration, and now data is live on the MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG24073823252 | 2024-08-20 | 2024-08-21 | Cust 1221 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on CPCB + Local + SPCB server. | We took access to the system and now data is live on CPCB + Local + Common cloud server. | We took access to the system and now data is live on CPCB + Local + Common cloud server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFNG24074143253 | 2024-08-21 | 2024-08-23 | Cust 1187 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data upload issue. | Now issue has been resolved and data is live on CPCB server. | Now issue has been resolved and data is live on CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG24074663254 | 2024-08-22 | 2024-08-24 | Cust 708 | NAGPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | LCD of Multiparameter Display(1) | Display fade issue reported by customer. Replacement received at Nagpur Branch. Call for the commissioning of same. | New received Display were not working on Acrylic. | Closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24075443255 | 2024-08-23 | 2024-08-26 | Cust 1223 | NAGPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | optex series(optex series) | Visit required for commissioning of Dissolve Oxygen Sensor | Visited the site on 23.08.24 Commissioning of DO analyzer done installed at location Ctech 1&2 . | Commissioning of DO analyzer done | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24075443255 | 2024-08-23 | 2024-08-26 | Cust 1223 | NAGPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | optex series(optex series) | Visit required for commissioning of Dissolve Oxygen Sensor | Visited the site on 23.08.24 Commissioning of DO analyzer done installed at location Ctech 1&2 . | Commissioning of DO analyzer done | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24075503256 | 2024-08-23 | 2024-08-24 | Cust 1187 | NAGPUR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | RS 485 card not working & CPCB connectivity issue of Both The Analyzer. | RS 485 card replaced at site. Nexa aap router configuration done & taken Analyzer into service. | Closed | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24075673257 | 2024-08-23 | 2024-08-24 | Cust 1187 | NAGPUR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | FM make RS485 card(1) | RS 485 data not coming till the Black Box. | Available card checked & found not working. M/s SMS will arrange the card & give call for the commissioning of same. | Closed | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24077143258 | 2024-08-26 | 2024-08-27 | Cust 1187 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Central CPCB data upload issue. | now issue has been resolved and data is live on Central CPCB for 75 MLD STP- Outlet (75-ANT-502) | now issue has been resolved and data is live on Central CPCB for 75 MLD STP- Outlet (75-ANT-502) | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG24083973259 | 2024-09-05 | 2024-09-06 | Cust 1224 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB Server. | There was customer side issue data is live on CPCB Server. | There was customer side issue data is live on CPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFNG24086453260 | 2024-09-10 | 2024-09-10 | Cust 1207 | NAGPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | TOC ANALYSER | FM 101 TOC Analyser(FM 101 TOC) | Commissioning and inauguration of TOC | commissioning & Inauguration completed successfully | commissioning & Inauguration completed successfully | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24087403261 | 2024-09-12 | 2024-09-12 | Cust 1221 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on CPCB + Local + common cloud server. | we took access to the system and now data is live on CPCB + Local + Common cloud server. | we took access to the system and now data is live on CPCB + Local + Common cloud server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFNG24091023262 | 2024-09-20 | 2024-09-23 | Cust 1185 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data uploaded issue. | Found network issue and data now the data is live of the server. | Found network issue and data now the data is live of the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG24093453263 | 2024-09-25 | 2024-09-28 | Cust 1195 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data upload issue. | we took access of the Blackbox and observed that the internet was not pinging till the BB hence As of now we are closing this complaint and will reopen it again once the internet issue is resolved from the coutomer end. | As of now we are closing this complaint and will reopen it again once the internet issue is resolved from the coutomer end. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG24093603264 | 2024-09-25 | 2024-09-28 | Cust 1209 | NAGPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | TOC ANALYSER | FM 101 TOC Analyser(FM 101 TOC) | Scheduled PG test | Visited the site from 19 to 21 PG Test for TOC analyzer carried out | TOC analyzer PG test carried out successfully | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24094053265 | 2024-09-26 | 2024-09-28 | Cust 1186 | NAGPUR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | LCD of Multiparameter Display(1) | Visit required to check the MPA installled AWR & ETP location | visited the site on 26.09.24 Had discussion with MSPGCL person visited the site and checked the reported problem @ ETP & AWR MPA | Reported issues @ MPA analyzer checked | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24094053265 | 2024-09-26 | 2024-09-28 | Cust 1186 | NAGPUR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | LCD of Multiparameter Display(1) | Visit required to check the MPA installled AWR & ETP location | visited the site on 26.09.24 Had discussion with MSPGCL person visited the site and checked the reported problem @ ETP & AWR MPA | Reported issues @ MPA analyzer checked | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24094293266 | 2024-09-26 | 2024-09-27 | Cust 1187 | NAGPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | VISI pro DO(VISI pro DO) | Re-Commissioning of Aquamax basic DO analyzer. | Re-Commissioning of Aquamax basic DO analyzer is completed successfully. | closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24094573267 | 2024-09-27 | 2024-10-04 | Cust 555 | NAGPUR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Smartpro 8966 not getting powered ON. | Complaint wrongly logged under Critical type. New Complaint under correct category is being logged. | Complaint wrongly logged under Critical type. New Complaint under correct category is being logged. | YES | No Trip | Core | Y | Product / equipment failure due to which customers plant or part of plant is stopped which leaads to production loss. | Analytical - CHAKAN
AFNG24096853268 | 2024-09-30 | 2024-10-06 | Cust 1209 | NAGPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | TOC ANALYSER | FM 101 TOC Analyser(FM 101 TOC) | Completing PG Test formalities | PG Test completed successfully | closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24096863269 | 2024-09-30 | 2024-10-05 | Cust 1190 | NAGPUR | Analytical - CHAKAN | CLOSED | DIGITAL CONEECTIVITY | MULTIPARAMETER ANALYSER | FM make RS485 card(1) | Data not available on FM Cloud Server. | Reported connectivity issue has been checked & resolved successfully with the help of our Support Team at pune. | Closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24096933270 | 2024-09-30 | 2024-10-05 | Cust 1205 | NAGPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Erratic readings of Mulltiparameter Analyzer. | Reported issue of Erratic reading has been checked & resolved successfully over telphonic support. | Closed | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24096953271 | 2024-09-30 | 2024-10-03 | Cust 1190 | NAGPUR | Analytical - CHAKAN | CLOSED | DIGITAL CONEECTIVITY | MULTIPARAMETER ANALYSER | RS485 card-FM of Multiaparameter (1) | Call from Mr. Mahendra (M) 9343647571 on 30.09.24 that data is not uploading on cpcb server for Multiparameter and Ph issue | online with MR. Mahendra Support provided to calibrated the pH sensor online and Multiparameter data not available checked issue with internet observed | online with MR. Mahendra Support provided to calibrated the pH sensor online and Multiparameter data not available checked issue with internet observed | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFNG24098453272 | 2024-10-03 | 2024-10-04 | Cust 1187 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | CPCB Data uploaded issue. | Data is now live on the server. | Data is now live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG24098573273 | 2024-10-03 | 2024-10-05 | Cust 708 | NAGPUR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Display fade issue reported by customer. | New received Front door assembly with the Display & taken Analyzer into service. Online working of Analyzer observed & found ok. | Closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24098803274 | 2024-10-04 | 2024-10-07 | Cust 1195 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB + MPPCB data upload issue. | Now internet side issue is resolved and data is live on CPCB + MPPCB server. | Now internet side issue is resolved and data is live on CPCB + MPPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG24099873275 | 2024-10-07 | 2024-10-09 | Cust 1186 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | CPCB connectivity. | CPCB connectivity established successfully. | Closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24100063276 | 2024-10-07 | 2024-10-10 | Cust 1186 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB connectivity needs to be done. | we have completed the CPCB connectivity and now the data is live on both sites. CPCB & MPCB. | we have completed the CPCB connectivity and now the data is live on both sites. CPCB &amp; MPCB. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG24100243277 | 2024-10-08 | 2024-11-04 | Cust 1225 | NAGPUR | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Need to Commission the Multiparameter Analyser at ETP at Quantum Industries. - | Installation Guidelines provided. | Installation Guidelines provided to customer as per discussion with Nishant Tank | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFNG24101483278 | 2024-10-10 | 2024-10-10 | Cust 1221 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on CPCB + Local + SPCB | we took access to the system and now data is live on CPCB + Local + Common cloud servers. | we took access to the system and now data is live on CPCB + Local + Common cloud servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFNG24101893279 | 2024-10-11 | 2024-10-12 | Cust 1186 | NAGPUR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Call from user that the ETP analyzer data not loading on CPCB server | visited the site on 11.10.24 reported issue attended successfully | visited the reported issue checked and attended | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFNG24106373280 | 2024-10-21 | 2024-10-24 | Cust 1215 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue. | As of now we are closing this complaint and will reopen it once the hardware side issue is resolved for Stack 2 CEM. | As of now we are closing this complaint and will reopen it once the hardware side issue is resolved for Stack 2 CEM. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG24106493281 | 2024-10-21 | 2024-10-25 | Cust 497 | NAGPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC visit | Maintenance activities carried out at site & taken Analyzer into service. Online working performance observed & found ok. | Closed | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFNG24107613282 | 2024-10-23 | 2024-10-26 | Cust 777 | NAGPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Smartpro-con) | Commissioning call for the Conductivity Tranmitter. | Installation checked , Calibration verified & taken Analyzer into service. Successfully commissioned the Analyzer. | Commissioning done at site. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24108113283 | 2024-10-24 | 2024-10-26 | Cust 1081 | NAGPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | VISI pro DO(VISI pro DO) | Replacement received at Nagpur office. The same need to carry at site & commission it. | New sensor installation done & taken into service. Online working performance observed & found ok. | Commissioning done at site. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24108333284 | 2024-10-24 | 2024-10-26 | Cust 1190 | NAGPUR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | DO Transmitters | VISI pro DO(VISI pro DO) | Feed water DO reading is zero. | 02 Nos Sensor at Unit1 & Unit 4 replaced. Sample lines charged & taken Analyzer into service. Online working performance observed & found ok. | Commissioning done at site. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24108943285 | 2024-10-26 | 2024-10-30 | Cust 1190 | NAGPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Smartpro-con) | Site readiness confirmed. call for the Commissioning call of Conductivity Transmitters. | Conductivity Transmitters commissioned successfully & taken into service. | Closed. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24110023286 | 2024-10-28 | 2024-10-28 | Cust 1195 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera streaming issue on CPCB server. | We have accessed the system, and the camera streaming is now live on the CPCB server. | We have accessed the system, and the camera streaming is now live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFNG24113093287 | 2024-11-05 | 2024-11-10 | Cust 1219 | NAGPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning of Multiparameter analyzer. | Commissioning of Multiparameter analyzer is completed successfully. | closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24114033288 | 2024-11-07 | 2024-11-07 | Cust 1190 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera streaming issue on CPCB server. | The camera streaming is now live on the CPCB server. | The camera streaming is now live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFNG24114283289 | 2024-11-07 | 2024-11-08 | Cust 1187 | NAGPUR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | FM make RS485 card(1) | Reported issue of RS 485 not working. | Reported issue of issue in new supplied card has been resolved successfully. Now the data is uploading on to the server. | Closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24114493290 | 2024-11-08 | 2024-11-08 | Cust 1221 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on CPCB + Local + common cloud server. | We took access to the system and now data is live on CPCB + Local + Common cloud servers. | We took access to the system and now data is live on CPCB + Local + Common cloud servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFNG24114513291 | 2024-11-08 | 2024-11-08 | Cust 1204 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | The new version of DAS needs to be updated. | We have updated to the new DAS version 4.6, and the data is now live on the CPCB server. | We have updated to the new DAS version 4.6, and the data is now live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFNG24115703292 | 2024-11-11 | 2024-11-17 | Cust 1184 | NAGPUR | Analytical - CHAKAN | CLOSED | AMC | Aqua4Trans | Aqua4Trans(1) | Maintenance and calibration of ETP analyzers (pH, TSS & flow) | TSS sensor maintenance and calibration carried out. pH electrode not working. Customer will procure the same. | closed | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFNG24115703292 | 2024-11-11 | 2024-11-17 | Cust 1184 | NAGPUR | Analytical - CHAKAN | CLOSED | AMC | Aqua4Trans | Aqua4Trans(1) | Maintenance and calibration of ETP analyzers (pH, TSS & flow) | TSS sensor maintenance and calibration carried out. pH electrode not working. Customer will procure the same. | closed | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFNG24117833293 | 2024-11-14 | 2024-11-28 | Cust 1219 | NAGPUR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | TSS reading always showing zero reading. | After inspection observed that PD-2 Detector faulty & needs replacement. | Closed. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24118923294 | 2024-11-18 | 2024-11-18 | Cust 1195 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on CPCB and MPPCB. | We have accessed the system, and the data is now live on the CPCB server. Additionally, the camera streaming is also live on the CPCB server. | We have accessed the system, and the data is now live on the CPCB server. Additionally, the camera streaming is also live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFNG24119133295 | 2024-11-18 | 2024-11-18 | Cust 1221 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on CPCB + Local + Common cloud server. | As per telephonic conversation we took access to the system and now data is live on CPCB + Local + Common cloud servers. | As per telephonic conversation we took access to the system and now data is live on CPCB + Local + Common cloud servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFNG24121103296 | 2024-11-21 | 2024-11-22 | Cust 1210 | NAGPUR | DCS - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | MICROCON+ER SERIES | CONTROLLER MODULE (ER-201)(ER-201) | Data is not uploading to CPCB and MPCB servers | Issue has been resolved and data from stack is now live on the CPCB and MPCB server. | Issue has been resolved and data from stack is now live on the CPCB and MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFNG24123263297 | 2024-11-25 | 2024-11-28 | Cust 1187 | NAGPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Touchpad of display unit not functioning properly. | Display Touch found not functioning & needs replacement. | Closed. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24123273298 | 2024-11-25 | 2024-11-28 | Cust 1187 | NAGPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | mA output not matching with the DCS reading. | After reset mA output issue resolved successfully. | Closed. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24125193299 | 2024-11-28 | 2024-12-03 | Cust 1207 | NAGPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | TOC ANALYSER | FM 101 TOC Analyser(FM 101 TOC) | Training to NSL team | Visited the site Training classroom and on site done for NSL personnel. | Visited the site Training classroom and on site done for NSL personne | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24150023300 | 2025-01-13 | 2025-01-21 | Cust 497 | NAGPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC visit. | Maintenance activities carried out & taken analyzer into service. Online working performance observed & found ok. | AMC visit done. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFNG24150763301 | 2025-01-14 | 2025-01-17 | Cust 1225 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Blackbox Commissioning | Here while commissioning there was an issue with the BB and now we completed the Blackbox (BB) configuration with the new BB and ETP data is coming to the Blackbox. | Here while commissioning there was an issue with the BB and now we completed the Blackbox (BB) configuration with the new BB and ETP data is coming to the Blackbox. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG24150963302 | 2025-01-15 | 2025-01-17 | Cust 497 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data uploading issue on MPCB server. | Issue has been resolved and data from ETP is now live on the MPCB server. Also we have upgraded Enviroconnect DAS version to the latest. | Issue has been resolved and data from ETP is now live on the MPCB server. Also we have upgraded Enviroconnect DAS version to the latest. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFNG24151353303 | 2025-01-15 | 2025-03-04 | Cust 1225 | NAGPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning call of Multiparameter Analyzer. | Analyser Commissioned Successfully. Connectivity established & complete data is availalble on cloud | Analyser commissioned & connected to cloud successfully | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFNG24153173304 | 2025-01-20 | 2025-01-21 | Cust 1226 | NAGPUR | Codel Products - CHAKAN | CLOSED | CALIBRATION JOB | GCEM 40 Series New Gas Analyser | GCEM 40xx series Head Only(GCEM 40xx Head 9) | Analyzer is showing D-C error | Attended the issue by reset and re-installation of analyzer backup. | closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24153173304 | 2025-01-20 | 2025-01-21 | Cust 1226 | NAGPUR | Codel Products - CHAKAN | CLOSED | CALIBRATION JOB | GCEM 40 Series New Gas Analyser | GCEM 40xx series Head Only(GCEM 40xx Head 9) | Air supply failure | Instrument air supply to analyzer was not available. | closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24153483305 | 2025-01-20 | 2025-01-27 | Cust 555 | NAGPUR | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Smartpro-con) | Connection details of Cell not clear | Telephonic support extended to resolve the issue | Telephonic support extended to resolve the issue | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFNG24153563306 | 2025-01-20 | 2025-01-23 | Cust 1217 | NAGPUR | Analytical - CHAKAN | CLOSED | DIGITAL CONEECTIVITY | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Connectivity | Visited the site on 20.01.25 MPA data upto black box made avilable. System taken online by connectivity team on 21.01.25 connectivty done for MPA analyzer. Mr. Tausif visited the site | Visited the site on 20.01.25 MPA data upto black box made avilable. System taken online by connectivity team on 21.01.25 connectivty done for MPA analyzer. Mr. Tausif visited the site | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFNG24153643307 | 2025-01-20 | 2025-01-21 | Cust 526 | NAGPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | calibration of pH, Conductivity and ORP analyzers | calibration of conductivity meter is completed. | closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24153643307 | 2025-01-20 | 2025-01-21 | Cust 526 | NAGPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | calibration of pH, Conductivity and ORP analyzers | calibration of conductivity meter is completed. | closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24153643307 | 2025-01-20 | 2025-01-21 | Cust 526 | NAGPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | calibration of pH, Conductivity and ORP analyzers | calibration of pH meter is completed. | closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24153643307 | 2025-01-20 | 2025-01-21 | Cust 526 | NAGPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | calibration of pH, Conductivity and ORP analyzers | calibration of pH meter is completed. | closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24153643307 | 2025-01-20 | 2025-01-21 | Cust 526 | NAGPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | calibration of pH, Conductivity and ORP analyzers | calibration of pH meter is completed. | closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24153643307 | 2025-01-20 | 2025-01-21 | Cust 526 | NAGPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | calibration of pH, Conductivity and ORP analyzers | calibration of conductivity meter is completed. | closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24153903308 | 2025-01-21 | 2025-01-21 | Cust 1217 | NAGPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Connectivity issue. | Data coming till the Black Box. Connectivity team need to check the issue at the server end. | Closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24153913309 | 2025-01-21 | 2025-01-25 | Cust 1186 | NAGPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration call of ETP analyzer. | Maintenance done for all 4 ETP Analyzer & taken analyzer into service. Online working performance observed & found ok. | AMC visit done. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFNG24155653310 | 2025-01-24 | 2025-01-27 | Cust 1204 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on CPCB server. | We have not received the RAMC for the current year . so we are closing this complaint for now. | We have not received the RAMC for the current year . so we are closing this complaint for now. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFNG24156143311 | 2025-01-25 | 2025-01-26 | Cust 1186 | NAGPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Showing readings are on higher side. | Maintenance done & taken Analyzer into service. Online working performance observed & found ok. | Closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24157793312 | 2025-01-28 | 2025-01-31 | Cust 497 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data availability issue on MPCB server. | now the data availability is 100%. | now the data availability is 100%. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFNG24160933313 | 2025-01-31 | 2025-02-03 | Cust 1185 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | ZLD flow data not updated in online CPCB server | We took access of the Blackbox and observed that the internet was not pinging till the BB and now the same has been resolved and data is live now. | We took access of the Blackbox and observed that the internet was not pinging till the BB and now the same has been resolved and data is live now. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG24162203314 | 2025-02-01 | 2025-02-03 | Cust 1225 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | MPCB connectivity needs to be done. | We have completed the MPCB connectivity and Data is live on the MPCB server. | We have completed the MPCB connectivity and Data is live on the MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG24162773315 | 2025-02-03 | 2025-02-06 | Cust 1110 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | EnviroConnect software Re-installation. | We took access to the system and successfully completed the software re-installation and configuration. And now the data is live on the CPCB server. | We took access to the system and successfully completed the software re-installation and configuration. And now the data is live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFNG24163473316 | 2025-02-04 | 2025-04-28 | Cust 1207 | NAGPUR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | TOC ANALYSER | 24V Pump Motor/gearbox(017?018) | Analyzer pump not working. required replacement | Replaced the pump | attended and closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24163863317 | 2025-02-05 | 2025-02-07 | Cust 1215 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to to CPCB server. | the issue has been resolved from the hardware side as the data is showing live on the CPCB server. Also we changed channel number for stack 1. | the issue has been resolved from the hardware side as the data is showing live on the CPCB server. Also we changed channel number for stack 1. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFNG24166023318 | 2025-02-11 | 2025-02-28 | Cust 1227 | NAGPUR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Aqua4Trans | Turbidity Input card of Aqua4Trans(1) | Turbidity analyser failure. Please visit to resolve the issue. | Telephonic support extended | Telephonic support extended | YES | No Trip | Core | Y | Product / equipment failure due to which customers plant or part of plant is stopped which leaads to production loss. | Analytical - CHAKAN
AFNG24167613319 | 2025-02-13 | 2025-02-17 | Cust 1184 | NAGPUR | Analytical - CHAKAN | CLOSED | AMC | Aqua4Trans | Turbidity Input card of Aqua4Trans(1) | ETP Analyzers maintenance and calibration | TSS analyzer Maintenance and Calibration carried out. | closed | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFNG24167613319 | 2025-02-13 | 2025-02-17 | Cust 1184 | NAGPUR | Analytical - CHAKAN | CLOSED | AMC | Aqua4Trans | Turbidity Input card of Aqua4Trans(1) | ETP Analyzers maintenance and calibration | TSS analyzer Maintenance and Calibration carried out. | closed | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFNG24170713320 | 2025-02-21 | 2025-02-21 | Cust 1195 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera streaming issue on CPCB server. | Now the camera is live on the CPCB server. | Now the camera is live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFNG24171763321 | 2025-02-24 | 2025-03-01 | Cust 1189 | NAGPUR | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Maintenance and calibration of Multiparameter analyzer. | Maintenance and calibration of Multiparameter analyzer is completed. | closed | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFNG24172043322 | 2025-02-25 | 2025-02-28 | Cust 1192 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPP online data not going to CPCB RTDMS site | Issue resolved by the coutomer. | Issue resolved by the coutomer. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG24174533323 | 2025-02-28 | 2025-03-03 | Cust 1189 | NAGPUR | Analytical - CHAKAN | CLOSED | DIGITAL CONEECTIVITY | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Digital Connectivity needs to be done. | Cloud server connectivity has been completed and Data is live on the Cloud server. | Cloud server connectivity has been completed and Data is live on the Cloud server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG24177103324 | 2025-03-04 | 2025-03-05 | Cust 1197 | NAGPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC visit. | Maintenance activities carried out & taken Analyzer into service. Online working performance observed & found ok. | AMC visit done. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFNG24177123325 | 2025-03-04 | 2025-03-10 | Cust 1217 | NAGPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | VISI pro DO(VISI pro DO) | Visit required for commissioing of Visiwater | visited the site on 4.03.25 Commissioning of Dissolv Oxygen meter qty 02 nos & Flow Meter 01 no. Done | Visited the site commissioning of DO & Flow meter done | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24177573326 | 2025-03-05 | 2025-03-07 | Cust 497 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to CPCB and MPCB | Data is now live on the CPCB and MPCB servers. | Data is now live on the CPCB and MPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFNG24178003327 | 2025-03-05 | 2025-03-09 | Cust 1228 | NAGPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Call for the inspection/Calibration of Analyzer. | Inspection done & report submitted to M/s Adani VIPL site. Request them to upgrade the instrument. Our Sales team will share the praposal soon. | Inspection done. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFNG24181033328 | 2025-03-11 | 2025-03-18 | Cust 1184 | NAGPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | Digital Product | RTU301() | Commissioning and SCADA configuration of DM plant analyzers. | Commissioning and SCADA configuration of DM plant analyzers completed | commissioning completed successfully | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24181443329 | 2025-03-12 | 2025-03-14 | Cust 1187 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data uploded issue. | As of now we are closing this complaint and will reopen it again once the hardware issue has been resolved from coustomer end. | As of now we are closing this complaint and will reopen it again once the hardware issue has been resolved from coustomer end. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG24185683330 | 2025-03-24 | 2025-03-27 | Cust 1192 | NAGPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Visite Under EU. | Maintenance done, Calibration performed & taken analyzer into service. | Closed | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG24186293331 | 2025-03-25 | 2025-03-27 | Cust 1229 | NAGPUR | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Online support required for commissioning and calibration of pH transmitters | online with user for installation, commissioning pH sensor commissioned successfully. | online with user for installation, commissioning pH sensor commissioned successfully. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFNG24188063332 | 2025-03-28 | 2025-03-29 | Cust 1187 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data uploded issue. | ow the Data is live on the CPCB server. Please find the attached screenshot for your reference. | ow the Data is live on the CPCB server. Please find the attached screenshot for your reference. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG25002803333 | 2025-04-08 | 2025-04-13 | Cust 1209 | NAGPUR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | TOC ANALYSER | Protoc 300 24V Pump Motor/gearbox(17018) | Replacement of spares | Analyzer pump motor was detected faulty during last visit. replaced the same and make analyzer operational. | closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG25003743334 | 2025-04-10 | 2025-04-11 | Cust 1195 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data from flowmeter is not uploading to CPCB server also camera is not live. | Now the data from ETP flow is live on the CPCB server. Also the camera is streaming live on the server. | Now the data from ETP flow is live on the CPCB server. Also the camera is streaming live on the server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFNG25004093335 | 2025-04-10 | 2025-04-11 | Cust 1190 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | DATCON | DATCON(DATCON) | ETP data is not updating on the CPCB server. | We have updated the DATCon patch to the latest version so now the data is live on the CPCB server. | We have updated the DATCon patch to the latest version so now the data is live on the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG25004143336 | 2025-04-10 | 2025-04-11 | Cust 1209 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data uploded issue. | We took access and made some modifications in the modbus configuration so now the Data is live on the CPCB. | We took access and made some modifications in the modbus configuration so now the Data is live on the CPCB. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG25006483337 | 2025-04-17 | 2025-04-26 | Cust 1185 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | For Flow data is not updating correctly. | We took access and asked them to check the 4-20 mA counts and also made some modifications on the BB configuration so now the issue has been resolved | We took access and asked them to check the 4-20 mA counts and also made some modifications on the BB configuration so now the issue has been resolved | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG25006763338 | 2025-04-18 | 2025-04-22 | Cust 1219 | NAGPUR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Sample pump not working. | Sample pump replaced successfully & taken analyzer into service. Online working performance observed & found ok. | Closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG25008603339 | 2025-04-23 | 2025-04-26 | Cust 1210 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Flow data uploded issue. | We took access of the Blackbox for the Flow meter and completed the connectivity, and the data was live.But we have observed that the flow data has not been updating to the CPCB server since yesterday hence we have asked them ro check internet connectivity so as of now we are closing this complaint. | as of now we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG25009023340 | 2025-04-24 | 2025-04-26 | Cust 497 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data availability issue. | As of now, we are closing this complaint and will reopen this case once we get remote access of the system from coustomer end | As of now, we are closing this complaint and will reopen this case once we get remote access of the system from coustomer end | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG25009793341 | 2025-04-25 | 2025-05-02 | Cust 1209 | NAGPUR | Analytical - CHAKAN | CLOSED | AMC | TOC ANALYSER | FM 101 TOC Analyser(FM 101 TOC) | Scheduled Maintenance | Scheduled Maintenance done | attended and closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG25012933342 | 2025-04-30 | 2025-05-03 | Cust 1210 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Flow data upload issue. | We have updated the BB with the new patch and now the Data is live continously | We have updated the BB with the new patch and now the Data is live continously | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG25016263343 | 2025-05-06 | 2025-05-09 | Cust 1185 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | ZLD OUTLET DATA IS NOT UPLODE ON CPCB | We have updated the BB with the new patch as the BB was hanging so now the is live . Furthermore, we have successfully updated the internet IP details for the ZLD system as provided by their IT team. | We have updated the BB with the new patch as the BB was hanging so now the is live . Furthermore, we have successfully updated the internet IP details for the ZLD system as provided by their IT team. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG25017073344 | 2025-05-07 | 2025-05-09 | Cust 1209 | NAGPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | TOC ANALYSER | FM 101 TOC Analyser(FM 101 TOC) | Maintenance visit. | Maintenance activities carried out, Reagent prepared & taken analyzer into service. Online working performance observed & found ok. | Closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG25017243345 | 2025-05-08 | 2025-05-09 | Cust 1209 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data upload issue. | The issue has been resolved. Data from SIM was not going. | The issue has been resolved. Data from SIM was not going. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG25019493346 | 2025-05-14 | 2025-05-16 | Cust 1186 | NAGPUR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Troubleshooting of multiparameter analyzers installed at MSPGCL koradi 210MW. | All 03 nos analyzers Issues were attended, analyzer taken into service. maintenance and calibration carried out. | closed | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFNG25019493346 | 2025-05-14 | 2025-05-16 | Cust 1186 | NAGPUR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Troubleshooting of multiparameter analyzers installed at MSPGCL koradi 210MW. | All 03 nos analyzers Issues were attended, analyzer taken into service. maintenance and calibration carried out. | closed | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFNG25019493346 | 2025-05-14 | 2025-05-16 | Cust 1186 | NAGPUR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Troubleshooting of multiparameter analyzers installed at MSPGCL koradi 210MW. | All 03 nos analyzers Issues were attended, analyzer taken into service. maintenance and calibration carried out. | closed | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFNG25019973347 | 2025-05-15 | 2025-05-16 | Cust 1187 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data upload issue. | we took access of the black box and now the internet side issue has been resolved and data is live on the CPCB server. | we took access of the black box and now the internet side issue has been resolved and data is live on the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG25021393348 | 2025-05-19 | 2025-05-23 | Cust 1195 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera streaming issue on CPCB server | Found the camera video streamer folder missing/deleted from the PC, so we have re configured the same and the camera is streaming live on the CPCB server. Internet issue has been resolved and flowmeter data is reflecting on the CPCB server. | Found the camera video streamer folder missing/deleted from the PC, so we have re configured the same and the camera is streaming live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFNG25021513349 | 2025-05-19 | 2025-05-29 | Cust 1187 | NAGPUR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | AQUADO Analyzer | OXYSENS(OXYSENS) | DO analyzers with Oxysense sensors supplied at M/s SMS ltd., Wajuj plant is not responding. | Replacement material related to RF OA NO.-5526800134 has been booked through Mark Express Docket No-210141333 dt- 28.05.2025. | Replacement material of RF OA NO.-5526800134 has been booked through Mark Express Docket No-210141333 dt- 28.05.2025. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFNG25021513349 | 2025-05-19 | 2025-05-29 | Cust 1187 | NAGPUR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | AQUADO Analyzer | OXYSENS(OXYSENS) | DO analyzers with Oxysense sensors supplied at M/s SMS ltd., Wajuj plant is not responding. | Replacement material related to RF OA NO.-5526800134 has been booked through Mark Express Docket No-210141333 dt- 28.05.2025. | Replacement material of RF OA NO.-5526800134 has been booked through Mark Express Docket No-210141333 dt- 28.05.2025. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFNG25021753350 | 2025-05-19 | 2025-05-20 | Cust 1190 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | MPCB data upload issue. | we took access of the black box and now the time sync issue has been resolved and data is live on the MPCB server. | we took access of the black box and now the time sync issue has been resolved and data is live on the MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG25022073351 | 2025-05-20 | 2025-05-23 | Cust 1217 | NAGPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | COD, TSS and pH values improper reading. | COD & TSS calibration carried out. readings under observation. issue found in pH input card. FMPL will arrange for the pH input card and calibrator and visit again. | closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG25022273352 | 2025-05-20 | 2025-05-26 | Cust 1209 | NAGPUR | Analytical - CHAKAN | CLOSED | AMC | TOC ANALYSER | TOC 1950 Single StreamANALYSER ANATEL(1950 TOC) | Routine Maintenance | Routine Maintenance done | attended and closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG25030343353 | 2025-06-03 | 2025-06-04 | Cust 1215 | NAGPUR | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | RTDAS upgradation needs to be done. | Today we took access of the system and updated the RTDAS, now data is live on CPCB server | Today we took access of the system and updated the RTDAS, now data is live on CPCB server | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFNG25030543354 | 2025-06-03 | 2025-06-05 | Cust 1217 | NAGPUR | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | PH Input Card for Multiparameter(AWA-PH) | pH input card faulty. | pH input card replaced, Calibration carried out. analyzer is now working ok | closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNG25032563355 | 2025-06-09 | 2025-06-14 | Cust 497 | NAGPUR | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC visit. | maintenance done & taken MPA analyzer into service. Online working performance observed & found ok. | AMC visit done. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFNG25033043356 | 2025-06-10 | 2025-06-13 | Cust 1209 | NAGPUR | Analytical - CHAKAN | CLOSED | AMC | TOC ANALYSER | FM 101 TOC Analyser(FM 101 TOC) | Scheduled Maintenance | During the present visit, following maintenance activities were carried out. A. Top-up of Reagent B. Sample line cleaning C. Checking of physical condition of all the tubings and other parts of the Analyser | attended and closed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFNR24181113357 | 2025-03-12 | 2025-03-14 | Cust 1230 | NIGERIA | Steam System Division | CLOSED | INSTALLATION/ ERECTION | RTD sensor | RTD sensor (1) | Equipment packing list check Installation guidelines for CRM and sensor | Identification of material Installation guidelines | identifications of material | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFNR24181113357 | 2025-03-12 | 2025-03-14 | Cust 1230 | NIGERIA | Steam System Division | CLOSED | INSTALLATION/ ERECTION | RTD sensor | RTD sensor (1) | Equipment packing list check Installation guidelines for CRM and sensor | Identification of material Installation guidelines | identifications of material | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFNR24181113357 | 2025-03-12 | 2025-03-14 | Cust 1230 | NIGERIA | Steam System Division | CLOSED | INSTALLATION/ ERECTION | RTD sensor | RTD sensor (1) | Equipment packing list check Installation guidelines for CRM and sensor | Identification of material Installation guidelines | identifications of material | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFNR24181113357 | 2025-03-12 | 2025-03-14 | Cust 1230 | NIGERIA | Steam System Division | CLOSED | INSTALLATION/ ERECTION | RTD sensor | RTD sensor (1) | Equipment packing list check Installation guidelines for CRM and sensor | Identification of material Installation guidelines | identifications of material | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFNR24181113357 | 2025-03-12 | 2025-03-14 | Cust 1230 | NIGERIA | Steam System Division | CLOSED | INSTALLATION/ ERECTION | RTD sensor | RTD sensor (1) | Equipment packing list check Installation guidelines for CRM and sensor | Identification of material Installation guidelines | identifications of material | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFNR24181113357 | 2025-03-12 | 2025-03-14 | Cust 1230 | NIGERIA | Steam System Division | CLOSED | INSTALLATION/ ERECTION | RTD sensor | RTD sensor (1) | Equipment packing list check Installation guidelines for CRM and sensor | Identification of material Installation guidelines | identifications of material | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFNR24181113357 | 2025-03-12 | 2025-03-14 | Cust 1230 | NIGERIA | Steam System Division | CLOSED | INSTALLATION/ ERECTION | CRS | FJP - FLASH JET PUPM - SIZE 80NB(080FJP) | Equipment packing list check Installation guidelines for CRM and sensor | Identification of material Installation guidelines | identifications of material | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFPN23033843358 | 2023-06-02 | 2023-06-05 | Cust 809 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | data not updating on the cpcb/mpcb server. | As of now we are closing this complaint and will reopen once gets seamless access of the system. | As of now we are closing this complaint and will reopen once gets seamless access of the system. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23034043359 | 2023-06-02 | 2023-06-09 | Cust 1231 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Smartpro-con) | conductivity transmitter not working , sensor CS40 | visited the site cheched the actual conductivity of the water which is much higher for current sensr of 0.01 cell constant , customer to purchase he new correct cell constant | visited the site cheched the actual conductivity of the water which is much higher for current sensr of 0.01 cell constant , customer to purchase he new correct cell constant | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23035753360 | 2023-06-06 | 2023-06-07 | Cust 517 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading on CPCB server. | COM port changed, now data is live on server. | COM port changed, now data is live on server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23036043361 | 2023-06-07 | 2023-06-08 | Cust 809 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data uploaded issue. | Data is live on the server hence we are closing this complaint. | Data is live on the server hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23036263362 | 2023-06-07 | 2023-06-11 | Cust 811 | PUNE | Analytical - CHAKAN | CLOSED | DIGITAL CONEECTIVITY | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | connectivity visit | connectivity visit | CONNECTIVITY VISIT | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23036263362 | 2023-06-07 | 2023-06-11 | Cust 811 | PUNE | Analytical - CHAKAN | CLOSED | DIGITAL CONEECTIVITY | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | connectivity visit | CONNECTIVITY VISIT | CONNECTIVITY VISIT | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23036633363 | 2023-06-08 | 2023-06-29 | Cust 1232 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | DATCON | DATCON(DATCON) | CPCB data uploded issue. | Our development team has made some changes and now the data is live on both servers. | The issue has been resolved and data is live hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23036673364 | 2023-06-08 | 2023-06-09 | Cust 811 | PUNE | Analytical - CHAKAN | CLOSED | DIGITAL CONEECTIVITY | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Digital connectivity needs to be done. | We have completed the Digital connectivity and data is live on the cloud server. | Data is live hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23036813365 | 2023-06-08 | 2023-06-12 | Cust 1233 | PUNE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Need to install Scanner and connectivity of the same | Telephonic support provided and resolve the issue of calibration | Telephonic support provided and resolve the issue of calibration | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFPN23036833366 | 2023-06-08 | 2023-06-09 | Cust 809 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | need to do software installation and configuration | We took access to the new system and successfully completed the software installation and configuration.But found that the CPCB/MPCB URL is blocked in network so, asked customer to resolve the same. | We took access to the new system and successfully completed the software installation and configuration.But found that the CPCB/MPCB URL is blocked in network so, asked customer to resolve the same. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23037173367 | 2023-06-09 | 2023-06-11 | Cust 1233 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning of Scanner & Connectivity | Commissioning of Scanner & Connectivity | COMMISSIONING OF SCANNER | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFPN23037873368 | 2023-06-12 | 2023-06-15 | Cust 809 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | data not updating on the cpcb /mpcb server | Data is now live on the CPCB/MPCB server. services restarted and ports were closed so asked their IT team to open the ports and urls. | Data is now live on the CPCB/MPCB server. services restarted and ports were closed so asked their IT team to open the ports and urls. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23037933369 | 2023-06-12 | 2023-06-15 | Cust 623 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera streaming issue on CPCB server. | We took access and now the camera streaming is live on the server. | We took access and now the camera streaming is live on the server. Hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23038333370 | 2023-06-12 | 2023-06-14 | Cust 517 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | MPCB data uploaded issue. | The NTP port is blocked in the network so we manually configured the time and now data is uploading on the MPCB server. | Data is live hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23038673371 | 2023-06-13 | 2023-06-17 | Cust 1234 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Electrode | Chemotrode 120(HT238 760) | communication error | DO SENSOR COMMUNICATION ISSUE | DO SENSOR ISSUE VISIT | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23038743372 | 2023-06-13 | 2023-06-14 | Cust 904 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | CPU with LCD of Multiparameter Display(1) | Display is completely blurred and flickering. | Needs to send at chakan factory for repair. Factory address and contact given to customer. customer is ready to send the same. Repair will be done on chargeable basis. | CLOSED | NO | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN23038743372 | 2023-06-13 | 2023-06-14 | Cust 904 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | PH Sensor for Multiparameter(AWA-PH-SENS) | PH sensor is not working at all. | Will send offer to customer. Customer is ready to purchase. | CLOSED | YES | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN23039103373 | 2023-06-13 | 2023-06-17 | Cust 798 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Chlorine Analyzer-MAC 100 | MAC 100(1) | Site Address: Ghodawat consumer ltd A/P chipri Tal - shirol Dist kolhapur 416101 Contact Person - Mr. Tejas Deore +91 90671 91975 | CHLORINE ANALYSER ISSUE | VISIT FOR CHLORINE ANALYSER | YES | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN23040463374 | 2023-06-15 | 2023-06-15 | Cust 841 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | MPCB connectivity needs to be done. | MPCB connectivity has been completed. | MPCB connectivity has been completed. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23040863375 | 2023-06-16 | 2023-06-16 | Cust 1235 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data uploaded issue. | Not found any issue as the data is live on the server. | Not found any issue as the data is live on the server. Hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23040983376 | 2023-06-16 | 2023-06-21 | Cust 1236 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT DESIGN FAILURE | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | Calibration needs to be done . | Suggested to send at factory. | CLOSED | NO | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN23040983376 | 2023-06-16 | 2023-06-21 | Cust 1236 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT DESIGN FAILURE | AQUADO Analyzer | OXYSENS(OXYSENS) | DO sensor is not working. | Sales engineer proposed new model visitrace. | CLOSED | NO | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN23041113377 | 2023-06-16 | 2023-06-17 | Cust 811 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | DATA WA NOT UPDATING ON cpdms site since feb23 | suggested for internet availability. | CONNECTIVITY ISSUE SINCE FEB23 due to no internet on sim from feb | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23042003378 | 2023-06-19 | 2023-06-24 | Cust 1237 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | DUE TO UNAVAILABLE INTERNET ON PANNEL DATA IS NOT COMMUNICATING........ OBSERVED LAN CABLE INTERNET PROBLEM,MEDIA CONVERTER WAS FAULTY | DUE TO UNAVAILABLE INTERNET ON PANNEL DATA IS NOT COMMUNICATING........ OBSERVED LAN CABLE INTERNET PROBLEM,MEDIA CONVERTER WAS FAULTY | INTERNET ISSUE | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23042773379 | 2023-06-20 | 2023-06-22 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | software reinstallation , 1st time Under AMC | We took access to the system and did software reinstallation and configuration. As discussed , right now the industry is shut down so data will not be live | We took access to the system and did software reinstallation and configuration. As discussed , right now the industry is shut down so data will not be live | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23042903380 | 2023-06-20 | 2023-06-24 | Cust 1239 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC VISIT FOR MPA | AMC VISIT FOR MPA | AMC VISIT FOR MPA | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23043023381 | 2023-06-20 | 2023-06-21 | Cust 1240 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | Calibration of DO sensor | Calibration of DO sensor has done. | RESOLVED | NO | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN23043023381 | 2023-06-20 | 2023-06-21 | Cust 1240 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | Calibration of DO sensor | Calibration of DO sensor has done. | RESOLVED | NO | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN23043023381 | 2023-06-20 | 2023-06-21 | Cust 1240 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Electrode | Polilyte PRO 120(HT238 411) | Installation and Commissioning of PH electrode | Installation , commission and calibration done. | RESOLVED | NO | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN23043613382 | 2023-06-21 | 2023-06-24 | Cust 1239 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC VISIT FOR MPA | AMC VISIT FOR MPA, DETECTORS ARE VERY LOW NEED TO CHANGE FLOW CELL | AMC VISIT FOR MPA | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23043623383 | 2023-06-21 | 2023-06-24 | Cust 1237 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | CETP DATA UPLOAD ISSUE | CETP DATA UPDATING ISSUE FROM ECOLYSER SIDE...DATA LOCALY PROPERLY UPDATING FROM OUR SIDE | CETP ECOLYSER DATA NOT UPDATING...DATA UPDATING ISSUE IS FROM ECOLYSER SIDE | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23043943384 | 2023-06-22 | 2023-06-29 | Cust 1237 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | data not updating on the ecolyser server. | The issue is resolved by the ecolyser team and correct data is updated on the server. Values are cross checked with customer . | The issue is resolved by the ecolyser team and correct data is updated on the server. Values are cross checked with customer . | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23045213385 | 2023-06-24 | 2023-06-29 | Cust 811 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Analyser display not working | AMC Visit carried out | AMC Visit carried out | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN23047003386 | 2023-06-27 | 2023-06-29 | Cust 1241 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | cx1000 chnaged to cx2000 , so did the chnages at cpcb server | We have made changes in configuration at our end as per requirement. Right now data is live on the cpcb server .Cross checked all values with you. | We have made changes in configuration at our end as per requirement. Right now data is live on the cpcb server .Cross checked all values with you. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23050173387 | 2023-07-03 | 2023-07-06 | Cust 1237 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | data mismatch issue on ecolyser server. | The issue is resolved by the Ecolyser team and correct data is now updating on Ecolyser server.We have cross checked all values with you . | The issue is resolved by the Ecolyser team and correct data is now updating on Ecolyser server.We have cross checked all values with you . | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23050303388 | 2023-07-03 | 2023-07-05 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | industry turn on - data not updating on local/cpcb server. | We took access to the system and did the configuration. Data is now live on CPCB/SPCB server.As discussed ,the plant is again going to shut down . | We took access to the system and did the configuration. Data is now live on CPCB/SPCB server.As discussed ,the plant is again going to shut down . | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23051863389 | 2023-07-06 | 2023-07-08 | Cust 1232 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | comissioning of DO sensor | comissioning of DO sensor | COMMISSIONING OF DO SENSOR | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23051923390 | 2023-07-06 | 2023-07-07 | Cust 904 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | CPU with LCD of Multiparameter Display(1) | MPA display found faulty | Display sent to chakan factory for repair . | CLOSED | NO | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN23052643391 | 2023-07-08 | 2023-07-08 | Cust 1234 | PUNE | Analytical - CHAKAN | CLOSED | APPLICATION ISSUE | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | PH & DO reading not showing on display....shows error | COMMUNICTION SOLVED AFTER BAUDRATE 34800 | COMMUNICATION ERROR SOLVED AFTER BAURD RATE 34800 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23053323392 | 2023-07-10 | 2023-07-12 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB + MPCB data upload issue. | Data is live on all the servers, hence we are closing this complaint. | Data is live on all the servers, hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23053403393 | 2023-07-10 | 2023-07-13 | Cust 1239 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data upload issue. | commercially not cleared. | commercially not cleared. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23053413394 | 2023-07-10 | 2023-07-12 | Cust 517 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | MPCB data uploaded issue. | We have changed the timing of BB and now the data is live on the server. | Issue has been resolved and data is live hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23056413395 | 2023-07-14 | 2023-07-17 | Cust 1237 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Data is not uploaded on the Ecolyser server. | After restarting the black box, the data started uploading to the server. | Data is live hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23056453396 | 2023-07-14 | 2023-07-14 | Cust 809 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | data not updating on the cpcb/mpcb server | We took access to the system and checked that some CPCB files were deleted. .Right now the issue is resolved and data is live .das files added again. | We took access to the system and checked that some CPCB files were deleted. .Right now the issue is resolved and data is live . | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23056463397 | 2023-07-14 | 2023-07-17 | Cust 795 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera is not streaming on CPCB server. | The camera IP is not opening in the browser, suggested customer to check with IT Team. As of now closing this complaint as we have not received any reply from customer side. | The camera IP is not opening in the browser, suggested customer to check with IT Team. As of now closing this complaint as we have not received any reply from customer side. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23056583398 | 2023-07-14 | 2023-07-16 | Cust 1242 | PUNE | Analytical - CHAKAN | CLOSED | SALES PROMOTION | Aqua4Trans | Aqua4 trans-A4TW033000(1) | received confirmation from the Mr. Ganesh to plan visit to install the DO analyzer | COMISSIONG VISIT | COMMISSIONING VISIT | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23056743399 | 2023-07-15 | 2023-07-16 | Cust 1243 | PUNE | Analytical - CHAKAN | CLOSED | APPLICATION ISSUE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | BOD error | MPA BOD ERROR SOLVED | BOD ERROR SOLVED FOIR MPA | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23056753400 | 2023-07-15 | 2023-07-16 | Cust 1237 | PUNE | Analytical - CHAKAN | CLOSED | DIGITAL CONEECTIVITY | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | CONNECTIVITY TO SERVER & CETP | CETP CONNECTIVITY VISIT | CETP CONNECTIVITY VISIT | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23056763401 | 2023-07-15 | 2023-07-16 | Cust 1237 | PUNE | Analytical - CHAKAN | CLOSED | DIGITAL CONEECTIVITY | Aqua4Trans | Aqua4Trans(1) | CETP CONNECTIVITY ISSUE | CETP CONNECTIVITY | CETP CONNECTIVITY VISIT | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23057583402 | 2023-07-17 | 2023-07-23 | Cust 811 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Analyser cleaning pump issue,,,& data issue | TUBEING CHANGED ,CLEANING PUMP ISSUE SOLVED | MULTIPARA ANALYSER CLEANING PUMP ISSUE | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23058133403 | 2023-07-18 | 2023-07-21 | Cust 651 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading on CPCB server. | did not get access .customer will give access on monday. | did not get access .customer will give access on monday. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23058293404 | 2023-07-18 | 2023-07-23 | Cust 1244 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Multiparameter issue | Multiparameter issue | VISIT DONE | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23060993405 | 2023-07-24 | 2023-07-30 | Cust 719 | PUNE | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Proactive visit to Multiparamter CX 2000 upgradation | Multiparamter CX 2000 upgradation offer sent | Multiparamter CX 2000 upgradation offer sent | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFPN23061553406 | 2023-07-25 | 2023-07-26 | Cust 1243 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data mismatch issue on ECOLYSER server. | we took access of the BB and we are getting 0 readings only till the Black Box the same readings are reflected on the Ecolyser server and now the hardware issue has been resolved and the correct readings are reflected on the Ecolyser server. | Issue has been resolved hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23061593407 | 2023-07-25 | 2023-07-26 | Cust 841 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Router configuration needs to be done. | I took access to the router and completed the configuration so now the data is live on the CPCB server | Data is now live hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23062283408 | 2023-07-26 | 2023-07-28 | Cust 1239 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data uploaded issue. | Data is not coming till the system (checked on MODSCAN), as of now we are closing this complaint | Data is not coming till the system (checked on MODSCAN), as of now we are closing this complaint | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23062353409 | 2023-07-26 | 2023-07-26 | Cust 663 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | PFA VISIT REPORT | PFA VISIT DONE | PFA VISIT | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23062723410 | 2023-07-27 | 2023-07-27 | Cust 1242 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC visit for MPA | AMC VISIT FOR MPA DONE | AMC VISIT FOR MPA DONE | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN23062923411 | 2023-07-27 | 2023-07-28 | Cust 1237 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Data is not uploaded on Ecolyser server. | There is an issue from the customer network end so we have asked them to allow url from there network hence as of now we are closing this complaint., | There is an issue from the customer network end so we have asked them to allow url from there network hence as of now we are closing this complaint., | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23062973412 | 2023-07-27 | 2023-07-27 | Cust 526 | PUNE | Analytical - CHAKAN | CLOSED | APPLICATION ISSUE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(Ph-Trans-Smartpro) | PH meter not working issue. | Guided customer for ph meter issue on phone call,,,now problem is solved.ph meter is working properly | PH METER WORKING PROPRLY....GUIDED TO CUSTOMER ON CALL | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFPN23063583413 | 2023-07-28 | 2023-07-28 | Cust 1243 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Data is not uploaded on the Ecolyser server. | We took access of the BB and made some changes form the hardware side so now the data is live on the server. | The issue has been resolved hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23066663414 | 2023-08-01 | 2023-08-01 | Cust 1035 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera streaming issue on CPCB server. | Successfully did the camera configuration and now camera is streaming live on CPCB server. | Successfully did the camera configuration and now camera is streaming live on CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23067443415 | 2023-08-02 | 2023-08-06 | Cust 1239 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | data communication card | issue obsaerved in RS485 Card,,faulty.....,need to change with a new one | visit done for MPA...issue observed in RS485 card faulty....need to change with a new one | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23070013416 | 2023-08-07 | 2023-08-09 | Cust 811 | PUNE | Analytical - CHAKAN | CLOSED | DIGITAL CONEECTIVITY | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Cloud server connectivity needs to be done. | Cloud server connectivity has been completed and data is live. | Cloud server connectivity has been completed and data is live. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23070183417 | 2023-08-08 | 2023-08-10 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | data not updated on cpcb /kspcb server | data is continuously updating on the KSPCB server.Also , data is also updating on the CPCB server and it will take some time as data is heavy . | data is continuously updating on the KSPCB server.Also , data is also updating on the CPCB server and it will take some time as data is heavy . | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23070313418 | 2023-08-08 | 2023-08-14 | Cust 1245 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation needs to be done. | As of now we are closing this complaint, we will reopen this after the system access provided by customer | As of now we are closing this complaint, we will reopen this after the system access provided by customer | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23071273419 | 2023-08-10 | 2023-09-04 | Cust 811 | PUNE | Analytical - CHAKAN | CLOSED | DIGITAL CONEECTIVITY | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | CONNECTIVITY VISIT | CONNECTIVITY VISIT to fm cloud | data connectivity visit to fm cloud | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23071313420 | 2023-08-10 | 2023-08-11 | Cust 692 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC VISIT | AMC VISIT | AMC VISIT DONE | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN23074923421 | 2023-08-18 | 2023-08-21 | Cust 1245 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation needs to be done | Data is live on CPCB server, but in site the electricity issue occurs, hence as of now we are closing this complaint, we will reopen it after the same issue is resolves. | Data is live on CPCB server, but in site the electricity issue occurs, hence as of now we are closing this complaint, we will reopen it after the same issue is resolves. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23074993422 | 2023-08-18 | 2023-09-09 | Cust 1246 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Chlorine sensor | Chlorine sensor-Krohne(1) | received call from Swapnil Wanjari , Prasahant Kadu and Shantanu , to visit to plan the commissioning of the modbus data communication with RIL system integration. Free chlorine sensor & Mag IFC 050 modbus communication demo | Visited the site modbus communication of chlorine DPI and flow meter explained to the custoemr and practically shown | Visited the site modbus communication of chlorine DPI and flow meter explained to the custoemr and practically shown | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23076283423 | 2023-08-21 | 2023-08-21 | Cust 663 | PUNE | Analytical - CHAKAN | CLOSED | AMC | Ph Tranmitter Smartpro | ORP transmitter(1) | COMISSIONING | COMISSIONING VISIT | COMISSIONING VISIT | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23076773424 | 2023-08-21 | 2023-08-23 | Cust 1243 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | MPCB connectivity needs to be done. | We took access of the BB and completed the MPCB connectivity and data is live on the server. | We took access of the BB and completed the MPCB connectivity and data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23076953425 | 2023-08-22 | 2023-08-24 | Cust 1237 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | on the Ecolyser server data is not uploading. | The issue has been resolved hence we are closing this complaint. | The issue has been resolved hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23076973426 | 2023-08-22 | 2023-08-23 | Cust 1232 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | DATCON | DATCON(DATCON) | MPCB data uploaded issue. | Here not found any issue as the data is live on the server. | Here not found any issue as the data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23077503427 | 2023-08-23 | 2023-09-10 | Cust 803 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | received mail from Diageo to plan visit for AMC activity | received mail from Diageo to plan visit for AMC activity | AMC VISIT | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN23077533428 | 2023-08-23 | 2023-09-09 | Cust 795 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | received mail from Mr. Sandip Shinde to attend the MP problem under AMC | Amar visited the site AMC activity completed succesfully | Amar visited the site AMC activity completed succesfully | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN23077603429 | 2023-08-23 | 2023-08-25 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue. | Ok, Please share the access to your system once the hardware issue is resolved and ready to connect.As of now we are closing this complaint. | Ok, Please share the access to your system once the hardware issue is resolved and ready to connect.As of now we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23077723430 | 2023-08-23 | 2023-08-27 | Cust 795 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration visit for MPA | Calibration visit for MPA | Calibration visit done | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23077733431 | 2023-08-23 | 2023-08-25 | Cust 1247 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data uploaded issue. | As of now we are closing this complaint we will reopen it after hardware issue is resolves from customers end | As of now we are closing this complaint we will reopen it after hardware issue is resolves from customers end | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23077753432 | 2023-08-23 | 2023-08-24 | Cust 1237 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Data is not uploaded on the Ecolyser server. | The issue has been resolved hence we are closing this complaint. | The issue has been resolved hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23078683433 | 2023-08-25 | 2023-08-28 | Cust 904 | PUNE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | CPU with LCD of Multiparameter Display(1) | Installation of new display unit | Installation and configuration has been done. | CLOSED | YES | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN23078683433 | 2023-08-25 | 2023-08-28 | Cust 904 | PUNE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Aqua4Trans | RS485 card- of Aqua4Trans(1) | Reconnection of the transmitter. | Reconnection and modbus testing has been done. All the necessary parameter setting has also done. | CLOSED | NO | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN23078703434 | 2023-08-25 | 2023-08-28 | Cust 904 | PUNE | Analytical - CHAKAN | CLOSED | DIGITAL CONEECTIVITY | MULTIPARAMETER ANALYSER | RS485 card-FM of Multiaparameter (1) | Digital Connectivity of MPA | Digital connectivity of MPA has been done . All the parameters started Uploading on CPCB portal. | CLOSED | NO | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN23078773435 | 2023-08-25 | 2023-08-25 | Cust 803 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | data upload issue on CPCB/MPCB server | We took access to the system and checked. COM port was changed , so it is updated on the portal. | We took access to the system and checked. COM port was changed , so it is updated on the portal. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23079263436 | 2023-08-26 | 2023-08-27 | Cust 811 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Multiparameter AMC visit | amc visit done | AMC Visit done | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23079453437 | 2023-08-27 | 2023-08-28 | Cust 1248 | PUNE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration and troubleshooting | Zero Calibration has been done. Tubes arranged from local market and analyser working. | CLOSED | NO | Paid | Process | N |  | Analytical - CHAKAN
AFPN23080103438 | 2023-08-28 | 2023-09-10 | Cust 1249 | PUNE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Tranmitter Smartpro | Conductivity Electrode Cable(Cond-Elec-cbl) | conductivity issue | conductivity issue | VISIT DONE | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23080123439 | 2023-08-28 | 2023-08-29 | Cust 1035 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB + local new device addition. | Now industry is on OFF mode, we will reopen this complaint once the customer is start the industry. | Now industry is on OFF mode, we will reopen this complaint once the customer is start the industry. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23080453440 | 2023-08-28 | 2023-09-04 | Cust 1232 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC visit for multiparameter analyzer | AMC visit for multiparameter analyzer | visit done | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23081563441 | 2023-08-30 | 2023-09-11 | Cust 1250 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | received call from AT to plan visit to check the online monitoring system performance and proposal for new AMC | received call from AT to plan visit to check the online monitoring system performance and proposal for new AMC | VISIT DONE | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23082113442 | 2023-08-31 | 2023-09-12 | Cust 1237 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Data mismatch issue on Ecolyser server. | The issue was resolved by the Ecolyser team and now the correct readings are uploading. | The issue was resolved by the Ecolyser team and now the correct readings are uploading. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23082433443 | 2023-08-31 | 2023-09-04 | Cust 1237 | PUNE | Analytical - CHAKAN | CLOSED | DIGITAL CONEECTIVITY | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | CLOUD CONNECTIVITY VISIT | planning to visit to check data on modscan | DATA UPDATING ISSUE....planning to visit in some days to check data from hardware side | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFPN23083093444 | 2023-09-01 | 2023-09-04 | Cust 1232 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | DATCON | DATCON(DATCON) | CPCB and MPCB data uploaded issue. | The issue has been resolved hence we are closing this complaint. | The issue has been resolved hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23083913445 | 2023-09-04 | 2023-09-10 | Cust 803 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Maintenance under AMC visit | Maintenance under AMC visit | VISIT DONE | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN23083963446 | 2023-09-04 | 2023-09-07 | Cust 1251 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | mpcb login issue on eversense pc. | commercially not clear | commercially not clear | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23085253447 | 2023-09-07 | 2023-09-08 | Cust 1252 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data uploaded issue. | We took the access to the system and now data is live on CPCB server | We took the access to the system and now data is live on CPCB server | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23087233448 | 2023-09-11 | 2023-09-17 | Cust 1239 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | SPARE REPLACEMENT VISIT | SPARE REPLACEMENT VISIT | REPLACEMENT OF RS485 CARD FOR COMMUNICATION | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23087583449 | 2023-09-12 | 2023-09-13 | Cust 1239 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue. | Took access of the system and changed from RS485 to scanner and check the data is live and correct. | Took access of the system and changed from RS485 to scanner and check the data is live and correct. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23088033450 | 2023-09-13 | 2023-09-17 | Cust 1239 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | MPA CHECK VISIT | MPA CHECK VISIT | VISIT DONE | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23088113451 | 2023-09-13 | 2023-09-15 | Cust 803 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software crashed, need to reinstall software. | We took the access to the system and successfully did the software reinstallation + local + CPCB + MPCB configuration, and now data is liv eon all the servers. | We took the access to the system and successfully did the software reinstallation + local + CPCB + MPCB configuration, and now data is liv eon all the servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23088793452 | 2023-09-14 | 2023-09-23 | Cust 1237 | PUNE | Analytical - CHAKAN | CLOSED | DIGITAL CONEECTIVITY | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | connectivity visit | IP ADDRESS ISSUE FROM CUSTOMER SIDE...... | VISIT DONE..BUT IP ADRESS ISSUEFROM CUSTOMER SIDE | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23089033453 | 2023-09-15 | 2023-09-18 | Cust 1245 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation needs to be done | As of now, we are closing this complaint. We will reopen this, once you provide the correct PC name from the customer end | As of now, we are closing this complaint. We will reopen this, once you provide the correct PC name from the customer end | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23089213454 | 2023-09-16 | 2023-09-17 | Cust 1253 | PUNE | Analytical - CHAKAN | CLOSED | APPLICATION ISSUE | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | OXYGEN READING NOT SHOWING IN CETP WATER | DO SENSOR CHECKED ON AIR ,ON CLEAN WATER SHOWS READING...BUT NOT SHOWS ZERO READING ON CUSTOMERS AERATION CETP TANK | DO SENSOR CHECK VISIT | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23089443455 | 2023-09-16 | 2023-09-21 | Cust 1237 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Ecolyser data uploaded issue for ENDURANCE TECHNOLOGIES LIMITED-E92 | As of now, we are closing this and will reopen this case once the network issue is resolved from your side. | As of now, we are closing this and will reopen this case once the network issue is resolved from your side. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23090343456 | 2023-09-18 | 2023-09-21 | Cust 803 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB + CPCB Data upload issue. | As of now we are closing this complaint and we will reopen it after the hardware issue is resolved from customers end | As of now we are closing this complaint and we will reopen it after the hardware issue is resolved from customers end | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23090363457 | 2023-09-18 | 2023-09-18 | Cust 497 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data upload issue. | We took the access to the system and successfully did the software reinstallation + local + CPCB + MPCB configuration, and now data is live on all the servers. | We took the access to the system and successfully did the software reinstallation + local + CPCB + MPCB configuration, and now data is live on all the servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23091033458 | 2023-09-20 | 2023-09-22 | Cust 651 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation needs to be done. | successfully reinstalled the Software and data is live on LOCAL+CPCB+MSPCB | successfully reinstalled the Software and data is live on LOCAL+CPCB+MSPCB | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23091433459 | 2023-09-20 | 2023-09-23 | Cust 1239 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | MULTIPARAMETER SPARE REPLACEMENT VISIT MULTIPARAMETER SPARE REPLACEMENT VISIT..... MULTIPARAMETER WORKING FROM 2014 FREQUENTLY GETTING PROBLEMS,,, RECOMMENDED FOR NEW ADVANCED MULTIPARAMETER ANALYSR TO REPLACE, | MULTIPARAMETER SPARE REPLACEMENT VISIT..... MULTIPARAMETER WORKING FROM 2014 FREQUENTLY GETTING PROBLEMS,,, RECOMMENDED FOR NEW ADVANCED MULTIPARAMETER ANALYSR TO REPLACE, | MULTIPARAMETER SPARE REPLACEMENT VISIT..... MULTIPARAMETER WORKING FROM 2014 FREQUENTLY GETTING PROBLEMS,,, RECOMMENDED FOR NEW ADVANCED MULTIPARAMETER ANALYSR TO REPLACE, | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23091983460 | 2023-09-22 | 2023-09-25 | Cust 731 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue. | We took the access to the system and found that the data is not coming till the local system, hence as of now we are closing this complaint. | We took the access to the system and found that the data is not coming till the local system, hence as of now we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23092063461 | 2023-09-22 | 2023-09-22 | Cust 1239 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Scanner to RS485 configuration needs to change. | Took access of the system and successfully completed the scanner to RS485 configuration. | Took access of the system and successfully completed the scanner to RS485 configuration. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23092243462 | 2023-09-22 | 2023-10-03 | Cust 1237 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CETP data uploaded issue. | we took access of the BB and again we found the Ecolyser server URL was blocked in the customer network hence their IT team is working on it and will reopen this case once the network issue has been resolved from the customer's end. | here we found the Ecolyser server URL was blocked in the customer network hence their IT team is working on it and will reopen this case once the network issue has been resolved. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23092913463 | 2023-09-25 | 2023-10-13 | Cust 795 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | PC issue not turning ON...need software installation in new PC. | Took access of the system and completed software reinstallation data is live now on CPCB+MPCB+LOCAL | Took access of the system and completed software reinstallation data is live now on CPCB+MPCB+LOCAL | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23093083464 | 2023-09-25 | 2023-09-25 | Cust 651 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera Configuration as new camera was installed. | Camera configuration completed and camera is live now on CPCB server. | Camera configuration completed and camera is live now on CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23094393465 | 2023-09-27 | 2023-09-28 | Cust 1254 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data upload issue. | we took the access to the system and after restated enviroconnect services data is live on cpcb server. | we took the access to the system and after restated enviroconnect services data is live on cpcb server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23094533466 | 2023-09-27 | 2023-09-29 | Cust 1035 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue. Ruchika | Now hardware issue is resolves from customers end and now data is live on CPCB server. | Now hardware issue is resolves from customers end and now data is live on CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23096013467 | 2023-09-29 | 2023-09-29 | Cust 936 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | PC C drive memory is frequently showing storage is full because of our ENVIRO CONNECT software issue | We took access of the system and deleted the unwanted files and suggested customer to increase the c drive space to avoid such issues in future. | We took access of the system and deleted the unwanted files and suggested customer to increase the c drive space to avoid such issues in future. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23097013468 | 2023-09-30 | 2023-10-03 | Cust 497 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | ETP Data not updating on MPCB site--Data breaking on MPCB Server. | Took access of the system and updated patch 50 customers system. | Took access of the system and updated patch 50 customers system. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23099213469 | 2023-10-03 | 2023-10-03 | Cust 1255 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not streaming on CPCB server. | Took access of the system and checked the camera configuration. | Took access of the system and checked the camera configuration. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23099263470 | 2023-10-03 | 2023-10-05 | Cust 904 | PUNE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | LCD of Multiparameter Display(1) | Installation , commission of new display and calibration of analyzer | Installation of the display has been done. Zero calibration has been done. | Closed | NO | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN23099533471 | 2023-10-04 | 2023-10-13 | Cust 517 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software Installation in new PC. | Took access of the system and data is live on LOCAL+ CPCB.and for MPCB Configuration port and URL is block from customer side once port unblock will reopen the complaint. | Took access of the system and data is live on LOCAL+ CPCB.and for MPCB Configuration port and URL is block from customer side once port unblock will reopen the complaint. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23100113472 | 2023-10-05 | 2023-10-13 | Cust 517 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Maintenance visit under AMC and software re-installation | AMC visit done. Necessary maintenance activities carried out for Mutlipara. | AMC visit done. Necessary maintenance activities carried out for Mutlipara. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN23100333473 | 2023-10-05 | 2023-10-09 | Cust 904 | PUNE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Display of Smartpro(1) | Keys of Smart pro transmitter are not working. | Suggested to send Chakan factory for further analysis. Also provided proper address and point of contact to customer as well. | Closed | NO | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN23100343474 | 2023-10-05 | 2023-10-09 | Cust 904 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | pH input card-FM of Multiparameter(1) | MPH DO card failed. | Offer sent for MPH DO card. | Closed | NO | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN23100403475 | 2023-10-06 | 2023-10-06 | Cust 1245 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Reinstallation of software | Software reinstallation and Local + CPCB + MPCB connectivity. | Software reinstallation and Local + CPCB + MPCB connectivity. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23102613476 | 2023-10-10 | 2023-10-16 | Cust 803 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | DATA not updating through Modscan app | DATA not updating through Modscan app requested to send DDU unit & Multiparameter Scanner for repair | VISIT DONE | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23102943477 | 2023-10-11 | 2023-10-16 | Cust 632 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation. PC hard disc damaged and replaced with new one. | Took access of the system and successfully completed reinstallation of the software. Data is live now on Local+CPCB Server.Closing the complaint as of now as KSPCB Connectivity not confirmed by customer. | Took access of the system and successfully completed reinstallation of the software. Data is live now on Local+CPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23103233478 | 2023-10-12 | 2023-10-13 | Cust 623 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | We were facing ETP outlet flow meter communication problem | now the issue has been resolved and data is live on the server. | now the issue has been resolved and data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23103393479 | 2023-10-12 | 2023-10-13 | Cust 1256 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not displaying on CPCB. | Took access of the system and updated camera configuration. | Took access of the system and updated camera configuration. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23103863480 | 2023-10-13 | 2023-10-18 | Cust 1257 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | received call from Customer to attend the problem of the MP analyzer not getting power ON | Remote support given to the customer. Issue observed with RS485 card which loading the power supply. Same removed and made analyser functional. | Remote support given to the customer. Issue observed with RS485 card which loading the power supply. Same removed and made analyser functional. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFPN23105523481 | 2023-10-17 | 2023-10-17 | Cust 1258 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on CPCB. | Took access of the system and added video streaming folder in customers PC. | Took access of the system and added video streaming folder in customers PC. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23105793482 | 2023-10-17 | 2023-10-18 | Cust 1259 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | calibration visit for multiparameter | calibration visit for multiparameter | VISIT DONE | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23105963483 | 2023-10-17 | 2023-11-02 | Cust 1260 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | RS485 CARD for Multiparameter(AWA-RS485) | received call from Chidambar to plan urgent visit for commissioning of the RTY panel and to take the data on FM server and to replace RS485 card | customer postponed the visit due to audit | customer postponed the visit due to audit | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFPN23105973484 | 2023-10-17 | 2023-10-26 | Cust 1253 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | AQUADO Analyzer | Visiferm DO sensor(1) | received call from SMS waluj to attend the problem of the DO analyser | Visited the site and necessary exercise carried out for checking sensor healthiness and same found OK. | Visited the site and necessary exercise carried out for checking sensor healthiness and same found OK. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23106113485 | 2023-10-18 | 2023-10-18 | Cust 517 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB connectivity | MPCB connectivity is completed | MPCB connectivity is completed | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23106253486 | 2023-10-18 | 2023-10-18 | Cust 1261 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Local Host link not opening in customer system. | Took access of the system and changed PC Name. | Took access of the system and changed PC Name. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23106323487 | 2023-10-18 | 2023-10-19 | Cust 904 | PUNE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Analyser was reset and not showing the parameters. | All the parameter setting done and calibration also done. Analyzer taken into line and working fine. | Closed | NO | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN23106343488 | 2023-10-18 | 2023-10-20 | Cust 1262 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB + MPCB data upload issue. | now the issue is resolved and now data is live on MPCB server. | now the issue is resolved and now data is live on MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23106363489 | 2023-10-18 | 2023-10-30 | Cust 1263 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data upload issue. | we did software reinstallation + CPC + MPCB configuration, and now hardware issue is resoled from customers end and data is live on all the server. | we did software reinstallation + CPC + MPCB configuration, and now hardware issue is resoled from customers end and data is live on all the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23106443490 | 2023-10-18 | 2023-10-26 | Cust 1264 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Maintenance under AMC visit | AMC visit done. Necessary maintenance activities carried out for Mutlipara. and dust | AMC visit done. Necessary maintenance activities carried out for Mutlipara. and dust | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN23106913491 | 2023-10-19 | 2023-10-20 | Cust 731 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | New Flow Meter Connectivity for MPCB Portal. | Took access of the system and added new device and installed updated version. | Took access of the system and added new device and installed updated version. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23106983492 | 2023-10-19 | 2023-10-20 | Cust 803 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB online portal connectivity issue | We have checked on the MPCB server and data is continuously reflecting on the same. | We have checked on the MPCB server and data is continuously reflecting on the same. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23107143493 | 2023-10-20 | 2023-10-26 | Cust 603 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | receievd call from Bahubali and AMC po for visit to check the MP online monitoring system | AMC visit done. Necessary maintenance activities carried out for Mutlipara. | AMC visit done. Necessary maintenance activities carried out for Mutlipara. | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN23108623494 | 2023-10-25 | 2023-10-26 | Cust 1191 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | SPM value of stack is not polling to MPCB and CPCB | data is now reflecting on both CPCB and MPCB servers by changing protocol from codel to Modbus. | data is now reflecting on both CPCB and MPCB servers by changing protocol from codel to Modbus. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23108873495 | 2023-10-25 | 2023-11-20 | Cust 633 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Conductivity Tranmitter Smartpro | Conductivity Electrode Cable(Cond-Elec-cbl) | conductivity transmitter ( Aquacon2000 smart pro ) issue at SMC Maldives site | competion of work | completed the work | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFPN23109583496 | 2023-10-26 | 2023-11-22 | Cust 731 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB Data not updating. | Data updating to CPCB MPCB | Data updating to CPCB MPCB | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23109793497 | 2023-10-27 | 2023-10-30 | Cust 1265 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning of CX2000 on demo basis | Commission done. Instrument has been installed on demo basis. | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN23109853498 | 2023-10-27 | 2023-10-30 | Cust 1256 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB connectivity | Completed the CPCB connectivity. Data is live on the CPCB server | Completed the CPCB connectivity. Data is live on the CPCB server | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23109893499 | 2023-10-27 | 2023-10-29 | Cust 954 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | LCD of Multiparameter Display(1) | NEW DISPLAY GIVEN,,,,NOT GETTING ON AFTER GIVING 12 V directly .....purple light in display and D3 light gets on & off in display back side | Suggested to send MPA to Chakan Factory to check | display not getting ON so Suggsted to send MPA to Chakan factory to check | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23110003500 | 2023-10-27 | 2023-10-31 | Cust 793 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB connectivity | As of now we are closing this complaint and will reopen this case once we receive the internet details from your end. | As of now we are closing this complaint and will reopen this case once we receive the internet details from your end. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23111423501 | 2023-10-29 | 2023-11-07 | Cust 1266 | PUNE | Codel Products - CHAKAN | CLOSED | AMC | DCEM 2100 Dust Monitor | DCEM 2100 Complete Transceiver with cable(DCEM 2100 TSR With Cable) | AMC VISIT | AMC VISIT DONE | VISIT DONE | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN23112173502 | 2023-10-30 | 2023-11-07 | Cust 1267 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Call for commissioning | Call for commissioning | VISIT DONE | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23112833503 | 2023-10-31 | 2023-11-07 | Cust 603 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC | AMC VISIT DONE | visit done | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN23113193504 | 2023-10-31 | 2023-11-15 | Cust 1260 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | FM make RS485 card(1) | commissioning of RTU panel and RS 485 card in MP | visited the site Rs 485 catrd replaced , MP modbus communication completed , data availabe on FM server | visited the site Rs 485 card replaced and issue resolved successfully | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23113253505 | 2023-10-31 | 2023-11-03 | Cust 1256 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not showing on Local, CPCB and MPCB servers. | completed Local + CPCB + MPCB connectivity | completed Local + CPCB + MPCB connectivity | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23113333506 | 2023-10-31 | 2023-11-03 | Cust 603 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB and MPCB data uploaded issue. | As of now, we are closing this complaint and please revert once the hardware issue is resolved from your end. | As of now, we are closing this complaint and please revert once the hardware issue is resolved from your end. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23114373507 | 2023-11-01 | 2023-11-03 | Cust 1035 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation needs to be done. | We did software reinstallation + CPCB + MPCB configuration but currently data is not communicating with the local system, hence as of now we are closing this complaint. | We did software reinstallation + CPCB + MPCB configuration but currently data is not communicating with the local system, hence as of now we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23114443508 | 2023-11-01 | 2023-11-01 | Cust 603 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB KSPCB Server. | We took access of the system and data is live on CPCB and KSPCB Server. | We took access of the system and data is live on CPCB and KSPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23114723509 | 2023-11-01 | 2023-11-01 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue. | Issue is resolve from hardware side and now data is live on CPCB server. | Issue is resolve from hardware side and now data is live on CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23114943510 | 2023-11-02 | 2023-11-05 | Cust 1268 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | MPA ISSUE | MPA ISSUE SOLVED | visit done | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN23114953511 | 2023-11-02 | 2023-11-07 | Cust 1269 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Silica Analyser | Silica 8891(8891) | received mail from customer to plan the visit for commissioning the silica analyzer | visited the site checked all the components and requirments of reagets of silica discussed , customer to arrange DM water , and FM will arrange reagents | visited the site checked all the components and requirments of reagets of silica discussed , customer to arrange DM water , and FM will arrange reagents | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23115003512 | 2023-11-02 | 2023-11-03 | Cust 1270 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB and MPCB data uploded issue | As of now we are closing this complaint and will reopen once we get access of the system from the customer end. | As of now we are closing this complaint and will reopen once we get access of the system from the customer end. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23115033513 | 2023-11-02 | 2023-11-03 | Cust 603 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data uploaded issue. | now the Local Host is accessible and currently the analyser is in off condition so the data will not be uploading. | now the Local Host is accessible and currently the analyser is in off condition so the data will not be uploading. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23115143514 | 2023-11-02 | 2023-11-03 | Cust 1271 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation | We have reinstalled the software with the updated version and completed the Local + CPCB + KSPCB connectivity. | We have reinstalled the software with the updated version and completed the Local + CPCB + KSPCB connectivity. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23115513515 | 2023-11-03 | 2023-11-06 | Cust 1035 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB + CPCB data upload issue. | now data is live on CPCB + MPCB server, hence we are closing this complaint. | now data is live on CPCB + MPCB server, hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23115733516 | 2023-11-04 | 2023-11-05 | Cust 1272 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | servicing visit for MPA ,,,,, | servicing visit for MPA ,,,,, | visit done | YES | Paid | Core | N |  | Analytical - CHAKAN
AFPN23116293517 | 2023-11-06 | 2023-11-08 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation needs to be done. | We did software reinstallation + CPCB + MPCB configuration and now data is live on all the server. | We did software reinstallation + CPCB + MPCB configuration and now data is live on all the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23116303518 | 2023-11-06 | 2023-11-09 | Cust 1035 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation needs to be done. | As of now we are closing this complaint, we will reopen it after the internet issue is resolved from customers end. | As of now we are closing this complaint, we will reopen it after the internet issue is resolved from customers end. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23116593519 | 2023-11-06 | 2023-11-09 | Cust 1273 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation | We have reinstalled the software and completed the Local + CPCB + MPCB connectivity. Currently data is not reflecting on servers as the USB converter is not connected to the PC. | We have reinstalled the software and completed the Local + CPCB + MPCB connectivity. Currently data is not reflecting on servers as the USB converter is not connected to the PC. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23116633520 | 2023-11-06 | 2023-11-12 | Cust 1270 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | Current output card-FM(1) | current out put card problem... | current out put card problem... Replaced with a new one | AMC VISIT | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN23116833521 | 2023-11-07 | 2023-12-08 | Cust 1274 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | DO Transmitters | VISI pro DO(VISI pro DO) | Electrode not responding | Site attended with new electrode. Response found satisfactory | Site attended with new electrode. Response found satisfactory | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFPN23116883522 | 2023-11-07 | 2023-11-08 | Cust 1271 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | data not showing in RTDMS dash board | issue has been resolved and data is reflecting on the central portal | issue has been resolved and data is reflecting on the central portal | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23117043523 | 2023-11-07 | 2023-11-09 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | CPCB + MPCB data upload issue. | Successfully did software reinstallation + CPCB + MPCB configuration and now data is live on both the servers. | Successfully did software reinstallation + CPCB + MPCB configuration and now data is live on both the servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23117063524 | 2023-11-07 | 2023-11-20 | Cust 1269 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Silica Analyser | Silica 9210(9210) | Commissioning of AWA- SILICA under FAT | completion of work | completed the work | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23117063524 | 2023-11-07 | 2023-11-20 | Cust 1269 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Silica Analyser | Silica 9210(9210) | Commissioning of AWA- SILICA under FAT | Completion of work | completed the work | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23117113525 | 2023-11-07 | 2023-11-09 | Cust 1275 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB +MPCB data upload issue. | Took access to the system and changed the configuration and now same data is displaying on CPCB+LOCAL+MPCB. | Took access to the system and changed the configuration and now same data is displaying on CPCB+LOCAL+MPCB. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23117123526 | 2023-11-07 | 2023-11-07 | Cust 1261 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera streaming issue. | Camera streaming is now live on the CPCB server. | Camera streaming is now live on the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23117303527 | 2023-11-08 | 2023-11-18 | Cust 1240 | PUNE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Calibration of 02 Nos. PH sensor with smart pro transmitter. | Calibration has been done successfully. | Closed | NO | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN23117303527 | 2023-11-08 | 2023-11-18 | Cust 1240 | PUNE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Calibration of 02 Nos. PH sensor with smart pro transmitter. | Calibration has been done successfully. | Closed | NO | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN23117553528 | 2023-11-08 | 2023-11-09 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB + MPCB data upload issue. | Successfully did software reinstallation + CPCB + MPCB configuration and now data is live on both the servers. | Successfully did software reinstallation + CPCB + MPCB configuration and now data is live on both the servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23117563529 | 2023-11-08 | 2023-11-10 | Cust 1035 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB + MPCB data data uploading issue. | we took the access to your system and observed that the USB converter is still not connected to the PC as the com port is not detecting, hence as of now we are closing this complaint. | we took the access to your system and observed that the USB converter is still not connected to the PC as the com port is not detecting, hence as of now we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23117603530 | 2023-11-08 | 2023-11-09 | Cust 1035 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data mismatch with actual reading | We took the access to the system and checked the data on MODSCAN and as per calculations the same data is reflecting onto the CPCB server. | We took the access to the system and checked the data on MODSCAN and as per calculations the same data is reflecting onto the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23118203531 | 2023-11-09 | 2023-11-09 | Cust 731 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on CPCB server for Incineration Boiler DCS Through Device. | We took access of the system and camera is live on CPCB Server. | We took access of the system and camera is live on CPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23118333532 | 2023-11-09 | 2023-11-20 | Cust 1256 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC | completion of work | completed the work | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN23118343533 | 2023-11-09 | 2023-11-12 | Cust 1276 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC | amc visit for MPA | VISIT DONE | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN23118373534 | 2023-11-09 | 2023-11-09 | Cust 1276 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB and MPCB Server. | We took access of the system and data is live on CPCB and MPCB Server. | We took access of the system and data is live on CPCB and MPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23118393535 | 2023-11-09 | 2023-11-10 | Cust 1035 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB + MPCB data upload issue. | We took the access to your system and checked but data is not communicating with the local system (checked on MODSCAN), hence as of now we will close this complaint. | We took the access to your system and checked but data is not communicating with the local system (checked on MODSCAN), hence as of now we will close this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23118583536 | 2023-11-10 | 2023-11-11 | Cust 1001 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data mismatch issue. | we took the access to the system and checked, but data is not coming (checked on MODSCAN), so as of now we are closing this complaint. | we took the access to the system and checked, but data is not coming (checked on MODSCAN), so as of now we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23118763537 | 2023-11-10 | 2023-12-08 | Cust 1035 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | received PO from Suyog to attend the problem of the MP analyser, not working properly. | visit done. Necessary maintenance activities carried out for Mutlipara. and Dust | visit done. Necessary maintenance activities carried out for Mutlipara. and Dust | NO | Paid | Core | N |  | Analytical - CHAKAN
AFPN23118773538 | 2023-11-10 | 2023-11-12 | Cust 731 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | zenon power supply new replacement visit | zenon power supply new replacement visit | VISIT DONE | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23118783539 | 2023-11-10 | 2023-11-12 | Cust 1270 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue. | We took the access to your system but stack data is still not communicating with the local system (checked on MODSCAN), hence as of now we are closing this complaint. | We took the access to your system but stack data is still not communicating with the local system (checked on MODSCAN), hence as of now we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23120463540 | 2023-11-16 | 2023-11-18 | Cust 1277 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB + MPCB data upload issue. | now the COM port is detecting and data is live on CPCB + MPCB server. | now the COM port is detecting and data is live on CPCB + MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23120483541 | 2023-11-16 | 2023-12-15 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation | Data is live on CPCB and MPCB servers. | Data is live on CPCB and MPCB servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23120553542 | 2023-11-16 | 2023-11-18 | Cust 1278 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data uploaded issue. | now the issue has been resolved and data is live on the server. | now the issue has been resolved and data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23120823543 | 2023-11-17 | 2023-11-18 | Cust 1035 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation + CPCB + MPCB configuration needs to be done. | here is the issue is from hardware side, so as of now we are closing this complaint. | here is the issue is from hardware side, so as of now we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23120833544 | 2023-11-17 | 2023-11-18 | Cust 1035 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue. | as of now we are closing this complaint. | as of now we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23120843545 | 2023-11-17 | 2023-12-12 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue. | As of now we are cloisng this complaint as the data is live on the server. | As of now we are cloisng this complaint as the data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23120933546 | 2023-11-17 | 2023-11-18 | Cust 1256 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data mismatch issue. | We took the access to your system and checked the data on MODSCAN and as per calculations the same data is reflected on the CPCB server. | We took the access to your system and checked the data on MODSCAN and as per calculations the same data is reflected on the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23121163547 | 2023-11-18 | 2023-11-18 | Cust 1279 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB and MPCB data uploded issue. | now the data is live on the server. | now the data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23121183548 | 2023-11-18 | 2023-11-21 | Cust 1237 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Data Mismatch issue in Ecolyser server. | Now the issue has been resolved | Now the issue has been resolved | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23121873549 | 2023-11-20 | 2023-11-23 | Cust 1233 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | data not uploading to MPCB and CPCB servers | As of now we are closing this complaint. We will reopen this once the issue is resolved from the hardware side. | As of now we are closing this complaint. We will reopen this once the issue is resolved from the hardware side. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23121903550 | 2023-11-20 | 2023-11-21 | Cust 1280 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data upload issue | Data is now live on the MPCB server. | Data is now live on the MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23122133551 | 2023-11-20 | 2023-11-21 | Cust 1261 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera streaming is not live on the CPCB | Camera streaming is now live on the CPCB server. | Camera streaming is now live on the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23122143552 | 2023-11-20 | 2023-11-21 | Cust 1281 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc | completed the work | completed the work | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN23122203553 | 2023-11-20 | 2023-11-23 | Cust 603 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB and MPCB data uploaded issue. | Now the data is live on the server. | Now the data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23122523554 | 2023-11-21 | 2023-11-24 | Cust 1277 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Stack data not display on CPCB | There is issue from hardware side so as of now we are closing this complaint. | There is issue from hardware side so as of now we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23122793555 | 2023-11-21 | 2023-12-04 | Cust 1232 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | DATCON | DATCON(DATCON) | CPCB and MPCB connectivity issue. | Here we found internet issue from the coustomer end and curretky the data is live on the server so as of now we are closing this ompliant. | Here we found internet issue from the coustomer end and curretky the data is live on the server so as of now we are closing this ompliant. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23122803556 | 2023-11-21 | 2023-11-23 | Cust 603 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB and MPCB data uploaded issue. | Today we have checked and the data is continuously uploading on both servers. | Today we have checked and the data is continuously uploading on both servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23123033557 | 2023-11-22 | 2023-11-26 | Cust 623 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | SODIUM ANALYSER | SODIUM 9073 single channel ANLAYSER(9073 Sodium) | commissioning of analyser | completion of work | completed the work | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23123033557 | 2023-11-22 | 2023-11-26 | Cust 623 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Silica Analyser | Silica 9210(9210) | commissioning for silica analyser make-swan | completion of work | completed the work | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23123073558 | 2023-11-22 | 2023-11-23 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | FLOW READINGS NOT SHOWING IN MPCB SERVER.BUT CORRECTLY VALUES OF FLOW SHOWN ON CPCB SERVER. | Now, flow data is reflected on the MPCB as well as the CPCB server. | Now, flow data is reflected on the MPCB as well as the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23123153559 | 2023-11-22 | 2023-11-23 | Cust 1261 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on MPCB server. | Data is now live on the MPCB server. | Data is now live on the MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23123553560 | 2023-11-23 | 2023-12-08 | Cust 1282 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | received AMC po and payment from customer , need to plan for commissioning of MP and software | AMC visit done. Necessary maintenance activities carried out for Mutlipara. | AMC visit done. Necessary maintenance activities carried out for Mutlipara. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN23123713561 | 2023-11-23 | 2023-11-25 | Cust 1283 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating to MPCB Server. | Data is live now on LOCAL+CPCB+MPCB. | Data is live now on LOCAL+CPCB+MPCB. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23123763562 | 2023-11-23 | 2023-11-23 | Cust 1284 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | data is not updating in MPCB portal | data is now live on the MPCB server. | data is now live on the MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23123783563 | 2023-11-23 | 2023-11-24 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation + Local + CPCB + MPCB connectivity | We have reinstalled the software and completed the Local + CPCB + MPCB connectivity and data is now live on the same. | We have reinstalled the software and completed the Local + CPCB + MPCB connectivity and data is now live on the same. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23124013564 | 2023-11-24 | 2023-11-27 | Cust 809 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB + MPCB data upload issue. | We took access to the system and checked that some CPCB files were deleted. Right now the issue is resolved and data is live | We took access to the system and checked that some CPCB files were deleted. Right now the issue is resolved and data is live | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23124093565 | 2023-11-24 | 2023-11-27 | Cust 1285 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB/KSPCB connectivity issue | We have reinstalled the software and completed the Local + CPCB + KSPCB connectivity and data is now live on the same. | We have reinstalled the software and completed the Local + CPCB + KSPCB connectivity and data is now live on the same. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23124263566 | 2023-11-24 | 2023-11-25 | Cust 1284 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | STACK2_40TPH BOILER data not showing on CPCB and MPCB portals | Data from STACK1_120TPH BOILER and STACK2_40TPH BOILER is now live on the Local + CPCB + MPCB servers. | Data from STACK1_120TPH BOILER and STACK2_40TPH BOILER is now live on the Local + CPCB + MPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23124563567 | 2023-11-25 | 2023-11-28 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Reconfiguration of software -PIYU SYSTEM (Shri Vitthal SSK Ltd-SUGAR PLANT) | Took access to the system and completed reconfiguration now data is live on MPCB+CPCB+LOCAL. | Took access to the system and completed reconfiguration now data is live on MPCB+CPCB+LOCAL. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23125623568 | 2023-11-27 | 2023-11-28 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on MPCB.==Siddhanath (Piyu system) | Took access to the system and completed reconfiguration now data is live on MPCB+CPCB+LOCAL. | Took access to the system and completed reconfiguration now data is live on MPCB+CPCB+LOCAL. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFPN23125633569 | 2023-11-27 | 2023-11-30 | Cust 1256 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation needs to be done. | We successfully did software reinstallation + CPCB + MPCB configuration and now data is live on both server. | We successfully did software reinstallation + CPCB + MPCB configuration and now data is live on both server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23125693570 | 2023-11-27 | 2023-11-30 | Cust 1286 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Localhost not accessible, MPCB + CPCB data upload issue. | Now localhost is accessible, and data is live on CPCB + MPCB server. | Now localhost is accessible, and data is live on CPCB + MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23126453571 | 2023-11-27 | 2023-11-30 | Cust 1287 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | DATCON | DATCON(DATCON) | KSPCB data uploaded issue. | The issue has been resolved and data is live on the KSPCB server. | http://cpdms.forbesmarshall.in:8080/enviroconnect | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23126873572 | 2023-11-28 | 2023-11-29 | Cust 1282 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation + Local + CPCB + MPCB connectivity. | We have reinstalled the software and completed the local + CPCB + MPCB connectivity and data is now live on the same servers. | We have reinstalled the software and completed the local + CPCB + MPCB connectivity and data is now live on the same servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23126903573 | 2023-11-28 | 2023-12-01 | Cust 1035 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data upload issue. | as of now we are closing this complaint and will reopen it after hardware issue is resolved from customer end | as of now we are closing this complaint and will reopen it after hardware issue is resolved from customer end | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23126913574 | 2023-11-28 | 2023-11-30 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB + MPCB data upload issue. | We have did change in configuration and now data is live on CPCB + MPCB server. | We have did change in configuration and now data is live on CPCB + MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23126943575 | 2023-11-28 | 2023-12-01 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue | Data is now live on the Local and CPCB server. | Data is now live on the Local and CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23126953576 | 2023-11-28 | 2023-12-04 | Cust 936 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Smartpro-con) | comissioning | completion of work | completed the work | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23126963577 | 2023-11-28 | 2023-12-04 | Cust 651 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | commisioning visit | completion of work | completed the work | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23126973578 | 2023-11-28 | 2023-12-04 | Cust 1274 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | arc DO to electrode | completion of work | completed the work | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23127003579 | 2023-11-28 | 2023-11-30 | Cust 1273 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on MPCB server. | Took access to the system and now data is live on MPCB. | Took access to the system and now data is live on MPCB. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23127033580 | 2023-11-28 | 2023-11-29 | Cust 1288 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Only PM value delay data updating on MPCB. | Took access to the system and did changed time and date and did changes in configuration now data is live on MPCB server. | Took access to the system and did changed time and date and did changes in configuration now data is live on MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23127043581 | 2023-11-28 | 2023-11-30 | Cust 1289 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Time not updating on CPCB and Local and MPCB connectivity. | Took access to the system and data is live on CPCB+MPCB+LOCAL. | Took access to the system and data is live on CPCB+MPCB+LOCAL. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23127143582 | 2023-11-28 | 2023-12-15 | Cust 1290 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | MPCB data uploaded issue., | now the Totalizer Flow values are uploading on the MPCB server. | now the Totalizer Flow values are uploading on the MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23127573583 | 2023-11-29 | 2023-11-30 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data upload issue. | We did some changes from configuration end and now data is live on MPCB server. | We did some changes from configuration end and now data is live on MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23128783584 | 2023-11-30 | 2023-12-01 | Cust 1291 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Flow data mismatch issue | we have checked on modscan and the same data is transmitted to CPCB. | we have checked on modscan and the same data is transmitted to CPCB. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23128913585 | 2023-11-30 | 2023-12-05 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data upload issue. | Data is live on MPCB server, hence we are closing this complaint. | Data is live on MPCB server, hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23128933586 | 2023-11-30 | 2023-12-05 | Cust 1292 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Eversense software reinstallation + CPCB + MPCB configuration needs to be done. | We did CPCB configuration and now data is live on CPCB server. | We did CPCB configuration and now data is live on CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23129953587 | 2023-12-01 | 2023-12-04 | Cust 1264 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data mismatch issue | data mismatch issue has been resolved. | data mismatch issue has been resolved. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23130123588 | 2023-12-01 | 2023-12-04 | Cust 1293 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | data connectivity | completion of work | completed the work | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23130183589 | 2023-12-01 | 2023-12-01 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Reinstallation of software. | Took access to the system and successfully completed installation now data is live on CPCB+MPCB+LOCAL | Took access to the system and successfully completed installation now data is live on CPCB+MPCB+LOCAL | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23130193590 | 2023-12-01 | 2023-12-01 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | PIYU SYSTEM --VITTHAL SUGAR MANUFACTURING LTD. Reinstallation of software. | reinstalled the software and completed Local + CPCB + MPCB connectivity and data is now live on the same servers. | reinstalled the software and completed Local + CPCB + MPCB connectivity and data is now live on the same servers. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFPN23130313591 | 2023-12-02 | 2023-12-08 | Cust 1294 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | ETP MPA DATA NOT UPDATING ON CPDMS SITE | Connectivity issue attended. | Connectivity issue attended. | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN23131073592 | 2023-12-04 | 2023-12-05 | Cust 1275 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on MSPCB and CPCB. | Took access to the system and data is live on CPCB and MPCB. | Took access to the system and data is live on CPCB and MPCB. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23131093593 | 2023-12-04 | 2023-12-05 | Cust 1275 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera streaming window closing and camera not playing on CPCB. | Took access to the system and Camera is live on CPCB. | Took access to the system and Camera is live on CPCB. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23131153594 | 2023-12-04 | 2023-12-06 | Cust 803 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating to MPCB Server. | Took access to the system and data is updating on MPCB Portal. | Took access to the system and data is updating on MPCB Portal. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23131203595 | 2023-12-04 | 2023-12-06 | Cust 1273 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data mismatch on LOCAL+MPCB | Took access to the system and data and correct data is displayed on MPCB and LOCAL. | Took access to the system and data and correct data is displayed on MPCB and LOCAL. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23131243596 | 2023-12-04 | 2023-12-07 | Cust 1295 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software installation + CPCB + MPCB configuration needs to be done | We did software installation + CPCB + MPCB configuration and now data is live on all the servers. | We did software installation + CPCB + MPCB configuration and now data is live on all the servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23131253597 | 2023-12-04 | 2023-12-05 | Cust 1296 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data mismatch issue | Data mismatch issue has been resolved. | Data mismatch issue has been resolved. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23131293598 | 2023-12-04 | 2023-12-06 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | stack data is not updating to server | Data from BOILER_105 TPH is live on the CPCB and MPCB servers. | Data from BOILER_105 TPH is live on the CPCB and MPCB servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23131433599 | 2023-12-04 | 2023-12-07 | Cust 988 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB and MPCB data uploded issue. | Data is now live on both servers. | Data is now live on both servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23131513600 | 2023-12-05 | 2023-12-15 | Cust 1035 | PUNE | Analytical - CHAKAN | CLOSED | SALES PROMOTION | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation needs to be done. | We took the access to the system and did software reinstallation + CPCB + MPCB configuration and now data is live on all the servers. | We took the access to the system and did software reinstallation + CPCB + MPCB configuration and now data is live on all the servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23131633601 | 2023-12-05 | 2023-12-07 | Cust 1245 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Zero data updating for PM on CPCB+LOCAL server. | We took access to the system and and live data is updating for PM on all server. | We took access to the system and and live data is updating for PM on all server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23132073602 | 2023-12-06 | 2023-12-08 | Cust 1289 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not Updating on MPCB Server. | Took access to the system and now data is updating on MPCB Server. | Took access to the system and now data is updating on MPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23132093603 | 2023-12-06 | 2023-12-08 | Cust 731 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Localhost timesync issue | Did some changes in configuration and data is live for all devices on CPCB and MPCB. | Did some changes in configuration and data is live for all devices on CPCB and MPCB. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23132183604 | 2023-12-06 | 2023-12-07 | Cust 1275 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data upload issue. | Data is now live on the MPCB server. | Data is now live on the MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23132523605 | 2023-12-07 | 2023-12-19 | Cust 1297 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Reinstallation of software. | Took access to the system and data is live now on CPCB+LOCAL+MPCB. | Took access to the system and data is live now on CPCB+LOCAL+MPCB. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23132723606 | 2023-12-07 | 2023-12-14 | Cust 1275 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | DATCON | DATCON(DATCON) | CPCB and MPCB connectivity issue | now the Totalizer Flow values are uploading on the MPCB server. | now the Totalizer Flow values are uploading on the MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23132783607 | 2023-12-07 | 2023-12-11 | Cust 992 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB and MPCB data uploded issue. | now the data is live on the server. | now the data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23132953608 | 2023-12-08 | 2023-12-09 | Cust 651 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | annual maintenance | completed the work | completed the work | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN23133003609 | 2023-12-08 | 2023-12-08 | Cust 1277 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue. | Data is live on the CPCB server. | Data is live on the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23133053610 | 2023-12-08 | 2023-12-11 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Reinstallation of software | Software reinstallation and Local configuration completed. As of now we are closing this complaint. We will reopen this once device is connected to PC. | Software reinstallation and Local configuration completed. As of now we are closing this complaint. We will reopen this once device is connected to PC. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23133063611 | 2023-12-08 | 2023-12-08 | Cust 1298 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not live on CPCB. | Camera streaming is live now. | Camera streaming is live now. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23133133612 | 2023-12-08 | 2023-12-15 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software installation + MPCB configuration needs to be done. | We did software installation + local + MPCB configuration, now data is live on all the servers. | We did software installation + local + MPCB configuration, now data is live on all the servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23133153613 | 2023-12-08 | 2023-12-11 | Cust 1035 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | local host not accessible. | As of now we are closing this complaint, we will reopen it after getting proper PC access from customer | As of now we are closing this complaint, we will reopen it after getting proper PC access from customer | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23134143614 | 2023-12-11 | 2023-12-13 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB+LOCAL+MPCB. | Took access to the system and did configuration changes data is live now on CPCB+MPCB+LOCAL | Took access to the system and did configuration changes data is live now on CPCB+MPCB+LOCAL | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23134313615 | 2023-12-11 | 2023-12-14 | Cust 1035 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue for Dust. | As of now we are closing this complaint, and will reopen it after the hardware issue is resolved from customers end. | As of now we are closing this complaint, and will reopen it after the hardware issue is resolved from customers end. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23134413616 | 2023-12-11 | 2023-12-12 | Cust 1279 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not uploading to CPCB, MPCB and local servers | Data is live on CPCB, local and MPCB. | Data is live on CPCB and MPCB. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23134433617 | 2023-12-11 | 2023-12-11 | Cust 1298 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on cpcb | Camera streaming is live now. | Camera streaming is live now. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23134783618 | 2023-12-12 | 2023-12-14 | Cust 1237 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | data was not updating on server | data was not updating on server,,,,,connectivity done through software Team | VISIT DONE | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23134793619 | 2023-12-12 | 2023-12-13 | Cust 1292 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | MPCB data upload issue. | we took the access to your system and did software upgradation (UBUNTU) + CPCB + MPCB configuration. And now data is live on both servers. | we took the access to your system and did software upgradation (UBUNTU) + CPCB + MPCB configuration. And now data is live on both servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23134803620 | 2023-12-12 | 2023-12-15 | Cust 1245 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating for CPCB. | We took access to the system and found data is not updating on MODSCAN. | We took access to the system and found data is not updating on MODSCAN. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23135163621 | 2023-12-13 | 2023-12-14 | Cust 1239 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data uploading issue | Now data is reflecting on the MPCB server. | Now data is reflecting on the MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23135213622 | 2023-12-13 | 2023-12-15 | Cust 1299 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB+LOCAL+MPCB and services not working. | We took access to the system and now data is live on CPCB+MPCB+LOCAL. | We took access to the system and now data is live on CPCB+MPCB+LOCAL. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23135263623 | 2023-12-13 | 2023-12-15 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation | Software installation has been completed and Data is not coming till the modscan so As of now we are closing this complaint and will reopen this case once the hardware issue has been resolved from the customer end. | As of now we are closing this complaint and will reopen this case once the hardware issue has been resolved from the customer end. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23135743624 | 2023-12-14 | 2023-12-15 | Cust 1289 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data was not updating stack1 Gas. | We took access to the system and now Stack1 GAS data is live now. | We took access to the system and now Stack1 GAS data is live now. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23135763625 | 2023-12-14 | 2023-12-15 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | FLOW VALUE ISSUE ON MPCB SERVER | Data from ETP Outlet_Flow is now available on the MPCB server. | Data from ETP Outlet_Flow is now available on the MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23135773626 | 2023-12-14 | 2023-12-17 | Cust 1237 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | Aqua4Trans | Aqua4Trans(1) | FOR ECOLYSER CONNECTIVITY VISIT | FOR ECOLYSER CONNECTIVITY VISIT | VISIT DONE | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23135783627 | 2023-12-14 | 2023-12-17 | Cust 1237 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | ecolyser connectivity visit | connectivity visit | visit done | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23136093628 | 2023-12-14 | 2023-12-15 | Cust 1293 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB and MPCB data upload issue | Data is now live on CPCB server. | Data is now live on CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23136443629 | 2023-12-15 | 2023-12-15 | Cust 1233 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on CPCB Server. | We took access to the system and now camera is live on CPCB Server. | We took access to the system and now camera is live on CPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23136473630 | 2023-12-15 | 2023-12-15 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on MPCB Server. | We took access to the system and data is now updating on MPCB Server. | We took access to the system and data is now updating on MPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23136623631 | 2023-12-15 | 2023-12-15 | Cust 1256 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB connectivity needs to be done. | CPCB Connectivity for Stack_2 has been completed and data is live on the server. | CPCB Connectivity for Stack_2 has been completed and data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23137003632 | 2023-12-16 | 2023-12-17 | Cust 811 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | COD.BOD,TSS READINGS SHOWS ZERO | COD.BOD,TSS READINGS SHOWS ZERO problem Solved | VISIT DONE | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23137863633 | 2023-12-18 | 2023-12-21 | Cust 731 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not updating on CPCB Server. | We took access to the system and Camera is updating on CPCB Server. | We took access to the system and Camera is updating on CPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23138093634 | 2023-12-18 | 2023-12-21 | Cust 1261 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Both Cameras not playing in CPCB Server. | Took access to the system and both cameras are playing on CPCB Server. | Took access to the system and both cameras are playing on CPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23138173635 | 2023-12-18 | 2023-12-20 | Cust 1300 | PUNE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | CPU with LCD of Multiparameter Display(1) | Display Failed at site | Display sent to chakan factory for further analysis. | Closed | NO | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN23138203636 | 2023-12-18 | 2023-12-20 | Cust 1301 | PUNE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | AQUADO Analyzer | Visiwater DO sensor(1) | calibration | completed the work | completed the work | NO | Paid | Core | N |  | Analytical - CHAKAN
AFPN23138203636 | 2023-12-18 | 2023-12-20 | Cust 1301 | PUNE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | AQUADO Analyzer | Visiwater DO sensor(1) | calibration | completed the work | completed the work | NO | Paid | Core | N |  | Analytical - CHAKAN
AFPN23139243637 | 2023-12-20 | 2023-12-25 | Cust 575 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | pH Transmitter Aquamon | Ph Transmitter Aquamon 2000(PH2000) | Commissioning of 1 ORP sensor. | Commissioning has been done successfully. | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN23139243637 | 2023-12-20 | 2023-12-25 | Cust 575 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | pH Transmitter Aquamon | Ph Transmitter Aquamon 2000(PH2000) | Commissioning of 7 Nos. Ph transmitter. | Commissioning has been done successfully. | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN23139243637 | 2023-12-20 | 2023-12-25 | Cust 575 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | pH Transmitter Aquamon | Ph Transmitter Aquamon 2000(PH2000) | Commissioning of 7 Nos. Ph transmitter. | Commissioning has been done successfully. | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN23139243637 | 2023-12-20 | 2023-12-25 | Cust 575 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | pH Transmitter Aquamon | Ph Transmitter Aquamon 2000(PH2000) | Commissioning of 7 Nos. Ph transmitter. | Commissioning has been done successfully. | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN23139243637 | 2023-12-20 | 2023-12-25 | Cust 575 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | pH Transmitter Aquamon | Ph Transmitter Aquamon 2000(PH2000) | Commissioning of 7 Nos. Ph transmitter. | Commissioning has been done successfully. | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN23139243637 | 2023-12-20 | 2023-12-25 | Cust 575 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | pH Transmitter Aquamon | Ph Transmitter Aquamon 2000(PH2000) | Commissioning of 7 Nos. Ph transmitter. | Commissioning has been done successfully. | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN23139243637 | 2023-12-20 | 2023-12-25 | Cust 575 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | pH Transmitter Aquamon | Ph Transmitter Aquamon 2000(PH2000) | Commissioning of 7 Nos. Ph transmitter. | Commissioning has been done successfully. | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN23139243637 | 2023-12-20 | 2023-12-25 | Cust 575 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | pH Transmitter Aquamon | Ph Transmitter Aquamon 2000(PH2000) | Commissioning of 7 Nos. Ph transmitter. | Commissioning has been done successfully. | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN23139243637 | 2023-12-20 | 2023-12-25 | Cust 575 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Transmitter Aquacon | Aquacon 2000(AQCN 2000 p) | Commissioning of Auacon 2000 at RO plant | Commission done successfully. | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN23139243637 | 2023-12-20 | 2023-12-25 | Cust 575 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Transmitter Aquacon | Aquacon 2000(AQCN 2000 p) | Commissioning of Auacon 2000 at RO plant | Commission done successfully. | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN23139243637 | 2023-12-20 | 2023-12-25 | Cust 575 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Transmitter Aquacon | Aquacon 2000(AQCN 2000 p) | Commissioning of Auacon 2000 at RO plant | Commission done successfully. | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN23139243637 | 2023-12-20 | 2023-12-25 | Cust 575 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Transmitter Aquacon | Aquacon 2000(AQCN 2000 p) | Commissioning of Auacon 2000 at RO plant | Commission done successfully. | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN23139243637 | 2023-12-20 | 2023-12-25 | Cust 575 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Transmitter Aquacon | Aquacon 2000(AQCN 2000 p) | Commissioning of Auacon 2000 at RO plant | Commission done successfully. | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN23139243637 | 2023-12-20 | 2023-12-25 | Cust 575 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Transmitter Aquacon | Aquacon 2000(AQCN 2000 p) | Commissioning of Auacon 2000 at RO plant | Commissioning has been done successfully. | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN23139243637 | 2023-12-20 | 2023-12-25 | Cust 575 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Transmitter Aquacon | Aquacon 2000(AQCN 2000 p) | Commissioning of Auacon 2000 at RO plant | Commissioning has been done successfully. | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN23139343638 | 2023-12-20 | 2023-12-22 | Cust 1302 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on MPCB server. | Took access to the system and data is now updating on MPCB SERVER | Took access to the system and data is now updating on MPCB SERVER | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23139433639 | 2023-12-20 | 2023-12-23 | Cust 1303 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Reconfiguration of software. | Took access to the system and successfully completed the reconfiguration and now data is live on CPCB+LOCAL+MPCB | Took access to the system and successfully completed the reconfiguration and now data is live on CPCB+LOCAL+MPCB | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23140283640 | 2023-12-22 | 2023-12-22 | Cust 1035 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation + CPCB +MPCB configuration needs to be done. | we took the access to the system and did software reinstallation + local CPCB configuration and now data is live on CPCB + MPCB server. | we took the access to the system and did software reinstallation + local CPCB configuration and now data is live on CPCB + MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23140313641 | 2023-12-22 | 2023-12-22 | Cust 1035 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB + MPCB data upload issue. | After restart Apache services now data is live on CPCB + MPCB server. | After restart Apache services now data is live on CPCB + MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23140333642 | 2023-12-22 | 2023-12-23 | Cust 1273 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Flow data not updating on MPCB. | Took access to the system and now data is updating on MPCB Server | Took access to the system and now data is updating on MPCB Server | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23140403643 | 2023-12-22 | 2023-12-23 | Cust 1302 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on MPCB. | We took access to the system and now data is updating on MPCB Server. | We took access to the system and now data is updating on MPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23140493644 | 2023-12-22 | 2023-12-22 | Cust 1251 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB and MPCB data uploded issue. | Now the data is live on the MPCB server. | Now the data is live on the MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23142283645 | 2023-12-26 | 2023-12-31 | Cust 1304 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4 trans-A4TW033000(1) | Commissioning | pH. Turbidity commissioned. | pH. Turbidity commissioned. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFPN23142423646 | 2023-12-26 | 2023-12-29 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Cane Agro Energy ltd. :- (EVERSENSE TO ENVIROCONNECT) | We took access of the system and now data is live on CPCB+LOCAL | We took access of the system and now data is live on CPCB+LOCAL | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23142833647 | 2023-12-27 | 2023-12-29 | Cust 1035 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | ETP data upload issue. | we took access to the system and now the issue has been resolved and data is live on the MPCB server. | we took access to the system and now the issue has been resolved and data is live on the MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23142893648 | 2023-12-27 | 2023-12-29 | Cust 1295 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data upload issue. | we took access to your system and now data is live on the MPCB server. | we took access to your system and now data is live on the MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23142913649 | 2023-12-27 | 2023-12-28 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Flow value not uploading on MPCB | Data from the flow meter is now reflecting on the MPCB server. | Data from the flow meter is now reflecting on the MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23142933650 | 2023-12-27 | 2023-12-28 | Cust 1035 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | harddisk storage issue. | Storage issue has been resolved. | Storage issue has been resolved. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23143133651 | 2023-12-27 | 2023-12-28 | Cust 731 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera time sync issue | We took access to the system and time sync issue has been resolved now. | We took access to the system and time sync issue has been resolved now. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23143143652 | 2023-12-27 | 2023-12-30 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Scanner to RS485 configuration needs to be change | Scanner to RS485 connection change. ( We have made changes as per request. | Scanner to RS485 connection change. ( We have made changes as per request. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23143173653 | 2023-12-27 | 2023-12-28 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Stack data is not uploading on CPCB | Data is now live on CPCB and MPCB servers. | Data is now live on CPCB and MPCB servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23143733654 | 2023-12-28 | 2023-12-30 | Cust 1305 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Reconfiguration of software. | Took access to the system and successfully completed reconfiguration and data is live on MPCB+LOCAL+CPCB | Took access to the system and successfully completed reconfiguration and data is live on MPCB+LOCAL+CPCB | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23144033655 | 2023-12-28 | 2023-12-31 | Cust 517 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Display issue | Visit done and attended the issue of Multipara, Dust and GAs | Visit done and attended the issue of Multipara, Dust and GAs | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN23144053656 | 2023-12-28 | 2023-12-31 | Cust 1242 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | PH high readings | AMC visit done. Necessary maintenance activities carried out for Mutlipara. | AMC visit done. Necessary maintenance activities carried out for Mutlipara. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN23144083657 | 2023-12-28 | 2023-12-28 | Cust 1035 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB and MPCB data uploading issue | Data from STACK2, ETP1 and ETP FLOW is live. | Data from STACK2, ETP1 and ETP FLOW is live on the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23144093658 | 2023-12-28 | 2023-12-30 | Cust 1264 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data uploading issue. | Data is now live on the CPCB server. | Data is now live on the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23144103659 | 2023-12-28 | 2023-12-30 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Flow data uploading issue. | Data is live on the MPCB server. | Data is live on the MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23144303660 | 2023-12-29 | 2023-12-30 | Cust 992 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB + MPCB data upload issue. | Now hardware side issue is solved and data is live on CPCB + MPCB server. | Now hardware side issue is solved and data is live on CPCB + MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23144363661 | 2023-12-29 | 2023-12-30 | Cust 1240 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Electrode | Polilyte Plus VP 120(HT242 428) | Commissioning of PH Sensor | Commissioning has been done successfully. | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN23144363661 | 2023-12-29 | 2023-12-30 | Cust 1240 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | PH input card(1) | Commissioning of Aqua4trans transmitter. | Commissioning has been done. | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN23144383662 | 2023-12-29 | 2023-12-31 | Cust 803 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC visit on urgent basis requested as MP not showing any reading | visited the site MP analyzer issue resolved , mom attached | visited the site MP analyzer issue resolved , mom attached | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN23144413663 | 2023-12-29 | 2023-12-30 | Cust 1306 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Gas Cogen data not updating on CPCB and Local | Took access to the system and did configuration changes now data is live on CPCB+LOCAL | Took access to the system and did configuration changes now data is live on CPCB+LOCAL | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23145183664 | 2023-12-30 | 2023-12-30 | Cust 1279 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on MPCB | Took access to the system and did configuration changes now data is live on CPCB+LOCAL | Took access to the system and did configuration changes now data is live on CPCB+LOCAL | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23145263665 | 2023-12-30 | 2023-12-30 | Cust 1293 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | data not uploading on the MPCB server. | Data from ETP Outlet and Stack 1_50 TPH Boiler is now live on CPCB, MPCB and Local servers. | Data from ETP Outlet and Stack 1_50 TPH Boiler is now live on CPCB, MPCB and Local servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23145353666 | 2023-12-30 | 2024-01-02 | Cust 798 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4Trans(1) | Commissioning of Aqua4trans transmitter. | Commission done. | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN23145353666 | 2023-12-30 | 2024-01-02 | Cust 798 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Chlorine sensor | CHLORINE SENSOR(1) | Commissioning of FCL chlorine Sensor | Commissioning Done. | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN23145363667 | 2023-12-30 | 2024-01-02 | Cust 1307 | PUNE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | Calibration job of 2 Conductivity based TDS transmitter | Calibration done successfully. | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN23145363667 | 2023-12-30 | 2024-01-02 | Cust 1307 | PUNE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | Calibration job of 2 Conductivity based TDS transmitter | Calibration done successfully. | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN23145433668 | 2023-12-30 | 2023-12-30 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | data not uploading on MPCB issue | Data is live on the MPCB server. | Data is live on the MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23146983669 | 2024-01-02 | 2024-01-05 | Cust 1279 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB Data is not updating. | We took access to the system and data is live on MPCB. | We took access to the system and data is live on MPCB. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23147013670 | 2024-01-02 | 2024-01-05 | Cust 1280 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Reconfiguration of Software. | We took access to the system and successfully completed reconfiguration and data is live now on CPCB+MPCB+LOCAL. | We took access to the system and successfully completed reconfiguration and data is live now on CPCB+MPCB+LOCAL. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23147163671 | 2024-01-02 | 2024-01-09 | Cust 1308 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB DATA UPLODED ISSUE. | Issue has been resolved and data is live. | Issue has been resolved and data is live. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23148253672 | 2024-01-04 | 2024-01-05 | Cust 1002 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data mismatch issue | Data mismatch issue has been resolved. | Data mismatch issue has been resolved. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23148313673 | 2024-01-04 | 2024-01-05 | Cust 1002 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Unable to login to MPCB Portal. | Took access to the system ad found credentials were blocked. Customer has takenup issue with MPCB | Took access to the system ad found credentials were blocked. Customer has takenup issue with MPCB | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23148423674 | 2024-01-04 | 2024-01-08 | Cust 1259 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | EUT online connectivity | completion of work | completed the work | NO | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23148833675 | 2024-01-05 | 2024-01-06 | Cust 642 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Conductivity Input card of Aqua4trans(1) | Commissioning of Conductivity based TDS with Aqua4trans | Commission done | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN23148833675 | 2024-01-05 | 2024-01-06 | Cust 642 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Conductivity Input card of Aqua4trans(1) | Commissioning of Conductivity based TDS with Aqua4trans | Commission done successfully. | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN23148833675 | 2024-01-05 | 2024-01-06 | Cust 642 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Electrode | Forbes Marshall Cell Model 8312, K=1(FM 8312-1.5) | Commissioning of 4 Conductivity sensor | Commission done. | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN23148833675 | 2024-01-05 | 2024-01-06 | Cust 642 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Electrode | Forbes Marshall Cell Model 8312, K=1(FM 8312-1.5) | Commissioning of 4 Conductivity sensor | Sensor found faulty and needs to give replacement. | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN23148833675 | 2024-01-05 | 2024-01-06 | Cust 642 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Electrode | Forbes Marshall Cell Model 8312, K=1(FM 8312-1.5) | Commissioning of 4 Conductivity sensor | Commission done. | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN23148833675 | 2024-01-05 | 2024-01-06 | Cust 642 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Electrode | Forbes Marshall Cell Model 8312, K=1(FM 8312-1.5) | Commissioning of 4 Conductivity sensor | Commission done | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN23148843676 | 2024-01-05 | 2024-01-08 | Cust 712 | PUNE | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | AQUADO Analyzer | AQUAMAX DOT(1) | Installation and erection | Guidelines given for the same. | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN23148843676 | 2024-01-05 | 2024-01-08 | Cust 712 | PUNE | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | AQUADO Analyzer | Visiwater DO sensor(1) | installation | Guidelines given for the same. | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN23148983677 | 2024-01-05 | 2024-01-05 | Cust 1306 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on CPCB Server. | We took access to the system and now camera is live on CPCB Server. | We took access to the system and now camera is live on CPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23148993678 | 2024-01-05 | 2024-01-05 | Cust 731 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on CPCB Server and time and date not matching wit PC time and date. | We took access to the system and now camera is live on CPCB Server and time and date is matching with PC time. | We took access to the system and now camera is live on CPCB Server and time and date is matching with PC time. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23149123679 | 2024-01-05 | 2024-01-05 | Cust 1035 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not getting uploaded to MPCB server. | Data is live on CPCB and MPCB servers. | Data is live on CPCB and MPCB servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23149133680 | 2024-01-05 | 2024-01-05 | Cust 1256 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not updating on Local server | Data is live on the Local + CPCB + MPCB servers. | Data is live on the Local + CPCB + MPCB servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23149603681 | 2024-01-06 | 2024-01-16 | Cust 603 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | mA Output card for Multiparameter(AWA-4-20mA) | Kamlesh Bodkhe RE: regd. mismatch of TSS parameter on display and online screen. | discussed with customer , problem identified , issue is with the mA output card, they need to replace it , mail sent to customer is attached | discussed with customer , problem identified , issue is with the mA output card, they need to replace it , mail sent to customer is attached | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFPN23150403682 | 2024-01-08 | 2024-01-11 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Piyu System - Kolhapur Sugars Mills data not updating on CPCB+LOCAL+MPCB | We took access to the system and now data is live on CPCB+MPCB+LOCAL | We took access to the system and now data is live on CPCB+MPCB+LOCAL | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23150493683 | 2024-01-08 | 2024-01-15 | Cust 1309 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4Trans(1) | commissioning | completed the work | completed the work | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23150813684 | 2024-01-09 | 2024-01-11 | Cust 1266 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | DATCON | DATCON(DATCON) | Data is not updating on CPCB and MPCB server | Data is now live on CPCB and MPCB servers. | Data is now live on CPCB and MPCB servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23151193685 | 2024-01-09 | 2024-01-12 | Cust 803 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB+LOCAL+MPCB | We took access to the system and now data is live on CPCB+MPCB+LOCAL | We took access to the system and now data is live on CPCB+MPCB+LOCAL | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23151213686 | 2024-01-09 | 2024-01-12 | Cust 1308 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on MPCB. | We took access to the system and now data is live on MPCB | We took access to the system and now data is live on MPCB | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23151423687 | 2024-01-10 | 2024-01-12 | Cust 1310 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB+MPCB | We took access to the system and now data is live on CPCB+MPCB | We took access to the system and now data is live on CPCB+MPCB | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23151513688 | 2024-01-10 | 2024-01-16 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation | We have re-installed the software with the updated version and completed Local + CPCB + MPCB connectivity and data is live on the same servers. | We have re-installed the software with the updated version and completed Local + CPCB + MPCB connectivity and data is live on the same servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23151523689 | 2024-01-10 | 2024-01-12 | Cust 803 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB+LOCAL+MPCB. | We took access to the system and now data is live on CPCB+MPCB+LOCAL | We took access to the system and now data is live on CPCB+MPCB+LOCAL | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23152183690 | 2024-01-11 | 2024-01-12 | Cust 731 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera date and time not matching with PC time and Date. | Took access to the system and now Camera timing is matching with PC timing. | Took access to the system and now Camera timing is matching with PC timing. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23152693691 | 2024-01-12 | 2024-01-15 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data upload issue. | we took the access to your system and now the issue has been resolved and data is live on the MPCB server. | we took the access to your system and now the issue has been resolved and data is live on the MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23152903692 | 2024-01-12 | 2024-01-12 | Cust 1256 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not updating to CPCB and MPCB servers. | data is now live on the CPCB + Local + MPCB servers. | data is now live on the CPCB + Local + MPCB servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23152923693 | 2024-01-12 | 2024-01-15 | Cust 1281 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Ubuntu Software installation + local + CPCB + MPCB configuration needs to be done. | As of now we are closing this complaint, and will reopen it network side issue is resolve. | As of now we are closing this complaint, and will reopen it network side issue is resolve. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23152933694 | 2024-01-12 | 2024-01-15 | Cust 809 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data upload issue. | As of now we are closing this complaint, and will reopen it after the hardware side issue is resolved for Dust analyzer. | As of now we are closing this complaint, and will reopen it after the hardware side issue is resolved for Dust analyzer. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23153573695 | 2024-01-15 | 2024-01-18 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data mismatch issue. | Now the hardware side issue is resolved, and now data is live on CPCB + MPCB server. | Now the hardware side issue is resolved, and now data is live on CPCB + MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23153703696 | 2024-01-15 | 2024-01-15 | Cust 1311 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4Trans(1) | FOR COMISSIONING VISIT | AQU4TRANS FOR COMMISIONING VISIT IS OK | AQU4TRANS TRANSMITTER RUNNING OK | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFPN23153733697 | 2024-01-15 | 2024-01-21 | Cust 517 | PUNE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | FOR CALLIBERATING THE SYSTEM | need spare for analyser running operational | completed the work | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN23153753698 | 2024-01-15 | 2024-01-22 | Cust 1312 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Electrode | Polilyte HTVP 120(HT238 428) | Commissioning of PH sensor | Commissioning has been done over phone and Branch engineer Mr. prashant Mahurkar has visited found that system working good. | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN23153873699 | 2024-01-15 | 2024-01-18 | Cust 1035 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | CPCB + MPCB data upload issue. | As of now we are closing this complaint, and will reopen it after the hardware side issue is resolved. | As of now we are closing this complaint, and will reopen it after the hardware side issue is resolved. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23153913700 | 2024-01-15 | 2024-01-17 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not updating on CPCB server | Data is now live on the CPCB server. | Data is now live on the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23154743701 | 2024-01-16 | 2024-01-18 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | need to change ranges on Local server | Ranges on local server are changed as per requirement | Ranges on local server are changed as per requirement | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23154953702 | 2024-01-16 | 2024-01-21 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | photo detector (low signal) | customer need to change photo detector in future | completed the work | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23155283703 | 2024-01-17 | 2024-01-22 | Cust 1035 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software re-installation | We have reinstalled the software and completed Local + CPCB + MPCB connectivity but currently data is not coming till the modscan hence as of now we are closing this complaint. We will reopen this once the issue is resolved from the hardware side. | As of now we are closing this complaint. We will reopen this once the issue is resolved from the hardware side. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23155313704 | 2024-01-17 | 2024-01-19 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB and MPCB data upload issue | As of now we are closing this complaint. We will reopen this once the access is provided by customer's end. | As of now we are closing this complaint. We will reopen this once the access is provided by customer's end. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23155643705 | 2024-01-18 | 2024-01-29 | Cust 1313 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Smartpro Electronics Assembly(Smartpro-Elc-Assembly) | ORP commissioning | ORP transmitter commissioned. | ORP transmitter commissioned. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23155693706 | 2024-01-18 | 2024-01-19 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data upload issue. | data is now live on the MPCB server. | data is now live on the MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23155853707 | 2024-01-18 | 2024-01-19 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading on CPCB and MPCB server. | We took access to the system and now the data is transmitting to the Local + CPCB and MPCB servers. | We took access to the system and now the data is transmitting to the Local + CPCB and MPCB servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23156313708 | 2024-01-19 | 2024-01-19 | Cust 603 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Need to add new device and parameter on CPCB and KSPCB. | Took access to the system and added PM data to the CPCB and KSPCB | Took access to the system and added PM data to the CPCB and KSPCB | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23156363709 | 2024-01-19 | 2024-01-19 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | data is not getting updated to the MPCB server. | data is now live on the MPCB server. | data is now live on the MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23156383710 | 2024-01-19 | 2024-01-19 | Cust 1256 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB configuration needs to be done. | we took access to your system did MPCB configuration for Boiler_40_TPH, and now data is live on the MPCB server. | we took access to your system did MPCB configuration for Boiler_40_TPH, and now data is live on the MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23157913711 | 2024-01-23 | 2024-01-31 | Cust 1233 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB + CPCB configuration needs to be done. | we took access to your system and completed CPCB + MPCB connectivity and now data is live on both servers. | we took access to your system and completed CPCB + MPCB connectivity and now data is live on both servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23158073712 | 2024-01-23 | 2024-01-25 | Cust 1306 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Lagoon camera not playing on CPCB. | We took access to the system and now camera is playing on CPCB Server | We took access to the system and now camera is playing on CPCB Server | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23158093713 | 2024-01-23 | 2024-01-25 | Cust 1258 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Date not updating on MPCB | We took access to the system and now data is updating on MPCB Server | We took access to the system and now data is updating on MPCB Server | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23158283714 | 2024-01-23 | 2024-01-29 | Cust 1314 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | multi parameter analyser RS485 O/P is NOT received | suggested to replaced the falcon mini scanner | completed the work | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23158313715 | 2024-01-23 | 2024-01-29 | Cust 1302 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | received call from Suyog for chargeable visit to attend the problem of the MP analyzer not getting power ON | visited the site problem addressed , issue found in Rs485 modbus card | visited the site problem addressed , issue found in Rs485 modbus card | NO | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23158333716 | 2024-01-23 | 2024-01-25 | Cust 1315 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB+LOCAL+MPCB | We took access to the system and now data is updating on LOCAL+CPCB+MPCB Server | We took access to the system and now data is updating on LOCAL+CPCB+MPCB Server | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23158693717 | 2024-01-24 | 2024-01-25 | Cust 1261 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on MPCB Portal | We took access to the system and now data is updating on MPCB Server | We took access to the system and now data is updating on MPCB Server | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23158713718 | 2024-01-24 | 2024-01-29 | Cust 1244 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on CPCB Server | Camera streaming is now live on the CPCB server. | Camera streaming is now live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23158773719 | 2024-01-24 | 2024-01-29 | Cust 1268 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | FOUND DISPLAY TOUCHPAD ISSUE. & SPCU FOUND HANGING HAVING SIGNAL DELAY.. | AMC visit done. Necessary maintenance activities carried out for Mutlipara. | AMC visit done. Necessary maintenance activities carried out for Mutlipara. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFPN23158783720 | 2024-01-24 | 2024-01-29 | Cust 1232 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | DISPLAY TOUCHPAD ISSUE. & SPCU IS HANGING FOUND DELAY IN SIGNAL | AMC visit done. Necessary maintenance activities carried out for Mutlipara. | AMC visit done. Necessary maintenance activities carried out for Mutlipara. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFPN23158933721 | 2024-01-24 | 2024-01-29 | Cust 1255 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | LCD of Multiparameter Display(1) | MP analyzer display not getting power ON | visited the site checked , problem found with the Rs485 Modbus card , communication changed with the mA output cards and scanner | visited the site checked , problem found with the Rs485 Modbus card , communication changed with the mA output cards and scanner , mom attached | NO | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23160443722 | 2024-01-29 | 2024-02-01 | Cust 1035 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not live on CPCB | Camera is now live on the CPCB server | Camera is now live on the CPCB server | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23161113723 | 2024-01-30 | 2024-02-02 | Cust 1266 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Stack data not updating on CPCB and RTDMS Portal. | We took access to the system and now stack data is live on CPCB and RTDMS Server. | We took access to the system and now stack data is live on CPCB and RTDMS Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23161133724 | 2024-01-30 | 2024-02-02 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Stack data is not updating to CPCB and MPCB. need to change configuration from RS485 to scanner. | We have changed the configuration and data is now live on the CPCB and MPCB servers. | We have changed the configuration and data is now live on the CPCB and MPCB servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23162343725 | 2024-01-31 | 2024-02-02 | Cust 1315 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB and MPCB | We took access to the system and data is live on CPCB+LOCAL+MPCB | We took access to the system and data is live on CPCB+LOCAL+MPCB | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23162493726 | 2024-01-31 | 2024-01-31 | Cust 803 | PUNE | Steam System Division | CLOSED | QUALITY PROBLEM | PRV - ARCA | ARCA-POSITIONER - PRESSURE REDUCSNG VALVE - SIZE 80NB(080ARCAPOS) | 080NB MEASURETROL - DOUBLE GASKETS INSTALLED INSIDE THE ACTUATOR WHICH MIGHT RESULT IN LEAKAGE | The same has been taken up with QC, R&D and MFG for a CAPA | The same has been taken up with QC, R&amp;amp;D and MFG for a CAPA | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFPN23162593727 | 2024-01-31 | 2024-01-31 | Cust 803 | PUNE | Steam System Division | CLOSED | QUALITY PROBLEM | PRV - ARCA | ARCA-POSITIONER - PRESSURE REDUCSNG VALVE - SIZE 80NB(080ARCAPOS) | 080NB MEASURETROL - MAIN PCB FAILURE. UNIT NOT POWER ON - 24VDC POWER NOT REACHING FURTHER. LED NOT GLOWING | During the visit the PCB has been changed and the failed PCB has been collected from site for RCA / CAPA - Abhijeet Deshpande / Jaykumar Prasad / Jitesh Patil / Jayant Sowani | During the visit the PCB has been changed and the failed PCB has been collected from site for RCA / CAPA - Abhijeet Deshpande / Jaykumar Prasad / Jitesh Patil / Jayant Sowani | YES | No Trip | Process | Y | Product / equipment failure due to which customers plant or part of plant is stopped which leaads to production loss. | Analytical - CHAKAN
AFPN23162693728 | 2024-01-31 | 2024-01-31 | Cust 803 | PUNE | Steam System Division | CLOSED | QUALITY PROBLEM | PRV - ARCA | ARCA-POSITIONER - PRESSURE REDUCSNG VALVE - SIZE 80NB(080ARCAPOS) | 080NB MEASURETROL - THE METAL LIFTING BRACKET WAS FOUND LOOSE INSIDE THE ACTUATOR | SAME WAS RETIGHTENED TO MAKE THE UNIT READY FOR COMMISSIONING | SAME WAS RETIGHTENED TO MAKE THE UNIT READY FOR COMMISSIONING | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFPN23162713729 | 2024-01-31 | 2024-02-02 | Cust 1315 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB and MPCB Data not updating for distillery unit. | We took access to the system and data is now live on CPCB and MPCB. | We took access to the system and data is now live on CPCB and MPCB. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23162763730 | 2024-01-31 | 2024-01-31 | Cust 803 | PUNE | Steam System Division | CLOSED | MANUFACTURING ISSUES | PRV - ARCA | ARCA-POSITIONER - PRESSURE REDUCSNG VALVE - SIZE 80NB(080ARCAPOS) | 080NB MEASURETROL - JUNCTION BOX (AS CURRENT LENGTH TOO SHORT) AND I&C CHART TO BE PART OF LOOSE MEASURETROL SUPPLIES | RND, QC AND MFG TO REVERT BACK ON THE CAPA | RND, QC AND MFG TO REVERT BACK ON THE CAPA | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFPN23163013731 | 2024-01-31 | 2024-02-05 | Cust 1233 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Spares needs to be replaced | AMC visit done. Necessary maintenance activities carried out for Mutlipara. DUst and flow meter | AMC visit done. Necessary maintenance activities carried out for Mutlipara. DUst and flow meter | NO | Paid | Core | N |  | Analytical - CHAKAN
AFPN23164243732 | 2024-02-01 | 2024-02-02 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not reflecting to MPCB server | We took access to the system and checked, now data is live on the MPCB server. | We took access to the system and checked, now data is live on the MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23164513733 | 2024-02-01 | 2024-02-02 | Cust 1233 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB and MPCB data uploaded issue. | Issue has been resolved and now the data is live on Both the sites | Issue has been resolved and now the data is live on Both the sites | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23164663734 | 2024-02-02 | 2024-02-06 | Cust 623 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Silica Analyser | Silica 8891(8891) | visit for feedback of analyser,inspection & calibration | completed the work | completed the work | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23164753735 | 2024-02-02 | 2024-02-24 | Cust 811 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | not communicating with server , online reading not getting transfer | visited the site issue of data transmission resolved | visited the site issue of data transmission resolved | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN23164833736 | 2024-02-02 | 2024-02-02 | Cust 1315 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Pc getting slow space issue and data not updating on MPCB. | We took access to the customers PC are deleted the log files and restarted the services and now data is live on CPCB+LOCAL+MPCB | We took access to the customers PC are deleted the log files and restarted the services and now data is live on CPCB+LOCAL+MPCB | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23164893737 | 2024-02-02 | 2024-02-02 | Cust 1315 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Jarandeshwar sugar CPCB MPCB Data is not updating. | We took access to the customers PC and now data is live on CPCB+LOCAL+MPCB | We took access to the customers PC and now data is live on CPCB+LOCAL+MPCB | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23165653738 | 2024-02-05 | 2024-02-06 | Cust 1256 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Addition of new device on CPCB and MPCB | We took access to the system and now new device data is live on LOCAL+CPCB+MPCB | We took access to the system and now new device data is live on LOCAL+CPCB+MPCB | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23165853739 | 2024-02-05 | 2024-02-07 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB and MPCB Server. == Loknete Marutrao Ghule Patil Dnyaneshwar Sahakari Sakhar Karkhana | We took access to the system and now data is updating on CPCB and MPCB Server. | We took access to the system and now data is updating on CPCB and MPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23165863740 | 2024-02-05 | 2024-02-08 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on MPCB Server.== Vishwas Rao Naik Sahakari Sakhar Karkhana Ltd | We took access to the system and now data is live on MPCB Server. | We took access to the system and now data is live on MPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23166303741 | 2024-02-06 | 2024-02-09 | Cust 1261 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB and MPCB portal for Distillery unit. | We took access to the system and now updating on CPCB and MPCB portal for Distillery unit. | We took access to the system and now updating on CPCB and MPCB portal for Distillery unit. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23166323742 | 2024-02-06 | 2024-02-09 | Cust 1261 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB and MPCB Portal | We took access to the system and now updating on CPCB and MPCB portal | We took access to the system and now updating on CPCB and MPCB portal | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23166333743 | 2024-02-06 | 2024-02-13 | Cust 867 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | AQUAMAX DOT(1) | DO Analyzer not functioning at Raheja , Kharadi (60003276) project. During Power on , its display is showing continuously 0 ppm. Site engineer Mr. Ajay Katre is available at site. Kindly contact him, Contact no.: 7875085169 | COMPLETED | completed the work | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23166543744 | 2024-02-06 | 2024-02-09 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data upload issue. | We took the access to your system and did some changes from configuration end, and now data is live on MPCB server. | We took the access to your system and did some changes from configuration end, and now data is live on MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23166563745 | 2024-02-06 | 2024-02-09 | Cust 1316 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera is not live on CPCB server | Camera is now live on the CPCB server. | Camera is now live on the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23166753746 | 2024-02-07 | 2024-02-20 | Cust 1252 | PUNE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | Flow cell for Multiparameter(AWA-FLOWCELL) | Flow cell was completely contaminated. | Flow cell was opened and cleaned with tissues | Closed | NO | Paid | Process | N |  | Analytical - CHAKAN
AFPN23166753746 | 2024-02-07 | 2024-02-20 | Cust 1252 | PUNE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | PH1 filter(1) | Both the filters are PH1&PH2 are completely failed | Offer sent for the same. | Closed | NO | Paid | Process | N |  | Analytical - CHAKAN
AFPN23166753746 | 2024-02-07 | 2024-02-20 | Cust 1252 | PUNE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | PH Sensor for Multiparameter(AWA-PH-SENS) | Commissioning of new PH sensor | Commission has been done. | Closed | YES | Paid | Process | N |  | Analytical - CHAKAN
AFPN23167343747 | 2024-02-08 | 2024-02-11 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data upload issue. | We took the access to your system, and now data is live on the MPCB server for Multiparameter. | We took the access to your system, and now data is live on the MPCB server for Multiparameter. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23167363748 | 2024-02-08 | 2024-02-09 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not updating to the MPCB server | The issue has been resolved and data is now live on the MPCB + CPCB + Local server. | The issue has been resolved and data is now live on the MPCB + CPCB + Local server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23167373749 | 2024-02-08 | 2024-02-09 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | PIYU SYSTEM==>Shri Vitthal SSK Ltd MPCB data not sync with CPCB Data | We took access to the system and now data is live on MPCB Server. | We took access to the system and now data is live on MPCB Server. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFPN23167633750 | 2024-02-08 | 2024-02-09 | Cust 1315 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on MPCB Server. | We took access to the system and now data is live on MPCB Server. | We took access to the system and now data is live on MPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23167643751 | 2024-02-08 | 2024-02-09 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Flowmeter data is not updating on local and MPCB server | We took access to the system and did some changes from configuration end and data is now live on the Local and MPCB server. | We took access to the system and did some changes from configuration end and data is now live on the Local and MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23167653752 | 2024-02-08 | 2024-02-09 | Cust 1282 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Unable to login CPCB and MPCB server | now the issue has been resolved and the CPCB and MPCB servers are now accessible in your system. | now the issue has been resolved and the CPCB and MPCB servers are now accessible in your system. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23168993753 | 2024-02-12 | 2024-02-19 | Cust 1269 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | commissinoing | all transmitter commissioning is completed | completed the work | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23169003754 | 2024-02-12 | 2024-02-19 | Cust 1317 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | annual maintain ace contract | AMC | completed the work | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN23169103755 | 2024-02-12 | 2024-02-15 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB + local data upload issue. | now stack data is live on CPCB + local server. | now stack data is live on CPCB + local server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23169593756 | 2024-02-13 | 2024-02-16 | Cust 991 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data was not updating on MPCB Server. | We took access to the system and now data is live on MPCB Server. | We took access to the system and now data is live on MPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23170153757 | 2024-02-14 | 2024-02-16 | Cust 1237 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | data is not uploading Cloud server. | Data is showing live on the MPCB server. | Data is showing live on the MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23170253758 | 2024-02-14 | 2024-02-16 | Cust 1035 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera is not live on CPCB server. | Camera streaming is live on the CPCB server. | Camera streaming is live on the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23170683759 | 2024-02-15 | 2024-02-16 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not updating on MPCB Portal | We took access to the system and now data is live on MPCB Portal. | We took access to the system and now data is live on MPCB Portal. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23170723760 | 2024-02-15 | 2024-02-16 | Cust 1314 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | CPCB + MPCB data upload issue. | we took the access to your system, and changed the configuration of the multiparameter (from RS485 to Scanner). And now multiparameter data is live on CPCB server. | we took the access to your system, and changed the configuration of the multiparameter (from RS485 to Scanner). And now multiparameter data is live on CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23170933761 | 2024-02-16 | 2024-02-16 | Cust 1261 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Stack data not updating on MPCB Server. | We took access to the system and now data is live on MPCB Portal. | We took access to the system and now data is live on MPCB Portal. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23170963762 | 2024-02-16 | 2024-02-16 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Piyu system - Vitthal Sahakari Sakhar Karkhana Ltd Sugar Data not updating on MPCB Portal | We took access to the system and now data is live on MPCB Portal. | We took access to the system and now data is live on MPCB Portal. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23171203763 | 2024-02-16 | 2024-02-16 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on MPCB Server. | We took access to the system and now data is live on MPCB Portal. | We took access to the system and now data is live on MPCB Portal. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23171233764 | 2024-02-16 | 2024-02-16 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data upload issue. | Data is now live on the MPCB server | Data is now live on the MPCB server | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23171243765 | 2024-02-16 | 2024-02-16 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data upload issue. | We took access to the system and now data is reflecting to the Local and MPCB server. Please check and revert. | We took access to the system and now data is reflecting to the Local and MPCB server. Please check and revert. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23172433766 | 2024-02-19 | 2024-03-08 | Cust 1314 | PUNE | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | MULTIPARAMETER ANALYSER | 24 VDC SMPS (1) | , no any earthing point at smps , not anywhere mention power connection terminal ( +ve & -ve ), | arrange anorher module | completed the work | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23172583767 | 2024-02-19 | 2024-02-23 | Cust 1255 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not updating on Local + CPCB + MPCB servers. | We took access to the system and checked, now data is live on the Local + CPCB + MPCB server. | We took access to the system and checked, now data is live on the Local + CPCB + MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23172853768 | 2024-02-20 | 2024-02-22 | Cust 1240 | PUNE | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | Aqua4Trans | PH input card(1) | Troubleshooting of PH sensor installation at process tank. | Liquid level needs to be maintained at tank in such a way that PH sensor should be dipped at all the times. | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN23173123769 | 2024-02-20 | 2024-02-23 | Cust 1035 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | SUYOG ENTERPRISES==Yogeshwari Sugar data was not updating on CPCB + LOCAL+MPCB. | Pramod Patil visited the site and found for flowmeter there is display issue from hardware side and data is not coming till MODSCAN. Data is live for all other devices on CPCB + LOCAL +MPCB server. | Pramod Patil visited the site and found for flowmeter there is display issue from hardware side and data is not coming till MODSCAN. Data is live for all other devices on CPCB + LOCAL +MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23173193770 | 2024-02-20 | 2024-02-23 | Cust 1237 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Ecolyser server data uploaded issue. | Now the data is live on the Ecolyser server. | Now the data is live on the Ecolyser server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23173743771 | 2024-02-21 | 2024-02-23 | Cust 1244 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | DATA NOT UPDATING ON LOCAL +CPCB +MPCB | We took access to the system and found Port was not detected due to lose connection now data is live on LOCAL+CPCB+MPCB. | We took access to the system and found Port was not detected due to lose connection now data is live on LOCAL+CPCB+MPCB. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23173823772 | 2024-02-21 | 2024-02-22 | Cust 712 | PUNE | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | AQUADO Analyzer | AQUAMAX DOT(1) | Installation And erection of Aquamac DO systems | Installation Guidelines given to customer. All the drawings and documents submitted. Site survey has also done and location of installation has also decided on mutual understanding. | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN23173823772 | 2024-02-21 | 2024-02-22 | Cust 712 | PUNE | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | AQUADO Analyzer | AQUAMAX DOT(1) | Installation And erection of Aquamac DO systems | Installation Guidelines given to customer. All the drawings and documents submitted. Site survey has also done and location of installation has also decided on mutual understanding. | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN23173993773 | 2024-02-21 | 2024-02-23 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not updating to MPCB portal. | We took access to the system and data is now live on the MPCB and Local server. | We took access to the system and data is now live on the MPCB and Local server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23174003774 | 2024-02-21 | 2024-02-23 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation | We have reinstalled the software and completed Local + CPCB + MPCB configuration and data is live on the CPCB and Local server. Currently, access is not provided by customer hence we will reopen this complaint once the access is provided for MPCB connectivity. | We have reinstalled the software and completed Local + CPCB + MPCB configuration and data is live on the CPCB and Local server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23174083775 | 2024-02-21 | 2024-02-26 | Cust 1232 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC call for Multiparameter | AMC WORK IS DONE | complete work | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN23174093776 | 2024-02-21 | 2024-02-26 | Cust 1239 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC call for spares replacement | FLOW CELL UNIT WITH STEAM REPLACE | complete work | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN23174543777 | 2024-02-22 | 2024-02-23 | Cust 1315 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB and MPCB Server. | We took access to the system and now data is live on LOCAL+CPCB+MPCB Server. | We took access to the system and now data is live on LOCAL+CPCB+MPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23174563778 | 2024-02-22 | 2024-02-23 | Cust 1290 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Cameras not updating on CPCB Server | We took access to the system and now both cameras are live on CPCB Server. | We took access to the system and now both cameras are live on CPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23176363779 | 2024-02-26 | 2024-03-03 | Cust 1264 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | FOR ANALYSER AMC VISIT | DUST & MUTIPARAMETER ANALYSER AMC WORK HAS BEEN DONE. | complete work | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN23176383780 | 2024-02-26 | 2024-03-03 | Cust 1239 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | FOR AMC VISIT | CHARGEBLE VISIT HAS BEEN DONE | complete work | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN23176483781 | 2024-02-26 | 2024-02-29 | Cust 1308 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not updating to the MPCB server. | We will reopen this system once the access is provided by the customer. | We will reopen this system once the access is provided by the customer. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23176493782 | 2024-02-26 | 2024-02-29 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data issue. | completed Local + CPCB + MPCB connectivvity and data is live on the same servers | We have reinstalled the software and completed Local + CPCB + MPCB connectivvity and data is live on the same servers | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23176613783 | 2024-02-26 | 2024-02-29 | Cust 1282 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | data not updating on MPCB portal | We took access to the system and data is now live on MPCB Portal. | We took access to the system and data is now live on MPCB Portal. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23176653784 | 2024-02-26 | 2024-02-29 | Cust 1308 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | data is not updating to the MPCB server. | There was an internet issue from customer side. guided customer on call. | There was an internet issue from customer side. guided customer on call. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23176863785 | 2024-02-26 | 2024-02-29 | Cust 809 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data upload issue. | We took the access to your system and found that ApacheTOMCAT service is stopped in customers local PC. After restarting the same, data is now live on the MPCB server. | We took the access to your system and found that ApacheTOMCAT service is stopped in customers local PC. After restarting the same, data is now live on the MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23176893786 | 2024-02-26 | 2024-02-29 | Cust 603 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB + MPCB data upload issue. | Data is now live. | Data is now live. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23177143787 | 2024-02-27 | 2024-02-29 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on MPCB Portal | We took access to your system and data is live on MPCB Portal | We took access to your system and data is live on MPCB Portal | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23177443788 | 2024-02-27 | 2024-03-01 | Cust 1035 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not updating to the CPCB and MPCB server | Data is now live on the CPCB + MPCB + Local servers. | Data is now live on the CPCB + MPCB + Local servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23178233789 | 2024-02-28 | 2024-02-29 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not updating on MPCB Portal | We took access to the system and now data is updating on MPCB Portal | We took access to the system and now data is updating on MPCB Portal | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23178243790 | 2024-02-28 | 2024-03-01 | Cust 1297 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on MPCB Server. | As of now we are closing this complaint and we will reopen once we get remort access of the system. | As of now we are closing this complaint, and will reopen it once the system is accessible. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23179273791 | 2024-02-29 | 2024-03-04 | Cust 1245 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | received confirmation to viist the site under AMC to attend the problem of the Magnetic flowmeter a nd online monitoring system for distillery and sugar | communication problem resolved , enviroconnect software installed and communication checked on modscan found OK , data live on CPCB & MPCB . | rahul viisted the site and issue resolved succsfully. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN23179283792 | 2024-02-29 | 2024-03-04 | Cust 1245 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Unit 2 PC not working , also MP not communicating with the software | iseue addressed , ph calibration completed b, need to replace acid cleaning pump head. | rahul viisted the site and issue resolved succsfully. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN23180013793 | 2024-02-29 | 2024-03-01 | Cust 1035 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not updating to CPCB and MPCB | Data is now live on the CPCB and MPCB servers. | Data is now live on the CPCB and MPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23180073794 | 2024-02-29 | 2024-03-01 | Cust 1318 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software re-installation | We have re-installed the software with an updated version and data is now live on the CPCB server. | We have re-installed the software with an updated version and data is now live on the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23180093795 | 2024-02-29 | 2024-02-29 | Cust 795 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on CPCB Server. | We took access to the system and now data is live on CPCB Server. | We took access to the system and now data is live on CPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23181373796 | 2024-03-03 | 2024-03-03 | Cust 1292 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC VISIT REQIURE DOR DUST & MULTIPAMETER ANALYSER | DUST & MULTIPAMETER ANALYSER WORK DONE | complete work | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN23182523797 | 2024-03-06 | 2024-03-08 | Cust 1261 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on MPCB Server. | We took access to the system and data is live on MPCB Server. | We took access to the system and data is live on MPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23182883798 | 2024-03-07 | 2024-03-08 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not updating to MPCB server | We took access to the system and completed the MPCB connectivity and the data is live on the same for all devices. | We took access to the system and completed the MPCB connectivity and the data is live on the same for all devices. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23183073799 | 2024-03-08 | 2024-03-08 | Cust 1306 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | DATE and TIME synch issue and data not updating on MPCB Server. | We took access to the system and Date and Time sync issue has been resolved and data is live on LOCAL+CPCB+MPCB. | We took access to the system and Date and Time sync issue has been resolved and data is live on LOCAL+CPCB+MPCB. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23183513800 | 2024-03-09 | 2024-03-09 | Cust 1035 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software re-installation | We have reinstalled the software with the updated version and data is now live on the Local + CPCB + MPCB servers. | We have reinstalled the software with the updated version and data is now live on the Local + CPCB + MPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23184123801 | 2024-03-11 | 2024-03-15 | Cust 811 | PUNE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | black box data break in running condition | CHECKED DATA ON MODSCAN SOFTWARE, ONLINE DATA CONNECTIVITY ESTABLISHED | completed the work | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23184423802 | 2024-03-11 | 2024-03-14 | Cust 712 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | AQUAMAX DOT(1) | Commissioning of 2 nos. Aquamax Dot transmitter with sensor. | Commissioning Done. | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN23184423802 | 2024-03-11 | 2024-03-14 | Cust 712 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | AQUAMAX DOT(1) | Commissioning of 2 nos. Aquamax Dot transmitter with sensor. | Commissioning Done | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN23184723803 | 2024-03-12 | 2024-03-15 | Cust 632 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software re-installation | We have re-installed the software with updated version and completed Local + CPCB connectivity and the data is live on the same servers. | We have re-installed the software with updated version and completed Local + CPCB connectivity and the data is live on the same servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23184793804 | 2024-03-12 | 2024-03-17 | Cust 811 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | EUTS invoice submission | analyzer data not show MPCB server now resolve the problem | complete work | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN23185193805 | 2024-03-13 | 2024-03-14 | Cust 1273 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on MPCB Server. | We took access of the system and data is now live on MPCB Server. | We took access of the system and data is now live on MPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23186223806 | 2024-03-15 | 2024-03-18 | Cust 803 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue. | Now Stack and Flow data is live on the server. | Now Stack and Flow data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23186243807 | 2024-03-15 | 2024-03-15 | Cust 632 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera is not streaming on CPCB server | Camera is streaming live on the CPCB server | Camera is streaming live on the CPCB server | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23186303808 | 2024-03-15 | 2024-03-15 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | configuration needs to be change from RS485 to scanner | We made changes as per requirement and Data is now live on the CPCB server. | We made changes as per requirement and Data is now live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23186383809 | 2024-03-15 | 2024-03-15 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data upload issue. | Data is now live on the MPCB server. | Data is now live on the MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23186583810 | 2024-03-16 | 2024-03-18 | Cust 1319 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | DO and pH electrode commissioning (newly purchased by customer) | completed the work | completed the work | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23186613811 | 2024-03-16 | 2024-03-18 | Cust 603 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Analyser display blank issue (ONLINE SUPPORT & SOLVED THE ISSUE) | DISPLAY PROBLEM & GIVES ONLINE SUPPORT | completed the work | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23188013812 | 2024-03-18 | 2024-03-18 | Cust 1299 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on LOCAL+CPCB+MPCB | We took access of the system and data is live on CPCB+LOCAL+MPCB Server. | We took access of the system and data is live on CPCB+LOCAL+MPCB Server. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFPN23188633813 | 2024-03-19 | 2024-03-21 | Cust 1285 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | DATCON | DATCON(DATCON) | CPCB and MPCB data uploaded issue. | As the scanner is changed from 24VDC to 230VAC so now we have changed the configuration as per the new scanner and data is live on both servers | As the scanner is changed from 24VDC to 230VAC so now we have changed the configuration as per the new scanner and data is live on both servers | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23188673814 | 2024-03-19 | 2024-03-22 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data upload issue. | As of now we are closing this complaint, and will reopen it once the internet side issue is resolved. | As of now we are closing this complaint, and will reopen it once the internet side issue is resolved. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23188913815 | 2024-03-19 | 2024-03-30 | Cust 1255 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commisioning of RS-485 in the supply. | visited the site problem addressed , rs 485 card replaced | visited the site problem addressed , rs 485 card replaced | NO | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23189743816 | 2024-03-20 | 2024-03-22 | Cust 1299 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB+MPCB Server | We took access to the system and data is live on CPCB+LOCAL+MPCB. | We took access to the system and data is live on CPCB+LOCAL+MPCB. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23191233817 | 2024-03-22 | 2024-04-25 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Local + MPCB configuration needs to be done. | Internet issue from customer side, we will reopen this complaint once the internet issue is resolved. | Internet issue from customer side, we will reopen this complaint once the internet issue is resolved. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23191363818 | 2024-03-22 | 2024-03-22 | Cust 1255 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on CPCB Server. | We took access to the system and Camera is live on CPCB. | We took access to the system and Camera is live on CPCB. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23191393819 | 2024-03-22 | 2024-03-22 | Cust 795 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | camera not playing on CPCB server. | We took access to the system and Camera is live on CPCB. | We took access to the system and Camera is live on CPCB. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23192943820 | 2024-03-26 | 2024-03-29 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB connectivity needs to be done. | we took the access to your system and now the issue has been resolved and data is live on Local + CPCB + MPCB server | we took the access to your system and now the issue has been resolved and data is live on Local + CPCB + MPCB server | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23193283821 | 2024-03-26 | 2024-03-28 | Cust 1035 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not updating to CPCB and MPCB server | Data is now live on the Local + CPCB + MPCB servers. | Data is now live on the Local + CPCB + MPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23193753822 | 2024-03-27 | 2024-03-29 | Cust 651 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Black Box configuration and Cloud server connectivity need to be done. | Black Box configuration and cloud server connectivity has been completed and data is live on the server. | Black Box configuration and cloud server connectivity has been completed and data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23193843823 | 2024-03-27 | 2024-03-29 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB connectivity | We have completed CPCB and Local connectivity for STACK_30TPH and data is live on the same servers. | We have completed CPCB and Local connectivity for STACK_30TPH and data is live on the same servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23194033824 | 2024-03-27 | 2024-03-30 | Cust 651 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | commissioning of cx-2000 analyser | commissionig of analyser is completed | MPA-CX2000 SUCCESFULLY COMMISSIONED | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23194043825 | 2024-03-27 | 2024-03-30 | Cust 692 | PUNE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | MAINTENANCE OF THE ANALYSER | neccesary maintainance done,under AMC | AMC visit done. Necessary maintenance activities carried out for Mutlipara. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN23194073826 | 2024-03-27 | 2024-03-29 | Cust 1320 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Software installation, CPCB and MPCB connectivity | As of now we are closing this complaint. We will reopen this once the issue is resolved from the hardware side for ETP Outlet. | As of now we are closing this complaint. We will reopen this once the issue is resolved from the hardware side for ETP Outlet. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23194633827 | 2024-03-28 | 2024-03-30 | Cust 1320 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | received call from Mr. Yogesh Datir of valtech engineering , to attend the online monitoring system problem software issue no communication on Modscan | visited the site problem of the modbus communication addressed | visited the site problem of the modbus communication addressed | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN23194663828 | 2024-03-28 | 2024-06-24 | Cust 1237 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Ecolyser Data uploaded issue. | For the ENDURANCE TECHNOLOGIES LIMITED-E92 Site, we are getting successful responses in logs but the data is not reflecting on the server hence there is no issue from our end and the customer is checking this issue with the Ecolyser team hence from our end we are closing this complaint. | Customer is checking this issue with the Ecolyser team hence from our end we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23194693829 | 2024-03-28 | 2024-03-29 | Cust 1321 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on LOCAL+CPCB+MPCB | We took access to the system and now data is live on LOCAL+CPCB+MPCB | We took access to the system and now data is live on LOCAL+CPCB+MPCB | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23194813830 | 2024-03-28 | 2024-03-30 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to Local, CPCB and MPCB servers. | CPCB + MPCB server url and FTP ports are blocked in customers network. we will reopen this complaint once the issue is resolved from customers end. | CPCB + MPCB server url and FTP ports are blocked in customers network. we will reopen this complaint once the issue is resolved from customers end. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23195303831 | 2024-03-29 | 2024-04-01 | Cust 988 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation needs to be done. | we took the access to your system and completed the software reinstallation + local + CPCB + MPCB configuration. And now data is live on all the servers. | we took the access to your system and completed the software reinstallation + local + CPCB + MPCB configuration. And now data is live on all the servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN23195593832 | 2024-03-29 | 2024-03-29 | Cust 1322 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on MPCB Server. | We took access to the system and now data is live on MPCB Server. | We took access to the system and now data is live on MPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23195813833 | 2024-03-29 | 2024-03-30 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | flow data & camera not sync with local + CPCB & MPCB | We took access to the system and now data is live on the Local, CPCB and MPCB servers. | We took access to the system and now data is live on the Local, CPCB and MPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN23195823834 | 2024-03-29 | 2024-03-29 | Cust 811 | PUNE | Analytical - CHAKAN | CLOSED | DIGITAL CONEECTIVITY | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | DATA ARE NOT SHOW IN MPCB SERVER | Dear Sir, Today we have updated the Latest patch in both the BBs for the TATA motors Pimpri site so now the data is live on the server. We will observe the site for 2/3 days and get back to you. | complete work | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN23195993835 | 2024-03-29 | 2024-03-30 | Cust 1323 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | AQUAMAX DOT(1) | commissioning of aquamax DOT DO analyser | commissioning completed | Completed the site | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN23196503836 | 2024-03-30 | 2024-03-30 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB Server forLoknete Marutrao Ghule Patil DSSK -Distillery | We took access of the system and data was live there was no issue from server side . | We took access of the system and data was live there was no issue from server side . | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24000323837 | 2024-04-01 | 2024-04-05 | Cust 795 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera is not streaming on CPCB Server. | We took access of the system and found streaming window was closed. Now camera is live on CPCB Server. | We took access of the system and found streaming window was closed. Now camera is live on CPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24000333838 | 2024-04-01 | 2024-04-05 | Cust 1252 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Local host page is not accessible. | We took access to the system and found services were closed and now data is live on Server. | We took access to the system and found services were closed and now data is live on Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24000943839 | 2024-04-02 | 2024-04-11 | Cust 1252 | PUNE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | PH2 Filter(1) | Installation of PH2 and PH 1 filter | Installation of the PH1 & PH 2 filter has been done. Analyser started working. | Closed | YES | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN24001283840 | 2024-04-03 | 2024-04-05 | Cust 1282 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not getting uploaded to CPCB and MPCB servers. | As of now we are closing this complaint, we will reopen this once the issue is resolved from customer end. | As of now we are closing this complaint, we will reopen this once the issue is resolved from customer end. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24001313841 | 2024-04-03 | 2024-04-08 | Cust 803 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | maintenance of analyser has been done. | amc completed for MPA analyser | completed the work | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24002573842 | 2024-04-05 | 2024-04-09 | Cust 526 | PUNE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | checking healthiness & maintainance | maintenance is done & analyser is working normal | completed the work | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24002593843 | 2024-04-05 | 2024-04-05 | Cust 1302 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on MPCB Server. | There was time synch issue data is live on MPCB Server. | There was time synch issue data is live on MPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24002623844 | 2024-04-05 | 2024-04-05 | Cust 617 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | New industry addition and installation of software. | We took access to the system and data is live on LOCAL Server. | We took access to the system and data is live on LOCAL Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24002703845 | 2024-04-05 | 2024-04-08 | Cust 803 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue. | We took access to the system, and changed the configuration (scanner), and now data is live on the CPCB server (Multiparameter). | We took access to the system, and changed the configuration (scanner), and now data is live on the CPCB server (Multiparameter). | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24004123846 | 2024-04-09 | 2024-04-10 | Cust 1025 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data upload issue. | now all the Flow data is uploading on the MPCB server. | now all the Flow data is uploading on the MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24004703847 | 2024-04-10 | 2024-04-12 | Cust 731 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on CPCB Server. | Issue has been resolved. | Issue has been resolved. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24004713848 | 2024-04-10 | 2024-04-12 | Cust 1252 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not not updating to local and MPCB servers. | Data is now live on the Local + CPCB + MPCB servers. | Data is now live on the Local + CPCB + MPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24004793849 | 2024-04-10 | 2024-04-12 | Cust 1324 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | MPCB connectivity needs to be done. | we have completed the MPCB connectivity and now the Flow data is live on the MPCB server. | we have completed the MPCB connectivity and now the Flow data is live on the MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24005553850 | 2024-04-12 | 2024-04-15 | Cust 1242 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | CUSTMER REQUIRE AMC FOR MULTIPARAMETER ANALYSER | AMC VISIT FOR MULTIPARAMETER ANALYSER HAS BEEN DONE. | MULTIPARAMETER ANALYSER WORKING SATISFACTORY,PH ELECTRODE REQUIRED FOR SPARE. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFPN24005563851 | 2024-04-12 | 2024-04-15 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data upload issue | We took access to the system and data is now live on the Local and MPCB server. | We took access to the system and data is now live on the Local and MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24005573852 | 2024-04-12 | 2024-04-15 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data upload issue | Data is now live on the MPCB server. | Data is now live on the MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24005673853 | 2024-04-12 | 2024-04-22 | Cust 1325 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Smartpro-con) | received call from Tushar Patil for the Condensate return system supplied by Excell engineer | discussed over phone issue resolved successfully. , | discussed over phone issue resolved successfully. , | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFPN24005683854 | 2024-04-12 | 2024-04-15 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB + KSPCB connectivity needs to be done. | we took access to your system and completed the CPCB + KSPCB + local connectivity. And now data is live on all the servers. | we took access to your system and completed the CPCB + KSPCB + local connectivity. And now data is live on all the servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24006413855 | 2024-04-15 | 2024-04-18 | Cust 1233 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation + CPCB + MPCB configuration needs to be done. | today we took access to your system and now the hardware side issue is resolved and data is live on CPCB + MPCB server. | today we took access to your system and now the hardware side issue is resolved and data is live on CPCB + MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24006423856 | 2024-04-15 | 2024-04-15 | Cust 1326 | PUNE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Mismatching the values of the scanner with BMS | RS-485 card has to be connected with the scanner for matching values | customer satisfied with values and suggestions | NO | Paid | Core | N |  | Analytical - CHAKAN
AFPN24006843857 | 2024-04-16 | 2024-04-19 | Cust 803 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on MPCB Server. | We took access to the system now data is live on MPCB Server | We took access to the syste now data is live on MPCB Server | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24006873858 | 2024-04-16 | 2024-04-19 | Cust 1282 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software re-installation | We have re-installed the software with the updated version and completed Local + CPCB + MPCB connectivity and the data is live on the same servers. | We have re-installed the software with the updated version and completed Local + CPCB + MPCB connectivity and the data is live on the same servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24006923859 | 2024-04-16 | 2024-04-19 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on MPCB Server. Loknete Marutrao Ghule Patil Dnyaneshwar Sahakari Sakhar Karkhana Ltd | We took access to the system now data is live on MPCB Server | We took access to the system now data is live on MPCB Server | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFPN24007373860 | 2024-04-17 | 2024-04-19 | Cust 1252 | PUNE | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | MULTIPARAMETER ANALYSER | PH1 detector(1) | Ph1 and Ph2 detector circuit issue | Zero calibration has been done. Alignment has also done. Data is also uploading on CPCB portal. Only there is error on display but it doesn't hamper the data connectivity. | Closed | NO | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN24007703861 | 2024-04-18 | 2024-04-22 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | software re-installation. | Currently we are closing this complaint. We will reopen this once we receive access to the system. | Currently we are closing this complaint. We will reopen this once we receive access to the system. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24008693862 | 2024-04-20 | 2024-04-24 | Cust 811 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Connectivity issue | Visit done and sort out the issue of data disconnection to CPCB/MPCB | Visit done and sort out the issue of data disconnection to CPCB/MPCB | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24009313863 | 2024-04-22 | 2024-04-24 | Cust 617 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB Configuration. | We took access to the system and successfully completed MPCB configuration. Now data is live on MPCB Server. | We took access to the system and successfully completed MPCB configuration. Now data is live on MPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24009943864 | 2024-04-23 | 2024-04-26 | Cust 795 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating to CPCB and MPCB Portal. | As of now we are closing this complaint. We will reopen this complaint once we get access. | As of now we are closing this complaint. We will reopen this complaint once we get access. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24010063865 | 2024-04-23 | 2024-04-26 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data upload issue. | Data is live on the MPCB server. | Data is live on the MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24010103866 | 2024-04-23 | 2024-04-29 | Cust 809 | PUNE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | analyser under maintainance | maintain the analyser | completed the work | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24010123867 | 2024-04-23 | 2024-04-26 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data mismatch issue | Data is live on the CPCB server. | Data is live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24010593868 | 2024-04-24 | 2024-04-25 | Cust 1300 | PUNE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | LCD of Multiparameter Display(1) | Installation of LCD display and spares installment | Successfully done. | Closed | YES | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN24010683869 | 2024-04-24 | 2024-04-26 | Cust 1289 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on MPCB Server. | We took access to the system and now data is live on MPCB Server. | We took access to the system and now data is live on MPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24011103870 | 2024-04-24 | 2024-04-29 | Cust 1237 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Data connectivity issue at ECOLYSER server | Visit done .Regarding ecolyser connectivity, log recorded and sent to ECOLYSER team. | Visit done .Regarding ecolyser connectivity, log recorded and sent to ECOLYSER team. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24011613871 | 2024-04-25 | 2024-04-29 | Cust 564 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | conductivity transmitter showing zero reading continuously. | visited the site , checked both the cond tx found problem with the module, both the modules found dfauty , need to send it to factory for further repair / testing | visited the site , checked both the cond tx found problem with the module, both the modules found dfauty , need to send it to factory for further repair / testing | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24011653872 | 2024-04-25 | 2024-04-29 | Cust 1237 | PUNE | Analytical - CHAKAN | CLOSED | AMC | Aqua4Trans | Aqua4Trans(1) | Analyser connectivity to CETP (Ecolyser server) | Visit done .Regarding Ecolyser connectivity, log recorded and sent to ECOLYSER team. | Ecolyser connectivity checked and necessary software upgrade done. | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24011693873 | 2024-04-25 | 2024-04-26 | Cust 1245 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB configuration needs to be done. | We took access to your system and completed MPCB connectivity and currently ETP1 data is live on MPCB server. | We took access to your system and completed MPCB connectivity and currently ETP1 data is live on MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24011903874 | 2024-04-25 | 2024-04-25 | Cust 904 | PUNE | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | MULTIPARAMETER ANALYSER | Flow cell for Multiparameter(AWA-FLOWCELL) | Cleaning of Flow cell, Sample tubes, PH chamber, PH calibration | Analyzer started working satisfactory. | Closed | NO | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN24012043875 | 2024-04-26 | 2024-04-29 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB connectivity | We have completed CPCB connectivity and data is now live on the same. Camera is also streaming live on the CPCB. | We have completed CPCB connectivity and data is now live on the same. Camera is also streaming live on the CPCB. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24012393876 | 2024-04-26 | 2024-04-26 | Cust 809 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Black Box Commissioning | we have completed the Black Box configuration +CPCB+MPCB and Cloud server connectivity and now the data is live on the server. | we have completed the Black Box configuration +CPCB+MPCB and Cloud server connectivity and now the data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24012423877 | 2024-04-26 | 2024-04-26 | Cust 1237 | PUNE | Analytical - CHAKAN | CLOSED | DIGITAL CONEECTIVITY | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Digital connectivity needs to be done | we have completed the Cloud server connectivity for Endurance Technologies Ltd E71 plant and now the data is live on the server. | we have completed the Cloud server connectivity for Endurance Technologies Ltd E71 plant and now the data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24013303878 | 2024-04-29 | 2024-05-01 | Cust 1002 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not getting uploaded to CPCB and MPCB servers. | Data is now live on the CPCB and MPCB servers. | Data is now live on the CPCB and MPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24013503879 | 2024-04-29 | 2024-05-05 | Cust 1243 | PUNE | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Proactive visit for the discussion regarding Enhanced Uptime service offer. | Visited the site and necessary discussion carried out for Enhanced Uptime Solutions | Visited the site and necessary discussion carried out for Enhanced Uptime Solutions | NO | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24013633880 | 2024-04-29 | 2024-05-05 | Cust 651 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT UPGRADE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Software Upgrade | Visited the site and updated the software for pH. Calibration /verification done. | Visited the site and updated the software for pH. Calibration /verification done. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24013863881 | 2024-04-29 | 2024-05-01 | Cust 719 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | New CX 2000 commissioning | New CX 2000 commissioning done | New CX 2000 commissioning done | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFPN24013923882 | 2024-04-29 | 2024-05-01 | Cust 719 | PUNE | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | CX 2000 installation and guidelines | CX 2000 installation and guidelines given | CX 2000 installation and guidelines given | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFPN24014333883 | 2024-04-30 | 2024-05-03 | Cust 618 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | communication error of ETP | Data is not coming till modscan.As of now we are closing this complaint.we will reopen this once the issue is resolved from the hardware side. | Data is not coming till modscan.As of now we are closing this complaint.we will reopen this once ,when the issue is resolved from the hardware side. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24014393884 | 2024-04-30 | 2024-05-03 | Cust 1281 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | CPCB camera streaming issue. | the camera issue has been resolved from our side and the Camera streaming is live on the server. | the camera issue has been resolved from our side and the Camera streaming is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24014403885 | 2024-04-30 | 2024-05-08 | Cust 811 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Data Breaking issue on CPCB and MPCB server. | As per R&D had some changes in BB after that the Data has been updating continuously from the last 1 week. | As per R&amp;amp;D had some changes in BB after that the Data has been updating continuously from the last 1 week. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24014453886 | 2024-04-30 | 2024-05-03 | Cust 1327 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Data is not uploading to CPCB and MPCB server | Data is now live on the CPCB and MPCB server. | Data is now live on the CPCB and MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24015103887 | 2024-04-30 | 2024-05-01 | Cust 1252 | PUNE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration | Service Visit done | Service Visit done | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFPN24015473888 | 2024-05-01 | 2024-05-03 | Cust 811 | PUNE | Analytical - CHAKAN | CLOSED | DIGITAL CONEECTIVITY | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | BLACK BOX DATA IS NOT CONTINUOSLY RECEIVED CUSTOMER END. | BLACK BOX UPGRADATION WORK IS DONE NOW DATA IS CONTINUOSLY RECEIVED CUSTOMER. | BLACK BOX UPGRADATION WORK IS DONE NOW DATA IS CONTINUOSLY RECEIVED CUSTOMER. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFPN24015483889 | 2024-05-01 | 2024-05-06 | Cust 640 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | MULTIPARAMETER ANLAYSER READING ARE NOT PROPERLY RECEIVED, | Analyser maintenance has been done,clean the ph sensor | completed the work | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24015533890 | 2024-05-01 | 2024-05-20 | Cust 651 | PUNE | Analytical - CHAKAN | CLOSED | DIGITAL CONEECTIVITY | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Router Hanging issue. | We have updated the router firmware and now from the last 15 days, the data has been uploading continuously. | We have updated the router firmware and now from the last 15 days, the data has been uploading continuously. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24015543891 | 2024-05-01 | 2024-05-01 | Cust 1254 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Camera streaming issue on CPCB server. | now the camera streaming is live on the server. | now the camera streaming is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24016003892 | 2024-05-01 | 2024-05-05 | Cust 1328 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning visit | Commissioning activities completed for CX-2000. | Commissioning activities completed for CX-2000. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24016033893 | 2024-05-01 | 2024-05-06 | Cust 1329 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | DO Transmitters | VISI pro DO(VISI pro DO) | received mail from Alkyl AMines for the complaint of the DO transmitter not showing correct DO reading | visited the site problem of the DO sensor resolved succesfully | visited the site problem of the DO sensor resolved succesfully | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24016213894 | 2024-05-02 | 2024-05-03 | Cust 1035 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data upload issue | Data from COGEN GAS and COGEN DUST is now reflecting on the CPCB and MPCB servers. | Data from COGEN GAS and COGEN DUST is now reflecting on the CPCB and MPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24016233895 | 2024-05-02 | 2024-05-03 | Cust 1289 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | DATA NOT UPDATING ON MPCB SERVER. | We took access to your system and now data is live on MPCB Server. | We took access to your system and now data is live on MPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24016243896 | 2024-05-02 | 2024-05-03 | Cust 1330 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on MPCB Server. | We took access to your system and now data is live on MPCB Server. | We took access to your system and now data is live on MPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24016623897 | 2024-05-03 | 2024-05-06 | Cust 633 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Turbidity Meter | Krohne turbidity transmitter(1) | Turbidity meter & flowmeter commissioning and guide how to installation . | two flow meter primary commissioning work is done. | flowmeter primary commissioning work is done. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFPN24017393898 | 2024-05-06 | 2024-05-09 | Cust 1002 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating to CPCB and MPCB Server. | We took access to the system and now data is live on LOCAL+CPCB+MPCB Server. | We took access to the system and now data is live on LOCAL+CPCB+MPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24017653899 | 2024-05-06 | 2024-05-09 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | KSPCB connectivity | We have completed the KSPCB connectivity and the data is live on the same server. | We have completed the KSPCB connectivity and the data is live on the same server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24018063900 | 2024-05-07 | 2024-05-09 | Cust 1035 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not live on the CPCB and MPCB server. | Data is not coming till the modscan hence as of now we are closing this complaint. We will reopen this once the issue is resolved from the hardware side. | As of now we are closing this complaint. We will reopen this once the issue is resolved from the hardware side. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFPN24018083901 | 2024-05-07 | 2024-05-10 | Cust 1327 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to CPCB and MPCB server. | As of now we are closing this complaint. We will reopen this once the issue is resolved from the hardware side. | As of now we are closing this complaint. We will reopen this once the issue is resolved from the hardware side. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24018133902 | 2024-05-07 | 2024-05-08 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB + Local + MPCB connectivity of DSSK distillery division , data is not synching with Local server. | We restarted the services and data was live on CPCB+Local+MPCB server. | We restarted the services and data was live on CPCB+Local+MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24018493903 | 2024-05-08 | 2024-05-10 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB connectivity | We have completed MPCB connectivity for Stack_30TPH and data is live on the same server. | We have completed MPCB connectivity for Stack_30TPH and data is live on the same server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24018623904 | 2024-05-08 | 2024-05-10 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | KSPCB connectivity needs to be done. | We took access of the system and completed KSPCB connectivity, and now data is live on the KSPCB server. | We took access of the system and completed KSPCB connectivity, and now data is live on the KSPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24019133905 | 2024-05-09 | 2024-05-14 | Cust 1331 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Silica Analyser | Silica 8891(8891) | NEW SUPPLIED ANALYSER COMMISSIONING | COMMISSIONING COMPLETED & PENDING ISSUE ADRESS IN NEXT VISIT | completed the work | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24019213906 | 2024-05-10 | 2024-05-10 | Cust 1237 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Flow Data Mismatches issue on Ecolyser Server. | We have checked and whatever readings are showing at the Flow Display the same values are uploading to the server and the same is verified with the couteomer | We have checked and whatever readings are showing at the Flow Display the same values are uploading to the server and the same is verified with the couteomer | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24019373907 | 2024-05-10 | 2024-05-16 | Cust 1294 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Turbidity Meter | Optex turbidity sensor(1) | Commissioning | installation guidelines has been given | completed the work | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24019383908 | 2024-05-10 | 2024-05-10 | Cust 1249 | PUNE | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | Conductivity Electrode | Forbes Marshall Cell Model 8311, K=0.1(FM -8311) | Installation guidelines | Necessary guidelines given to the customer over phone. | Necessary guidelines given to the customer over phone. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFPN24019463909 | 2024-05-10 | 2024-05-10 | Cust 1235 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB connectivoty for COD parameter | Now the COD parameter data is live on the MPCB server. | Now the COD parameter data is live on the MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24020833910 | 2024-05-14 | 2024-05-15 | Cust 904 | PUNE | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Pro active visit for listing of spares and AMC proposal. | Analyzer is working good. | Closed | NO | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN24021363911 | 2024-05-15 | 2024-05-17 | Cust 1002 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB and MPCB server. | data is live now on LOCAL+CPCB+MPCB | data is live now on LOCAL+CPCB+MPCB | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24021763912 | 2024-05-16 | 2024-05-17 | Cust 1310 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on LOCAL+CPCB+MPCB Server. | data is live on LOCAL+CPCB+MPCB server. | data is live on LOCAL+CPCB+MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24022233913 | 2024-05-17 | 2024-05-28 | Cust 1331 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Silica Analyser | Silica 9210(9210) | analyser under commissioning reagents stopping problem observed | Replaced the spare of the analyser i,e reagent tubing & pumps | completed the work | YES | Warranty (Free) | Core | Y |  | Analytical - CHAKAN
AFPN24023023914 | 2024-05-20 | 2024-05-23 | Cust 651 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB connectivity needs to be done. | today we took access to your system and completed CPCB connectivity, and now data is live on CPCB server for ETP Outlet. | today we took access to your system and completed CPCB connectivity, and now data is live on CPCB server for ETP Outlet. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24023493915 | 2024-05-20 | 2024-05-23 | Cust 1244 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Reinstallation of software. | We took access to the system and successfully completed the installation and configuration of software. Now data is live on LOCAL+CPCB+MPCB. | We took access to the system and successfully completed the installation and configuration of software. Now data is live on LOCAL+CPCB+MPCB. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24024713916 | 2024-05-22 | 2024-05-24 | Cust 1233 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data uploaded issue. | we took access of your system and found that the Data is not communicating till the Local system hence this issue needs to be checked from the hardware side. | we took access of your system and found that the Data is not communicating till the Local system hence this issue needs to be checked from the hardware side. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24025023917 | 2024-05-23 | 2024-05-24 | Cust 795 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB and MPCB Server. | We took access to the system and now data is live on CPCB nad MPCB. | We took access to the system and now data is live on CPCB nad MPCB. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24025443918 | 2024-05-24 | 2024-05-27 | Cust 1284 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to CPCB and MPCB server | Data is now live on the Local, CPCB and MPCB servers. | Data is now live on the Local, CPCB and MPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24025553919 | 2024-05-24 | 2024-05-27 | Cust 809 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB + MPCB data upload issue. | the Data is live on the CPCB and MPCB servers and the same you can check in the screenshot below. | the Data is live on the CPCB and MPCB servers and the same you can check in the screenshot below. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24028193920 | 2024-05-28 | 2024-05-31 | Cust 795 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data breaking on MPCB Server. | Updated EnviroconectDAS and now data is live on server. | Updated EnviroconectDAS and now data is live on server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24028223921 | 2024-05-28 | 2024-05-30 | Cust 1002 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on MPCB Server. | Took access of the system and now data is updating on MPCB Server. | Took access of the system and now data is updating on MPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24028773922 | 2024-05-29 | 2024-05-31 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Shree Dudhaganga Vedganga Sahakari ==CPCB new configuration of STACK2 60TPH. | We took access to the system and configured the new Stack2 60TPH. | We took access to the system and configured the new Stack2 60TPH. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24029023923 | 2024-05-29 | 2024-05-31 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB configuration of Shree Dudhaganga Vedganga SSK | Took access to the system and now data is live on MPCB Server. | Took access to the system and now data is live on MPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24029473924 | 2024-05-30 | 2024-06-03 | Cust 633 | PUNE | Analytical - CHAKAN | CLOSED | APPLICATION ISSUE | Aqua4Trans | Turbidity Input card of Aqua4Trans(1) | CALIBRATE THE ANALYSER | Turbidity and pH analyser commissioned. | Turbidity and pH analyser commissioned. | NO | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24030963925 | 2024-05-31 | 2024-06-24 | Cust 1332 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | replacement of the spare parts | AMC visit done. Necessary maintenance activities carried out for Mutlipara. | AMC visit done. Necessary maintenance activities carried out for Mutlipara. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24032063926 | 2024-06-03 | 2024-06-08 | Cust 811 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4Trans(1) | connection and calibration done satisfied with the values | Commissioning work of aqua4trans sensor. | connection and calibration done satisfied with the values. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFPN24033013927 | 2024-06-05 | 2024-06-05 | Cust 803 | PUNE | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | Aqua4Trans | Aqua4Trans(1) | Installation and erection guidelines for Aqua4trans transmitter | Installation guidelines given to customer. | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN24033013927 | 2024-06-05 | 2024-06-05 | Cust 803 | PUNE | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | DO Transmitters | optex series(optex series) | Installation and erection work WATER QUALITY MONITORING SYSTEM,OPTICAL DO SENSORS FOR EXPLOSIVE ENVIRONMENTS, MODEL: VISIFERM MA 120 H3, WITH ODO CAP H3, RANGE: 4 PPB TO 25 PPM, INTERFACE: BLUETOOTH& HART, ATEX APPROVAL | Installation guidelines given to customer. | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN24033013927 | 2024-06-05 | 2024-06-05 | Cust 803 | PUNE | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | AQUADO Analyzer | Aquamax system(1) | Installation and erection work guidelines | Installation guidelines given to customer. | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN24033013927 | 2024-06-05 | 2024-06-05 | Cust 803 | PUNE | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | Conductivity Electrode | Forbes Marshall Cell Model 8312, K=1(FM 8312-1.5) | Installation and erection work. | Installation guidelines given to customer. | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN24033013927 | 2024-06-05 | 2024-06-05 | Cust 803 | PUNE | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | Ph Electrode | Polilyte Plus VP 120(HT242 428) | POLILYTE PLUS PHI VP 120 PT100 installation and guidelines | Installation guidelines given to customer. | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN24033253928 | 2024-06-05 | 2024-06-07 | Cust 1320 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Eversense software configuration and MPCB CPCB Connectivity need to be done. | We took access to the system and successfully completed the software configuration and now data is live on CPCB and MPCB Server. | We took access to the system and successfully completed the software configuration and now data is live on CPCB and MPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24033623929 | 2024-06-06 | 2024-06-11 | Cust 987 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPDMS APPROVAL IS PENDING FROM CUSTOMER AND THATS WHY ARE DATA ARE NOT RECEIVED | Here there is a Hardware issue we are closing this complaint. | Here there is a Hardware issue we are closing this complaint. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFPN24034153930 | 2024-06-07 | 2024-06-17 | Cust 640 | PUNE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | uptime offer discussion & spare offer | discussion on spare offer | completed the work | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24034193931 | 2024-06-07 | 2024-06-26 | Cust 811 | PUNE | Analytical - CHAKAN | CLOSED | DIGITAL CONEECTIVITY | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Connectivity is checked with black box | commissioning done | data arrived after commissioning | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFPN24035283932 | 2024-06-10 | 2024-06-14 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Display board configuration need to be done. | We have done the necessary configuration to post JSon on the provided IP (192.168.0.26), | We have done the necessary configuration to post JSon on the provided IP (192.168.0.26), | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFPN24036343933 | 2024-06-12 | 2024-06-24 | Cust 1294 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | LCD of Multiparameter Display(1) | Display error | Display replaced . | Display replaced . | NO | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24036453934 | 2024-06-13 | 2024-06-17 | Cust 1332 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Check the device wether in healthy condition | SUGGESTED THE MAINTANENCE | complete work | NO | Paid | Core | N |  | Analytical - CHAKAN
AFPN24036613935 | 2024-06-13 | 2024-06-14 | Cust 1035 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on CPCB server. | As of now we are closing this complaint. We will reopen this complaint once we get access. | As of now we are closing this complaint. We will reopen this complaint once we get access. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24036643936 | 2024-06-13 | 2024-06-14 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on MPCB server. | Data is live on CPCB + Local + MPCB server. | Data is live on CPCB + Local + MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24036693937 | 2024-06-13 | 2024-06-14 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on Local + MPCB server. | Now data is live on CPCB + Local + MPCB server. | Now data is live on CPCB + Local + MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24036893938 | 2024-06-13 | 2024-06-17 | Cust 1259 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | check the device condition and the values are in the limit | ISSUE NOTICED SUGGESTED REQUIREMENTS | complete work | NO | Paid | Core | N |  | Analytical - CHAKAN
AFPN24037183939 | 2024-06-14 | 2024-06-17 | Cust 1333 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Electrode | Polilyte HT 120(HT238 431) | Commissioning OF Ph sensor. | Commissioning Has been done. | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN24037183939 | 2024-06-14 | 2024-06-17 | Cust 1333 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Commissioning Of PH Meter . | Commissioning Has been done. | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN24037273940 | 2024-06-14 | 2024-06-17 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB connectivity needs to be done. | We took access of the system and completed CPCB connectivity and now data is live on CPCB server. | We took access of the system and completed CPCB connectivity and now data is live on CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24037283941 | 2024-06-14 | 2024-06-17 | Cust 1334 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on CPCB + MPCB server. | We have completed the display board configuration and now the data is live on the server. | We have completed the display board configuration and now the data is live on the server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24038003942 | 2024-06-17 | 2024-06-19 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | KSPCB connectivity needs to be done. | Today we took access to your system and completed KSPCB connectivity, and now data is live on KSPCB server. | Today we took access to your system and completed KSPCB connectivity, and now data is live on KSPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24039073943 | 2024-06-18 | 2024-07-09 | Cust 795 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data uploaded issue. | issue resolved . mail received from customer is attached. | issue resolved . mail received from customer is attached. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24039623944 | 2024-06-19 | 2024-06-20 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB connectivity needs to be done. | Now data is live on the CPCB server. | Now data is live on the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24039633945 | 2024-06-19 | 2024-06-21 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB camera connectivity needs to be done. | Industry is closed due to seasonal shutdown & LAN network is completely changed at site | Industry is closed due to seasonal shutdown &amp;amp; LAN network is completely changed at site | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24039783946 | 2024-06-19 | 2024-06-26 | Cust 803 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | parts replaced | after replacing the spare | completed the work | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24040093947 | 2024-06-20 | 2024-06-20 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | showing different values on MPCB & CPCB server | Now data is showing the same values on CPCB + Local + MPCB server. | Now data is showing the same values on CPCB + Local + MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24040843948 | 2024-06-22 | 2024-06-24 | Cust 1301 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4Trans(1) | commissioning of TDS | commissioning done | work completed | NO | Paid | Core | N |  | Analytical - CHAKAN
AFPN24043033949 | 2024-06-26 | 2024-06-29 | Cust 1266 | PUNE | Analytical - CHAKAN | CLOSED | SALES PROMOTION | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | receievd call from Mr. Desai to plan visit to clear the all balance outstanding cases and payments | visited the site all the commercial issue are cleared | visited the site all the commercial issue are cleared | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24043043950 | 2024-06-26 | 2024-06-29 | Cust 974 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB CONFIGURATION NEEDS TO BE AS PER THE NEW ANALYSER. | We have made the modifications and now the data is live on the server. | We have made the modifications and now the data is live on the server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24043433951 | 2024-06-27 | 2024-06-27 | Cust 1335 | PUNE | Steam System Division | CLOSED | QUALITY PROBLEM | PRV - MEASURETROL | FM MEASURETROL(SIZE 050NB) | 17Nos ELE - Have to change to upgraded design | Reference site visit by Jaykumar, Jitesh Patil, Parvez Momin and Vishwanath Deshpande - Upgraded the PCB with Zener Diode | Reference site visit by Jaykumar, Jitesh Patil, Parvez Momin and Vishwanath Deshpande - Upgraded the PCB with Zener Diode | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFPN24043473952 | 2024-06-27 | 2024-06-29 | Cust 1237 | PUNE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Aqua4Trans | Aqua4Trans(1) | sensor connection has broken , fixed | sensor wiring needs to be noticed | calibration has done | NO | Paid | Core | N |  | Analytical - CHAKAN
AFPN24043483953 | 2024-06-27 | 2024-06-27 | Cust 1335 | PUNE | Steam System Division | CLOSED | QUALITY PROBLEM | PRV - MEASURETROL | FM MEASURETROL(SIZE 100NB) | 2)Found LED wire pressed in between LED plate and gasket. 3)Found MS bolting on top cover. 4)2 No. PCB get short and failed due to water leakage. Replaced by new PCB. 8)AFR guages short supply and alternate AFR supplied. 9)Amphenol Connectors was broken. 10) Internal connectors was broken. 11) P1, P2 cable was broken and crimping problem. 12) Screw for PCB mounting spacer was missing. 13) Piezo-Pilot valve get failed due to water leakage. | Reference site visit by Jaykumar, Jitesh Patil, Parvez Momin and Vishwanath Deshpande - REPLACEMENT SCHEDULED | Reference site visit by Jaykumar, Jitesh Patil, Parvez Momin and Vishwanath Deshpande - REPLACEMENT SCHEDULED | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFPN24044223954 | 2024-06-28 | 2024-06-29 | Cust 809 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB + MPCB data upload issue. | we took access of the black box, and now the issue has been resolved and data is live on CPCB + MPCB server. | we took access of the black box, and now the issue has been resolved and data is live on CPCB + MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24046963955 | 2024-07-01 | 2024-08-29 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data upload issue. | As of now we are closing this complaint and will revert you once we get access from customer. | As of now we are closing this complaint and will revert you once we get access from customer. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24047543956 | 2024-07-02 | 2024-07-04 | Cust 1262 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Data is not updating to CPCB and MPCB servers. | As of now we are closing this complaint. we will reopen this once the issue is resolved from hardware side. | As of now we are closing this complaint. we will reopen this once the issue is resolved from hardware side. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24048443957 | 2024-07-04 | 2024-07-05 | Cust 803 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Conductivity Input card of Aqua4trans(1) | Commissioning of Aqua4trans transmitter | Commissioning has been done successfully. | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN24048443957 | 2024-07-04 | 2024-07-05 | Cust 803 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Electrode | Forbes Marshall Cell Model 8312, K=1(FM 8312-1.5) | Commissioning of Conductivity sensor. | Commissioning has been done successfully. | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN24048443957 | 2024-07-04 | 2024-07-05 | Cust 803 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Electrode | Polilyte Plus VP 120(HT242 428) | Commissioning of Ph electrode. | Commissioning has been done successfully. | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN24048443957 | 2024-07-04 | 2024-07-05 | Cust 803 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | Visiferm DO sensor(1) | Visiferm DO commission. | Commissioning has been done successfully. | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN24049043958 | 2024-07-05 | 2024-07-08 | Cust 1259 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | CPU with LCD of Multiparameter Display(1) | replacement of the display | commissioned the pcx display verified the device , working properly got the data device in healthy condition | commissioned and verified device | NO | Paid | Core | N |  | Analytical - CHAKAN
AFPN24050543959 | 2024-07-08 | 2024-07-15 | Cust 811 | PUNE | Analytical - CHAKAN | CLOSED | AMC | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | data breakdown | network settings has been changed physically , The blackbox has been uploaded with the new format of data files | device is working properly after changing network settings. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFPN24051463960 | 2024-07-10 | 2024-07-15 | Cust 1336 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Showing zero readings | work done problem has been solved | problem solved | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24052003961 | 2024-07-11 | 2024-07-15 | Cust 692 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Detector low | Multiple problem solved | problem solved | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24052353962 | 2024-07-11 | 2024-07-12 | Cust 803 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB+MPCB data upload issue. | Now hardware side issue is resolve and data is live on CPCB server. | Now hardware side issue is resolve and data is live on CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24053723963 | 2024-07-15 | 2024-07-19 | Cust 1337 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | LOCAL+MPCB data uploading issue. | We took access to the system and successfully completed the LOCAL+MPCB Configuration and now data is live. | We took access to the system and successfully completed the LOCAL+MPCB Configuration and now data is live. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24053763964 | 2024-07-15 | 2024-07-16 | Cust 809 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data mismatch issue. | We took access of the black box and observed that data coming till the black box and CPCB server values are matching for the COD parameter. | We took access of the black box and observed that data coming till the black box and CPCB server values are matching for the COD parameter. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24054953965 | 2024-07-17 | 2024-08-14 | Cust 1241 | PUNE | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Unable to upgrade the patch | FMBB dispatched RF OA NO.: 5525800246, MARK EXPRESS PRIVATE LIMITED vide LR No - 210104659 LR Dt - 22-JUL-2024 | FMBB dispatched RF OA NO.: 5525800246, MARK EXPRESS PRIVATE LIMITED vide LR No - 210104659 LR Dt - 22-JUL-2024 | YES | No Trip | Core | N |  | DCS - CHAKAN
AFPN24054973966 | 2024-07-17 | 2024-07-19 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data upload issue. | Data is live on the MPCB server. | Data is live on the MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24055483967 | 2024-07-18 | 2024-07-22 | Cust 809 | PUNE | Analytical - CHAKAN | CLOSED | DIGITAL CONEECTIVITY | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | data connectivity loss | problem has been solved | problem solved | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24056073968 | 2024-07-19 | 2024-07-19 | Cust 623 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera streaming issue on CPCB server. | Now the camera streaming is showing live on the CPCB server. | Now the camera streaming is showing live on the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24056203969 | 2024-07-19 | 2024-07-23 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data uploading issue. | As of now we are closing this complaint, we will reopen this once the issue is resolved from the hardware side. | As of now we are closing this complaint, we will reopen this once the issue is resolved from the hardware side. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24058053970 | 2024-07-23 | 2024-07-29 | Cust 1318 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Getting incorrect data | problem has been solved | problem solved | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24059153971 | 2024-07-25 | 2024-07-29 | Cust 1239 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | location of analyzer | problem has been solved | problem solved | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24059203972 | 2024-07-25 | 2024-07-26 | Cust 1338 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on CPCB Server. | We are closing this complaint as there is PC version issue from customer end we will reopen the complaint once issue resolved from customer end. | We are closing this complaint as there is PC version issue from customer end we will reopen the complaint once issue resolved from customer end. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24060073973 | 2024-07-26 | 2024-07-29 | Cust 811 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | data doesn't show on cpcb | work has been done | problem solved | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24063203974 | 2024-07-30 | 2024-08-02 | Cust 1338 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Cameras not playing on CPCB Server. | As of now we are closing the complaint we will reopen this complaint once issue is resolved from customer end | As of now we are closing the complaint we will reopen this complaint once issue is resolved from customer end | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24064473975 | 2024-07-31 | 2024-08-06 | Cust 1242 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Maintenance under AMC visit | AMC maintenance carried out for multiparameter analyser. | AMC maintenance carried out for multiparameter analyser. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24066613976 | 2024-08-05 | 2024-08-06 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on CPCB and MPCB server. | Now the data is live on CPCB & MPCB servers. | Now the data is live on CPCB &amp;amp; MPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24066913977 | 2024-08-05 | 2024-08-07 | Cust 1256 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Data not updating on Local server. | We took access to the system and now data is updating on CPCB and MPCB Server. | We took access to the system and now data is updating on CPCB and MPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24067213978 | 2024-08-06 | 2024-08-07 | Cust 712 | PUNE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | AQUADO Analyzer | AQUAMAX DOT(1) | Calibration of DOT installed at ETP. | Calibration has been done successfuly. Training has been done successfully. | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN24067213978 | 2024-08-06 | 2024-08-07 | Cust 712 | PUNE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | AQUADO Analyzer | AQUAMAX DOT(1) | Calibration of DOT installed at ETP. | Calibration of the sensor's were done successfully. | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN24067283979 | 2024-08-06 | 2024-08-08 | Cust 1338 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on CPCB Server. | We took access to the system and camera ports were changed due to which camera were not playing. We took access to the system and now both cameras are live on CPCB Server. | We took access to the system and camera ports were changed due to which camera were not playing. We took access to the system and now both cameras are live on CPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24067953980 | 2024-08-07 | 2024-08-09 | Cust 1339 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on MPCB server and camera streaming is not live on CPCB server. | We accessed the system and found that the camera at GSP Outlet Product is still not streaming outside the network.so issue is from the customer side . so we are closing this complaint and we will reopen this complaint once we get revert from customer side. | We accessed the system and found that the camera at GSP Outlet Product is still not streaming outside the network.so issue is from the customer side . so we are closing this complaint. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24068213981 | 2024-08-08 | 2024-08-10 | Cust 795 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Amc visit | annual maintenance & calibration | completed the work | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24068343982 | 2024-08-08 | 2024-09-04 | Cust 803 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Display board configuration needs to be done. | Customer has not release payment yet, so as of now we are closing this complaint. | Customer has not release payment yet, so as of now we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24070013983 | 2024-08-12 | 2024-08-14 | Cust 811 | PUNE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | values are fluctuating | visit done. Necessary maintenance activities carried out for Multipara,. | visit done. Necessary maintenance activities carried out for Multipara,. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24070023984 | 2024-08-12 | 2024-08-14 | Cust 1339 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera streaming issue on CPCB server. | There is a technical issue with the Biocompost 2 camera from the customer?s side. So we are closing this complaint for now. We will reopen it once we receive a response from the customer. | There is a technical issue with the Biocompost 2 camera from the customers side. So we are closing this complaint for now. We will reopen it once we receive a response from the customer. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24070443985 | 2024-08-13 | 2024-08-19 | Cust 1232 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | product fail | maintenance has been done | completed the work | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24071063986 | 2024-08-14 | 2024-08-19 | Cust 1340 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Spares needs to be installed and commissioned. | maintainance has been done,found multiple issue in the analyser , need to send the analyser to factory for repairing | completed the work | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24071693987 | 2024-08-16 | 2024-08-19 | Cust 517 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC of the device | AMC visit done. Necessary maintenance activities carried out for Mutlipara. | AMC visit done. Necessary maintenance activities carried out for Mutlipara. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24073803988 | 2024-08-20 | 2024-08-26 | Cust 1237 | PUNE | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Proactive visit for renewal of EUTs | Data has been arrived after changing the settings amc has conducted | data has arrived after the configuration | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24074763989 | 2024-08-22 | 2024-08-26 | Cust 1237 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | display not working | The device maintance has been done and changed device configuration for live data | Device maintance has checked and changed device settings for online communication. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFPN24075043990 | 2024-08-22 | 2024-08-26 | Cust 1260 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC of MP analyser and connectivity with Goa PCB | visited the site complete AMC activity for MP analyser , MOM attached | visited the site complete AMC activity for MP analyser , MOM attached | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24075483991 | 2024-08-23 | 2024-09-02 | Cust 952 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | pH meter problem | AS transmitter was not powering on asked customer to send factory for further rectification. | AS transmitter was not powering on asked customer to send factory for further rectification. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFPN24075563992 | 2024-08-23 | 2024-08-26 | Cust 1239 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | commissioning of the device. | The Commissioning of the device has been done. | Commissioning has been done successfully | YES | Paid | Core | N |  | Analytical - CHAKAN
AFPN24075983993 | 2024-08-24 | 2024-08-27 | Cust 1237 | PUNE | Analytical - CHAKAN | CLOSED | DIGITAL CONEECTIVITY | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | data connectivity | FM cloud server connectviity done for MP, Dust and FLow meter. | FM cloud server connectviity done. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24077183994 | 2024-08-26 | 2024-08-27 | Cust 1320 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Eversense software installation . | As discussed with the customer, there is an internet issue on their end. The customer will provide access to the distillery unit on Friday. So as of now we are closing the complaint and will reopen once we get the access. | The customer will provide access to the distillery unit on Friday. So as of now we are closing the complaint and will reopen once we get the access. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24077653995 | 2024-08-27 | 2024-09-10 | Cust 526 | PUNE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | CALIBRATION ISSUE & CALL SUPPORT TO CUSTOMER | conductivity value is not shown in meter | completed the work | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24078493996 | 2024-08-28 | 2024-08-30 | Cust 809 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Customer asked reports of the Dated 27/02/2024.of CPCB+MPCB | We have shared CPCB+MPCB REPORTS to the customer of the Dated 27/02/2024. | We have shared CPCB+MPCB REPORTS to the customer of the Dated 27/02/2024. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24078603997 | 2024-08-28 | 2024-08-29 | Cust 1320 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Data upload issue on MPCB server | We accessed the system and observed that data is not coming till the local system for the ETP Outlet, as (Checked on MODSCAN) So issue is from the hardware side so we are closing this complaint. | We accessed the system and observed that data is not coming till the local system for the ETP Outlet, as (Checked on MODSCAN) So issue is from the hardware side so we are closing this complaint. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24078943998 | 2024-08-29 | 2024-09-02 | Cust 904 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | PH2 Filter(1) | PH1 & PH 2 values are zero | Suggested to send the analyser at Pune for further testing. | Closed | NO | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN24079773999 | 2024-08-30 | 2024-09-04 | Cust 803 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Electrode | Easyferm Plus 120(HT238 643) | Commissioning of PH sesnor | Commissioning has been done. | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN24080344000 | 2024-08-31 | 2024-09-02 | Cust 1341 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | received call from Parshant Rajbhoj to support customer for the pH and conductivity transmitter | Visited the site and commissioning/calibration verification done for pH and Conductivity. | Visited the site and commissioning/calibration verification done for pH and Condictivity. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24083194001 | 2024-09-03 | 2024-09-06 | Cust 809 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Internet configuration needs to be done. | we took access of the black box and now issue has been resolved and data is live on CPCB + MPCB server. | we took access of the black box and now issue has been resolved and data is live on CPCB + MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24083964002 | 2024-09-05 | 2024-09-06 | Cust 795 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB Server. | There was customer side issue data is live on CPCB Server. | There was customer side issue data is live on CPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24084214003 | 2024-09-05 | 2024-09-09 | Cust 692 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | The Spares has replaced with old one | AMC visit done. Necessary maintenance activities carried out for Mutlipara. | AMC visit done. Necessary maintenance activities carried out for Mutlipara. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24084504004 | 2024-09-06 | 2024-09-06 | Cust 1313 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Demo/Testing Enviroconnect software installation needs to be done. | the demo was successful on their PC & we addressed all their queries regarding this local software. | the demo was successful on their PC &amp;amp; we addressed all their queries regarding this local software. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24085894005 | 2024-09-09 | 2024-09-09 | Cust 1239 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | BOD readings high. pH readings Low | Online support given and BOD and pH readings process calibration done. | Online support given and BOD and pH readings process calibration done. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFPN24085944006 | 2024-09-09 | 2024-09-12 | Cust 1320 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Data upload issue on CPCB & MPCB servers. | As of now we are closing this complaint and will reopen it once we get system access. | As of now we are closing this complaint and will reopen it once we get system access. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24086544007 | 2024-09-10 | 2024-09-18 | Cust 803 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Maintenance | visit done. Necessary maintenance activities carried out for Multipara,. | visit done. Necessary maintenance activities carried out for Multipara,. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFPN24086844008 | 2024-09-11 | 2024-09-13 | Cust 1239 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB AND MPCB connectivity needs to be done. | For now we are closing this complaint. | For now we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24087254009 | 2024-09-12 | 2024-09-12 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on MPCB server. | We took access to the system and now data is live on the MPCB server. | We took access to the system and now data is live on the MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24087744010 | 2024-09-13 | 2024-09-13 | Cust 731 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera is not live on the CPCB server | Camera is streaming live on the CPCB server. | Camera is streaming live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24087914011 | 2024-09-13 | 2024-09-13 | Cust 623 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | PM data not updating on CPCB Server. | Now the issue has been resolved and data is live on the CPCB and HSPCB server. | Now the issue has been resolved and data is live on the CPCB and HSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24087974012 | 2024-09-13 | 2024-09-20 | Cust 1301 | PUNE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | pH. ORP, Conductivity caloibration | Visit done calibration carried out. | Visit done calibration carried out. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFPN24088844013 | 2024-09-16 | 2024-09-19 | Cust 795 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data uploading issue | the data is live on the CPCB server. | the data is live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24090074014 | 2024-09-18 | 2024-09-20 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Display Board configuration needs to be done. | As of now we are closing this complaint and will reopen it once we get system access. | As of now we are closing this complaint and will reopen it once we get system access. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24090094015 | 2024-09-18 | 2024-09-18 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on CPCB + Local + MPCB servers. | We took access to the system and observed that data is not coming till the local system (Checked on MODSCAN). So issue is from the hardware side ,so we are closing this complaint. | We took access to the system and observed that data is not coming till the local system (Checked on MODSCAN). So issue is from the hardware side ,so we are closing this complaint. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24090764016 | 2024-09-19 | 2024-09-28 | Cust 623 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Common Chimney Dust - Device data breaking on CPCB Server. | Now the issue has been resolved and we have observed that for common chimney and Stack 2 there were two moxas connected simultaneously which is conflicting as we can not connect two masters parallelly. Hence to resolve this issue we have removed one moxa and now the data is continuously uploading. | Issue has been resolved. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24090784017 | 2024-09-19 | 2024-09-20 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on MPCB Server. | As of now we are closing this complaint and will reopen it once we get access. | As of now we are closing this complaint and will reopen it once we get access. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24091234018 | 2024-09-20 | 2024-09-27 | Cust 793 | PUNE | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Installation & Erection work. | Installation and erection work has been done. | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN24091904019 | 2024-09-23 | 2024-09-27 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data mismatch | Issue has been resolved and data is now live on the CPCB server. | Issue has been resolved and data is now live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24092804020 | 2024-09-24 | 2024-09-25 | Cust 1284 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on MPCB server. | We took access to the system and now data is live on MPCB Server. | We took access to the system and now data is live on MPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24092824021 | 2024-09-24 | 2024-09-27 | Cust 517 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Display Board data upload issue. | we have resolved the Display Board data transmission issue and now the data is uploading to the Display Board. | we have resolved the Display Board data transmission issue and now the data is uploading to the Display Board. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24092844022 | 2024-09-24 | 2024-09-28 | Cust 1242 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC has done with the replacement of new PH sensor | AMC done, PH sensor replaced | PH Sensor Replaced and AMc has been Done, Data is not Live Given instruction to thier IT team to manage the setup. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFPN24092914023 | 2024-09-24 | 2024-09-27 | Cust 1270 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Data is not uploading to CPCB and MPCB server. | As of now we are closing this complaint. We will reopen this once the issue is resolved from the hardware side. | As of now we are closing this complaint. We will reopen this once the issue is resolved from the hardware side. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24092954024 | 2024-09-24 | 2024-10-04 | Cust 803 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Analyser communication with display | viisted the iste addressed software issue related to the display file patch update | viisted the iste addressed software issue related to the display file patch update | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24093494025 | 2024-09-25 | 2024-09-28 | Cust 1242 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to the Local server. | As of now we are closing this complaint. We will reopen this once we get access to the system. | As of now we are closing this complaint. We will reopen this once we get access to the system. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24093514026 | 2024-09-25 | 2024-09-28 | Cust 1264 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB + MPCB + software reinstallation needs to be done. | As of now we are closing this complaint and will reopen it once we get system access. | As of now we are closing this complaint and will reopen it once we get system access. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24093524027 | 2024-09-25 | 2024-09-25 | Cust 1284 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on Local + MPCB servers | We took access to the system and now data is live on the Local + MPCB servers. | We took access to the system and now data is live on the Local + MPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24093584028 | 2024-09-25 | 2024-09-28 | Cust 803 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Display Data issue. | Display Board files are creating in their respective location on the customer's Local PC ( SS is attached for your reference) but the data is not displaying on the board so that needs to be checked at the site. Hence we are closing this compliant. | Display Board files are creating on the customer's Local PC but the data is not displaying on the board so that needs to be checked at the site.Hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24093834029 | 2024-09-26 | 2024-09-28 | Cust 803 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Display board configuration needs to be done. | Display Board files are creating in their respective location on the customer's Local PC ( SS is attached for your reference) but the data is not displaying on the board so that needs to be checked at the site. Hence we are closing this compliant. | Display Board files are creating on the customer's Local PC but the data is not displaying on the board so that needs to be checked at the site. Hence we are closing this compliant. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24094114030 | 2024-09-26 | 2024-10-04 | Cust 564 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Conductivity Tranmitter Smartpro | Conductivity Electrode Cable(Cond-Elec-cbl) | conductivity transmitter showing zero reading | visited the site addressed all the issues of ph and conductivity transmitter comp logged for quality purpose | visited the site addressed all the isssues of ph and conductivtu transmitter comp logge dfor quaity purpose | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24094114030 | 2024-09-26 | 2024-10-04 | Cust 564 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | conductivity transmitter showing zero reading | visited the site addressed all the issues of ph and conductivity transmitter comp logged for quality purpose | visited the site addressed all the isssues of ph and conductivtu transmitter comp logge dfor quaity purpose | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24094404031 | 2024-09-27 | 2024-09-28 | Cust 692 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB + RSPCB data upload issue. | We took the access to your system and found that EnviroconnectDAS service is stopped on your local PC. | We took the access to your system and found that EnviroconnectDAS service is stopped on your local PC. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24094504032 | 2024-09-27 | 2024-09-28 | Cust 811 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data upload issue. | we took access of the black box now the internet side issue is resolved and data is live on CPCB + MPCB server. | we took access of the black box now the internet side issue is resolved and data is live on CPCB + MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24095074033 | 2024-09-28 | 2024-09-28 | Cust 1278 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Blackbox Commissioning | we have completed the BLACK BOX Configuration + CPCB connectivity for the Belgaum Sugar site so now the data is live on the CPCB server. | we have completed the BLACK BOX Configuration + CPCB connectivity for the Belgaum Sugar site so now the data is live on the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24096444034 | 2024-09-30 | 2024-10-03 | Cust 1259 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB DATA UPLOAD FOR SKODA AUTO Volkswagen India Private Limited SKODA AUTO Volkswagen India Private Limited (formerly known as Volkswagen India Private Limited) | now the issue has been resolved and the data is live on the MPCB server. | now the issue has been resolved and the data is live on the MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24096454035 | 2024-09-30 | 2024-11-05 | Cust 1331 | PUNE | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | Silica Analyser | Hach Silica 5500 series(Hach Silica 5500 series) | AWA Silica sequencer issue | Material carried by FM engg. and installed at site. | Material carried by FM engg. and installed at site. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFPN24098014036 | 2024-10-02 | 2024-10-09 | Cust 793 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning of Multiparameter Analyser | Commission has been done successfully. | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN24098184037 | 2024-10-03 | 2024-10-16 | Cust 793 | PUNE | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Assembling was not proper | I have made it proper and then commissioned the instrument. | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN24098364038 | 2024-10-03 | 2024-10-04 | Cust 651 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB + MPCB connectivity needs to be done. | we took access of the black box and now the hardware side issue is resolved and data is live on the CPCB server and the same is verified by you. | we took access of the black box and now the hardware side issue is resolved and data is live on the CPCB server and the same is verified by you. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24098664039 | 2024-10-04 | 2024-11-18 | Cust 809 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB + MPCB DATA UPLOAD ISSUE. | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210119708 LR Dt - 11-NOV-2024 | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210119708 LR Dt - 11-NOV-2024 | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24098944040 | 2024-10-04 | 2024-10-10 | Cust 517 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC visit | AMC visit done. Necessary maintenance activities carried out for Mutlipara. | AMC visit done. Necessary maintenance activities carried out for Mutlipara. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24098954041 | 2024-10-04 | 2024-10-10 | Cust 1326 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Spares needs to installed | visit done. Necessary maintenance activities carried out for Multipara, with spare replacement | visit done. Necessary maintenance activities carried out for Multipara, with spare replacement | NO | Paid | Core | N |  | Analytical - CHAKAN
AFPN24098974042 | 2024-10-04 | 2024-10-10 | Cust 1319 | PUNE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | AQUADO Analyzer | OXYSENS(OXYSENS) | commissioning complete | commissining complete with Do analyser | completed the work | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24098984043 | 2024-10-04 | 2024-10-07 | Cust 1342 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | new transmiitter commisining | Commissioning done and working properly. | commission has been done working properly customer satisfied | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24099664044 | 2024-10-07 | 2024-10-10 | Cust 1270 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Re-installation of Eversense software. | As discussed with the customer, they will provide system access tomorrow. Therefore, we are closing this complaint for now. We will reopen it once we receive system access. | As discussed with the customer, they will provide system access tomorrow. Therefore, we are closing this complaint for now. We will reopen it once we receive system access. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24099684045 | 2024-10-07 | 2024-10-10 | Cust 991 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | DisplayBoard Configuration | As of now we are closing this complaint and will reopen it once twe get access from customer end. | As of now we are closing this complaint and will reopen it once twe get access from customer end. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24100524046 | 2024-10-08 | 2024-10-21 | Cust 1343 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | LCD of Multiparameter Display(1) | received confirmation from customer to attend the AMV visit | visited the site addressed all the problem of the MP , display needs repair for CMOS battery | visited the site addressed all the problem of the MP , display needs repair for CMOS battery | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24100654047 | 2024-10-09 | 2024-10-14 | Cust 564 | PUNE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Smartpro-con) | Value mismatch with sample need to calibrate with solution at site | PROBLEM SOLVED | completed the work | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24101434048 | 2024-10-10 | 2024-10-11 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | TSS data mismatch issue : Sri Basaveshwara Sugars Ltd (Sugar) | we took access to your system and made some modifications from configuration end, now TSS data is matching with the actual readings and the same is verified by you. | we took access to your system and made some modifications from configuration end, now TSS data is matching with the actual readings and the same is verified by you. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24101464049 | 2024-10-10 | 2024-10-11 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Display board configuration needs to be done. | Today we took access of your system and observed that the below IP address is not accessible in your network. so for now we are closing this complaint. | Today we took access of your system and observed that the below IP address is not accessible in your network. so for now we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24101504050 | 2024-10-10 | 2024-11-05 | Cust 1326 | PUNE | Analytical - CHAKAN | CLOSED | MANUFACTURING ISSUES | MULTIPARAMETER ANALYSER | CPU with LCD of Multiparameter Display(1) | Need to supply display with door panel | Short Supply material dispatched. | Short Supply material dispatched. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFPN24101534051 | 2024-10-10 | 2024-10-15 | Cust 1326 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | New spares replacement | visit done. Necessary maintenance activities carried out for Multipara,. | visit done. Necessary maintenance activities carried out for Multipara,. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24101854052 | 2024-10-11 | 2024-10-20 | Cust 1239 | PUNE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | The Values are showing zer0 | work done | problem solved | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24102684053 | 2024-10-14 | 2024-10-17 | Cust 1270 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | CPCB + MPCB configuration needs to be done. | We are closing this complaint for now. We will reopen it once we have access to a Windows-based PC system. | We are closing this complaint for now. We will reopen it once we have access to a Windows-based PC system. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24102844054 | 2024-10-14 | 2024-10-15 | Cust 731 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on CPCB server. | We took access to the system and downgraded the firefox version to 128.0.3 and now camera is live on CPCB server. | We took access to the system and downgraded the firefox version to 128.0.3 and now camera is live on CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24103384055 | 2024-10-15 | 2024-10-16 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Display board configuration needs to be done. | we took access of your system and completed Display Board configuration, now the JSON is generating | we took access of your system and completed Display Board configuration, now the JSON is generating | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24103634056 | 2024-10-15 | 2024-10-20 | Cust 1232 | PUNE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Tss became high detector fluctuating | work done | problem solved | NO | Paid | Core | N |  | Analytical - CHAKAN
AFPN24103804057 | 2024-10-15 | 2024-10-18 | Cust 617 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data breaking issue on MPCB server. | We took access to the system and updated the java and Patch. We kept the system under observation. | We took access to the system and updated the java and Patch. We kept the system under observation. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24104654058 | 2024-10-17 | 2024-10-21 | Cust 651 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | RS485 ISOLATOR(1) | received confirmation from Division with Scanner to update on the CX2K software version update and to check communication on Modscan combinely for MP and FM | visited the site MP software upgradation completed, further connectivity withe CPCB done | visited the site MP software upgradation completed, further connectivity withe CPCB done | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24104754059 | 2024-10-17 | 2024-10-18 | Cust 731 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on CPCB server. | We took access to the system and downgraded the version and now camera is live. | We took access to the system and downgraded the version and now camera is live. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24105064060 | 2024-10-18 | 2024-10-25 | Cust 1259 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Maintenance under EUT services | AMC visit done. Necessary maintenance activities carried out for Mutlipara. | AMC visit done. Necessary maintenance activities carried out for Mutlipara. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24105214061 | 2024-10-18 | 2024-10-21 | Cust 1271 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | FM make RS485 card(1) | AMC visit for MP and connectivity | viisted the ste AMC activity completed succesfully , mom atatched | viisted the ste AMC activity completed succesfully , mom atatched | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24105494062 | 2024-10-19 | 2024-10-27 | Cust 1344 | PUNE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | analyser calibration and data connectivity | AMC of the analyser | completed the work | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24106104063 | 2024-10-21 | 2024-11-18 | Cust 1345 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Electrode | Polilyte plus H S8 120(1) | At ambient temp. pH sensor works normally and as we are raising the temp. at 60 deg. C, the inner white solution starts leaking from the probe in droplets. After dropping in the buffer solution it forms a crystal-like structure. | The sensor Behaviour is normal as confirmed by the manufacturer. The life will depend on temperature & process conditions. Alternatively they may opt for other sensor with ceramic diaphragm but it may stuck & be sluggish | The sensor Behaviour is normal as confirmed by the mfg. The life will depend on temp. process conditions .They may opt for other sensor with ceramic diaphragm but it may stuck or be sluggish | YES | No Trip | Core | Y | Product / equipment failure impacting safety of equipment, process or people. | Analytical - CHAKAN
AFPN24106414064 | 2024-10-21 | 2024-10-24 | Cust 1346 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Localhost not accessible on CPCB server. | We took access to the system and and found services were stopped. Now localhost is accessible. | We took access to the system and and found services were stopped. Now localhost is accessible. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24107384065 | 2024-10-23 | 2024-10-23 | Cust 793 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | CX2000 commissioning MOM report was pending | Mom has been done. And training given to customer. | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN24107484066 | 2024-10-23 | 2024-10-25 | Cust 1270 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Data upload issue. | This issue is from the hardware side.so we are closing this complaint. | This issue is from the hardware side.so we are closing this complaint. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24107504067 | 2024-10-23 | 2024-11-03 | Cust 1331 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Silica Analyser | Silica 8891(8891) | AWA Silica sequencer failed | visited the site and attended the issue of Sequencer. | visited the site and attned the issue of Sequencer. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24107514068 | 2024-10-23 | 2024-10-24 | Cust 1264 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | EnviroConnect software re-installation and configuration needs to be done. | We accessed the system and successfully completed the re-installation of the EnviroConnect software.and completed the configuration for CPCB and Local Servers and now data is live on both CPCB and Local servers. | We accessed the system and successfully completed the re-installation of the EnviroConnect software. and now data is live on both CPCB and Local servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24107634069 | 2024-10-23 | 2024-10-25 | Cust 1344 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB CONNECTIVITY NEEDS TO BE DONE. | We have completed the CPCB connectivity and now the ETP and Stack data is live on the CPCB server | We have completed the CPCB connectivity and now the ETP and Stack data is live on the CPCB server | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24107644070 | 2024-10-23 | 2024-11-04 | Cust 603 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | calibration job | problem solved | problem solved | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24108124071 | 2024-10-24 | 2024-11-04 | Cust 1347 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Amc visit | problem not solved | problem not solved | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24108194072 | 2024-10-24 | 2024-10-25 | Cust 603 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on CPCB server | We have accessed the system, and the data is now live on the CPCB server. | We have accessed the system, and the data is now live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24108444073 | 2024-10-25 | 2024-11-03 | Cust 1313 | PUNE | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Precommissioning visit | visited the site and necessary guidelines given to the customer for CX-2000 analyser | visited the site and necessary guidelines given to the customer | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24108644074 | 2024-10-25 | 2024-10-25 | Cust 731 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Local host server not accessible . | We have successfully accessed the system, and the local host server is now accessible. The data is currently live. | We have successfully accessed the system, and the local host server is now accessible. The data is currently live. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24108674075 | 2024-10-25 | 2024-10-28 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | cpcb data upload issue. | As discussed, we took access of the system and now the hardware side issue is resolved and data is live on the CPCB server and the same is verified by you. | As discussed, we took access of the system and now the hardware side issue is resolved and data is live on the CPCB server and the same is verified by you. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24108744076 | 2024-10-25 | 2024-10-28 | Cust 1348 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | MPCB connectivity needs to be done. | We have completed the MPCB connectivity Flow data live on the MPCB server. Please find the attached screenshot for your reference. | We have completed the MPCB connectivity Flow data live on the MPCB server. Please find the attached screenshot for your reference. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24108774077 | 2024-10-25 | 2024-10-25 | Cust 603 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Local host server is not accessible. | The local server is now accessible, and the data is live. | The local server is now accessible, and the data is live. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24110464078 | 2024-10-29 | 2024-10-30 | Cust 603 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | local host system and MPCB/CPCB URL not opening | The MPCB and CPCB Url is blocked in the coutomers network so once they allow the same the data starts uploading to the server hence as of now we are closing this complaint. | The MPCB and CPCB Url is blocked in the coutomers network so once they allow the same the data starts uploading to the server hence as of now we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24110774079 | 2024-10-29 | 2024-10-29 | Cust 731 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera streaming issue on CPCB server. | We have accessed the system, and the camera streaming is now live on the CPCB server. | We have accessed the system, and the camera streaming is now live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24112114080 | 2024-11-03 | 2024-11-05 | Cust 517 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Maintenance under EUTs | Visited the site and necessary maintenance carried out at site for Multipara meter | Visited the site and necessary maintenance carried out at site for Multipara meter | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24112284081 | 2024-11-04 | 2024-11-07 | Cust 803 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on MPCB Server. | We took access to the system and now data is live on MPCB Server. | We took access to the system and now data is live on MPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24112344082 | 2024-11-04 | 2024-11-04 | Cust 731 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Local host server is not accessible. | We accessed the system and found that the Apache Tomcat service was stopped on customer's local PC. After restarting it, the local server is now accessible. | We accessed the system and found that the Apache Tomcat service was stopped on customer's local PC. After restarting it, the local server is now accessible. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24112364083 | 2024-11-04 | 2024-11-04 | Cust 731 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera streaming issue. | camera streaming is now live on the CPCB server. | camera streaming is now live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24112524084 | 2024-11-04 | 2024-11-04 | Cust 1264 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on CPCB and Local servers. | Now the data is live on CPCB and Local servers. | Now the data is live on CPCB and Local servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24112744085 | 2024-11-05 | 2024-11-08 | Cust 603 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB and MpCB data upload issue. | now the issue has been resolved and the data is live on the CPCB and MPCB servers. | now the issue has been resolved and the data is live on the CPCB and MPCB servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24112764086 | 2024-11-05 | 2024-11-08 | Cust 1237 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | On ecolyser portal. It shows an NA message for the L6 unit. | Commercially not clear. | Commercially not clear. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24112844087 | 2024-11-05 | 2024-11-07 | Cust 731 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Localhost is not accessible. | We took access to the system and found Apache services were stopped. Now localhost is accessible. | We took access to the system and found Apache services were stopped. Now localhost is accessible. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24112884088 | 2024-11-05 | 2024-11-08 | Cust 731 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue. | We have not received the PO from the customer, so we are closing this complaint. | We have not received the PO from the customer, so we are closing this complaint. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24113874089 | 2024-11-07 | 2024-11-11 | Cust 1259 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | PH1 detector(1) | Detector zero | closed | completed the work | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24114004090 | 2024-11-07 | 2024-11-08 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB + MPCB data upload issue. | we took access to your system and changed configuration for Dust analyzer from RS485 to Scanner. Now data is live on CPCB + MPCB server. | we took access to your system and changed configuration for Dust analyzer from RS485 to Scanner. Now data is live on CPCB + MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24114194091 | 2024-11-07 | 2024-11-11 | Cust 1349 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | wrong cable supplied conductivity cable supplied instead of PH | material will supplied on urjant basis | completed the work | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24114564092 | 2024-11-08 | 2024-11-14 | Cust 811 | PUNE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | PH Sensor for Multiparameter(AWA-PH-SENS) | sensor not responding to buffer solution | offer send for new sensor | completed the work | NO | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24115594093 | 2024-11-11 | 2024-11-14 | Cust 1264 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on CPCB server. | We have accessed the system, and the data for ETP 1 is now live on the CPCB server. And However, the data for the Stack 2_130 TPH boiler is not communicating with the local PC. We have checked using MODSCAN, Hence this issue is from the hardware side . | We have accessed the system, and the data for ETP 1 is now live on the CPCB server . | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24115624094 | 2024-11-11 | 2024-11-14 | Cust 1233 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation + CPCB + Local + MPCB configuration needs to be done. | we took access to your system and now the hardware isde issue is resolved and data is live on CPCB + local server. | we took access to your system and now the hardware isde issue is resolved and data is live on CPCB + local server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24115804095 | 2024-11-11 | 2024-11-14 | Cust 603 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data mismatch issue. | We took access to the system and found MODSCAN data and server data is matching. | We took access to the system and found MODSCAN data and server data is matching. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24116164096 | 2024-11-12 | 2024-11-15 | Cust 1273 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | O.S. get corrupted need to reinstall the software. | We took access to the system and found data is not coming till MODSCAN. | We took access to the system and found data is not coming till MODSCAN. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24116404097 | 2024-11-12 | 2024-11-13 | Cust 640 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | The CPCB website is not opening in browser | We accessed the system and noticed that the CPCB URL was blocked on the customer's PC network. The issue has now been resolved, and the URL is accessible in the network. | We accessed the system and noticed that the CPCB URL was blocked on the customer's PC network. The issue has now been resolved, and the URL is accessible in the network. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24116414098 | 2024-11-12 | 2024-12-23 | Cust 757 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Chlorine sensor | CHLORINE SENSOR(1) | Chlorine analyser showing constant reading without any response in reading . | Visited site. Displays replaced. Chlorine sensor wrong electrolyte was filled. Corrected the same & Unitworking fine. | Electrolyte replaced. Working FIne | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFPN24116574099 | 2024-11-12 | 2024-11-15 | Cust 1251 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB + MPCB connectivity needs to be done. | today we took access of the black box and completed the MPCB configuration, and now data is live on the MPCB server. | today we took access of the black box and completed the MPCB configuration, and now data is live on the MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24117084100 | 2024-11-13 | 2024-11-13 | Cust 1289 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | DDU reading not matching with CPCB and local servers. | The issue has been resolved, and the DDU data is now matching with the CPCB and local servers, which has also been verified by the customer. | The issue has been resolved, and the DDU data is now matching with the CPCB and local servers, which has also been verified by the customer. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24117124101 | 2024-11-13 | 2024-11-13 | Cust 731 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on CPCB + Local + MPCB servers. | Now the data is live on CPCB + Local + MPCB servers. | Now the data is live on CPCB + Local + MPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24117864102 | 2024-11-14 | 2024-11-17 | Cust 992 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue. | We have not received remote support AMC PO details . so we are closing this complaint. | We have not received remote support AMC PO details . so we are closing this complaint. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24117914103 | 2024-11-14 | 2024-11-15 | Cust 1296 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Data not updating on Local Server. | We took access to the system and found Data is not coming till MODSCAN. | We took access to the system and found Data is not coming till MODSCAN. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24117964104 | 2024-11-14 | 2024-11-20 | Cust 1268 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Amc | Spare has been replaced calibration has been done device is healthy | calibration has done | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24118064105 | 2024-11-14 | 2024-11-18 | Cust 1290 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Maintenance under AMC visit | work done | problem solved | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24118074106 | 2024-11-14 | 2024-11-18 | Cust 1276 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC visit | problem solved | problem solved | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24118244107 | 2024-11-15 | 2024-11-18 | Cust 1268 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data mismatch issue. | we took access to the system and now the hardware side issue is resolved and data is live on CPCB server and matching with the actual reading and the same is verified by you. | we took access to the system and now the hardware side issue is resolved and data is live on CPCB server and matching with the actual reading and the same is verified by you. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24118694108 | 2024-11-16 | 2024-11-23 | Cust 1270 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | reset the analyser now working normal | reset the analyser and for ready for sample readings | completed the work | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24118754109 | 2024-11-17 | 2024-11-22 | Cust 1266 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | CPCB and MPCB data uploded issue. | for the Lokmangal Distillery unit, we have successfully completed the software installation and established connectivity with CPCB and MPCB. However, the analysers are currently having some issues hence the data is not being reflected on the server. | We have completed the software installation and established connectivity with CPCB and MPCB. However, the analysers are currently having some issues, so the data is not being reflected on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24118914110 | 2024-11-18 | 2024-11-23 | Cust 1277 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | ph values on display doesn't match with Mpcb | amc work completed | completed the work | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24118944111 | 2024-11-18 | 2024-11-18 | Cust 1277 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera streaming issue on CPCB server. | We took access to the system and we have checked, the camera is not playing in the browser or in the VLC media player ,hence this issue is from customer side .so we are closing this complaint. | We took access to the system and we have checked, the camera is not playing in the browser or in the VLC media player ,hence this issue is from customer side .so we are closing this complaint. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24119074112 | 2024-11-18 | 2024-11-19 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue. | we took access to the black box and made some changes from configuration end, now ETP Recycle Flow Meter | we took access to the black box and made some changes from configuration end, now ETP Recycle Flow Meter | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24119374113 | 2024-11-18 | 2024-11-21 | Cust 1035 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on CPCB server. | We took access to the system and observed that for Raw Spent Wash the data is not communicating with the local system (Checked on MODSCAN), For Concentrated Spent Wash, the comp port is not being detected...so issue is from the hardware side.so we are closing this complaint. | for Raw Spent Wash the data is not communicating with the local system (Checked on MODSCAN), For Concentrated Spent Wash, the comp port is not being detected... | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24119394114 | 2024-11-18 | 2024-11-21 | Cust 1285 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Re-installation of EnviroConnect software. | We have successfully completed the installation and configuration on software but data is not coming till MODSCAN | We have successfully completed the installation and configuration on software but data is not coming till MODSCAN | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24119504115 | 2024-11-18 | 2024-11-26 | Cust 1256 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Xenon power supply faulty | AMC visit done. Necessary maintenance activities carried out for Mutlipara. | AMC visit done. Necessary maintenance activities carried out for Mutlipara. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24119714116 | 2024-11-19 | 2024-11-21 | Cust 1217 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(Ph-Trans-Smartpro) | commissioning of the system | checked the installation of all flowmeter | completed the work | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24119844117 | 2024-11-19 | 2024-11-22 | Cust 1244 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Re-installation of EnviroConnect software. | We took access to the system and observed that for Raw Spent Wash the data is not communicating with the local system (Checked on MODSCAN), For Concentrated Spent Wash, the comp port is not being detected. So this issue is from the hardware side .So we are closing this complaint. | For Concentrated Spent Wash, the comp port is not being detected. So this issue is from the hardware side .So we are closing this complaint. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24119924118 | 2024-11-19 | 2024-11-22 | Cust 603 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation needs to be done. | We have completed software reinstallation + CPCB configuration now data is live on cpcb server for stack & etp data is not coming till system for now we are closing this complaint. | We have completed software reinstallation + CPCB configuration now data is live on cpcb server for stack &amp;amp; etp data is not coming till system for now we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24119934119 | 2024-11-19 | 2024-11-22 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software re-installation | As of now we are closing this complaint. We will reopen this once we receive access to the system. | As of now we are closing this complaint. We will reopen this once we receive access to the system. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24119984120 | 2024-11-19 | 2024-11-22 | Cust 1270 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software installation needs to be done. | for the Lokmangal Distillery unit, we have successfully completed the software installation and established connectivity with CPCB and MPCB. However, the analysers are currently having some issues hence the data is not being reflected on the server. | we completed the software installation and established connectivity with CPCB and MPCB. However, the analysers are currently having some issues hence the data is not being reflected on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24120154121 | 2024-11-19 | 2024-11-22 | Cust 1270 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Re-installation of EnviroConnect Software. | we have successfully re-installed the EnviroConnect software for the sugar unit of Lokmangal Agro Ind Ltd. The data is now live on both the CPCB and local servers for the Effluent Flow. | we have successfully re-installed the EnviroConnect software for the sugar unit of Lokmangal Agro Ind Ltd. The data is now live on both the CPCB and local servers for the Effluent Flow. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24120284122 | 2024-11-19 | 2024-11-21 | Cust 1350 | PUNE | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | Ph Tranmitter Smartpro | ORP transmitter(1) | ORP transmitters are not working. | Transmitter keys are not working and sensors are also exhausted. Suggested to procure new set up. | Closed | YES | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN24120284122 | 2024-11-19 | 2024-11-21 | Cust 1350 | PUNE | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | Ph Tranmitter Smartpro | ORP transmitter(1) | ORP transmitters are not working. | Transmitter keys are not working and sensors are also exhausted. Suggested to procure new set up. | Closed | YES | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN24120284122 | 2024-11-19 | 2024-11-21 | Cust 1350 | PUNE | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | Ph Tranmitter Smartpro | ORP transmitter(1) | Problem in keys operation | Transmitter keys are not working and sensors are also exhausted. Suggested to procure new set up. | Closed | YES | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN24120284122 | 2024-11-19 | 2024-11-21 | Cust 1350 | PUNE | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | Ph Tranmitter Smartpro | ORP transmitter(1) | Problem in keys operation | Transmitter keys are not working and sensors are also exhausted. Suggested to procure new set up. | Closed | YES | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN24120284122 | 2024-11-19 | 2024-11-21 | Cust 1350 | PUNE | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | Conductivity Transmitter Aquacon | Conductivity Transmitter Aquamon 2000(COND-Trans Aqu-2000) | conductivity transmitter | Transmitter keys are not working and sensors are also exhausted. Suggested to procure new set up. | Closed | YES | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN24120284122 | 2024-11-19 | 2024-11-21 | Cust 1350 | PUNE | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | Conductivity Transmitter Aquacon | Conductivity Transmitter Aquamon 2000(COND-Trans Aqu-2000) | PH transmitter | Transmitter keys are not working and sensors are also exhausted. Suggested to procure new set up. | Closed | YES | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN24120724123 | 2024-11-20 | 2024-11-23 | Cust 1343 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Display has been replaced, Calibration done | Display has been replaced with new one and calibration has been done device is working properly | commission has been done working properly customer satisfied | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24120894124 | 2024-11-21 | 2024-11-25 | Cust 1244 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on CPCB + Local + MPCB servers. | We have accessed the system, and the data is now live on the CPB and Local servers for The Malegaon Sahakari Sakhar Karkhana Ltd., Sugar Industry. ( Malegaon Sahakari Sakhar Stack ) | We have accessed the system, and the data is now live on the CPB and Local servers for The Malegaon Sahakari Sakhar Karkhana Ltd., Sugar Industry. ( Malegaon Sahakari Sakhar Stack ) | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24121154125 | 2024-11-21 | 2024-11-22 | Cust 1285 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Reinstallation and configuration of the Enviroconnect software. | We have successfully completed the installation and reconfiguration of software and now data is live on LOCAL+CPCB+KSPCB server. | We have successfully completed the installation and reconfiguration of software and now data is live on LOCAL+CPCB+KSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24121544126 | 2024-11-22 | 2024-11-23 | Cust 1239 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning of the new cx2000 Anaslyzer | Commissioning has been done working properly | commission has been done working properly customer satisfied | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24121564127 | 2024-11-22 | 2024-11-25 | Cust 1244 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Re -installation of EnviroConnect software | We have accessed the system and successfully re-installed the EnviroConnect software, as well as completed the configuration for the Malegaon Sahakari Sakhar Karkhana Ltd. Sugar Industry (Malegaon Sahakari Sakhar K) (ETP1). However, we have noticed that the data for ETP1 is not communicating with the local system (as checked on MODSCAN).This issue is from the hardware side..so we are closing this complaint. | We have accessed the system and successfully re-installed the EnviroConnect software, However, we have noticed that the data for ETP1 is not communicating with the local system | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24122344128 | 2024-11-24 | 2024-11-27 | Cust 598 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | ONLINE DATA is not upload on server, NEED TO CONFIGURE | online support to customer and resolve with data connctivity | completed the work | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24122634129 | 2024-11-25 | 2024-11-28 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation | We have re-installed the software with the latest version and completed the local, CPCB and MPCB server connectivity. | We have re-installed the software with the latest version and completed the local, CPCB and MPCB server connectivity. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24122734130 | 2024-11-25 | 2024-11-28 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | data is not uploading to MPCB server. | Issue has been resolved and data is now live on the MPCB server. | Issue has been resolved and data is now live on the MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24122744131 | 2024-11-25 | 2024-11-26 | Cust 1233 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data upload issue. | we took access to your system and now the MPCB server URL is accessible in your network and data is live on MPCB server. | we took access to your system and now the MPCB server URL is accessible in your network and data is live on MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24122754132 | 2024-11-25 | 2024-11-26 | Cust 1275 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data upload issue. | today we took access to your system and completed MPCB connectivity, now data is live on MPCB server and the same is verified by you. | today we took access to your system and completed MPCB connectivity, now data is live on MPCB server and the same is verified by you. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24122864133 | 2024-11-25 | 2024-11-25 | Cust 1277 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on CPCB + Local + MPCB servers. | Now the data is live on the CPCB + Local + MPCB servers. | Now the data is live on the CPCB + Local + MPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24122954134 | 2024-11-25 | 2024-11-26 | Cust 1002 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB+LOCAL | We took access to the system and now data is live on CPCB + LOCAL + MPCB server. | We took access to the system and now data is live on CPCB + LOCAL + MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24122964135 | 2024-11-25 | 2024-11-26 | Cust 1276 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | CPCB and Hardware data mismatching | We took access to the system and checked data on MODSCAN and same data is updating on CPCB and LOCAL server. | We took access to the system and checked data on MODSCAN and same data is updating on CPCB and LOCAL server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24122974136 | 2024-11-25 | 2024-11-26 | Cust 1285 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | PM data mismatch on Hardware and CPCB. | We took access to the system and found correct STACK data was updating on CPCB +LOCAL server. | We took access to the system and found correct STACK data was updating on CPCB +LOCAL server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24123334137 | 2024-11-26 | 2024-11-29 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on MPCB server. | We have not received the AMC PO details for remote support, so we are closing this complaint. | We have not received the AMC PO details for remote support, so we are closing this complaint. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24123414138 | 2024-11-26 | 2024-11-28 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB and MPCB data uploading issue | We have updated the EnviroConnect software DAS version to the latest version and now the data is live on the Local, CPCB and MPCB server. | We have updated the EnviroConnect software DAS version to the latest version and now the data is live on the Local, CPCB and MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24123494139 | 2024-11-26 | 2024-12-03 | Cust 1326 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | CPU Display card for Multiparameter(AWA-CPU1) | Display configuration. | Visited the site and attended the issue of Analyser Display | Visited the site and attended the issue of Analyser Display | NO | Paid | Core | N |  | Analytical - CHAKAN
AFPN24123544140 | 2024-11-26 | 2024-11-29 | Cust 1299 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Re-installation ad configuration of software. | We are closing this complaint for now. We will reopen it once the system is accessible. | We are closing this complaint for now. We will reopen it once the system is accessible. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24123564141 | 2024-11-26 | 2024-11-26 | Cust 1244 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on CPCB + Local servers. | We took access to the system and now the data is live on CPCB + Local servers for the Malegaon Sahakari Sakhar Karkhana Ltd. Sugar Industry (Malegaon Sahakari Sakhar K) (ETP1). | We took access to the system and now the data is live on CPCB + Local servers for the Malegaon Sahakari Sakhar Karkhana Ltd. Sugar Industry (Malegaon Sahakari Sakhar K) (ETP1). | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24123584142 | 2024-11-26 | 2024-11-29 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation needs to be done. | yesterday we took access to your system and completed software reinstallation + CPCB configuration now data is live on CPCB + MPCB server. | yesterday we took access to your system and completed software reinstallation + CPCB configuration now data is live on CPCB + MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24123604143 | 2024-11-26 | 2024-11-29 | Cust 1035 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to local and CPCB server. | As of now we are closing this complaint. We will reopen this once the issue is resolved from the hardware side. | As of now we are closing this complaint. We will reopen this once the issue is resolved from the hardware side. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24123614144 | 2024-11-26 | 2024-11-26 | Cust 1266 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on MPCB server. | Now the data is live on the MPCB server. | Now the data is live on the MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24123884145 | 2024-11-27 | 2024-11-29 | Cust 1351 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB and KSPCB data uploaded issue. | The issue has been resolved and now the data is live on CPCB + KSPCB servers. | The issue has been resolved and now the data is live on CPCB + KSPCB servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24123894146 | 2024-11-27 | 2024-12-02 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | software re-installation | We have updated the EnviroConnect software DAS version to the latest version i.e. 4.6.0.0. Issue has been resolved and data is now live on the Local, CPCB and MPCB servers. | We have updated the EnviroConnect software DAS version to the latest version i.e. 4.6.0.0. Issue has been resolved and data is now live on the Local, CPCB and MPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24123954147 | 2024-11-27 | 2024-11-29 | Cust 1314 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Blackbox commissioning. | Yesterday we have completed Black configuration + CPCB and MPCB connectivity, so now the data is live on both sites | Yesterday we have completed Black configuration + CPCB and MPCB connectivity, so now the data is live on both sites | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24125014148 | 2024-11-28 | 2024-11-29 | Cust 1233 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue. | we took access to your system and now ETP_1_Outlet Flow data is coming till the local system and data is uploading on CPCB server | we took access to your system and now ETP_1_Outlet Flow data is coming till the local system and data is uploading on CPCB server | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24125044149 | 2024-11-28 | 2024-11-29 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB + MPCB data upload issue. | we took access of the black box and now the internet side issue is resolved and data is live on CPCB + MPCB server. | we took access of the black box and now the internet side issue is resolved and data is live on CPCB + MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24125144150 | 2024-11-28 | 2024-11-29 | Cust 991 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Reinstallation and configuration of software. | We took access to the system and completed the installation and configuration of software and now data is live on CPCB+LOCAL+MPCB Server. | We took access to the system and completed the installation and configuration of software and now data is live on CPCB+LOCAL+MPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24125384151 | 2024-11-29 | 2024-11-29 | Cust 1321 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue. | we took access of the system and made some modifications from configuration end for STACK1 (DCEM2100A), and now the issue has been resolved and data is live on CPCB + MPCB server | we took access of the system and made some modifications from configuration end for STACK1 (DCEM2100A), and now the issue has been resolved and data is live on CPCB + MPCB server | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24125464152 | 2024-11-29 | 2024-12-02 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Re-installation of EnviroConnect software. | We took access to the system and successfully completed the software re-installation and completed the CPCB + Local configuration. And now the data is live on CPCB + Local servers. | We took access to the system and successfully completed the software re-installation and completed the CPCB + Local configuration. And now the data is live on CPCB + Local servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24125524153 | 2024-11-29 | 2024-12-02 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue | We have changed configuration from RS485 to scanner as per requirement for ETP 1 and data is now live on the CPCB server. | We have changed configuration from RS485 to scanner as per requirement for ETP 1 and data is now live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24125734154 | 2024-11-29 | 2024-12-11 | Cust 1289 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | installation of MPA | job done | commissioning job done only connectivity part remaining due to mpcb confirmation. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFPN24125784155 | 2024-11-29 | 2024-12-02 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data uploading issue | As of now we are closing this complaint. We will reopen this once the internet issue is resolved from the customer end. | As of now we are closing this complaint. We will reopen this once the internet issue is resolved from the customer end. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24125804156 | 2024-11-29 | 2024-11-29 | Cust 1289 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on MPCB server. | We took access to the system and now data is live on the MPCB server. | We took access to the system and now data is live on the MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24126324157 | 2024-11-30 | 2024-12-11 | Cust 1295 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | 12 V POWER SUPPLY CARD for Multiparameter(AWA-PS1) | power supply card needs to be checked and others | offer submitted for power supply card and ph input card | completed the work | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24127864158 | 2024-12-02 | 2024-12-02 | Cust 1280 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on MPCB server. | Now the data is live on the MPCB server. | Now the data is live on the MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24128054159 | 2024-12-02 | 2024-12-04 | Cust 1284 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Stack 1_ 22 TPH Incineration Boiler device data is not updating on CPCB + MPCB server. | We took access to the system and observed data is not coming till MODSCAN . There is hardware side issue we are closing this complaint. We will re-open once the issue is resolved from Hardware end. | We took access to the system and observed data is not coming till MODSCAN . There is hardware side issue we are closing this complaint. We will re-open once the issue is resolved from Hardware end. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24128184160 | 2024-12-02 | 2024-12-02 | Cust 1244 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | local host server not accessible. | We took access to the system and now the local host server is accessible and data is live on the Local server. | We took access to the system and now the local host server is accessible and data is live on the Local server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24128194161 | 2024-12-02 | 2024-12-02 | Cust 1244 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | To change the permissible range for Flow on the local host server. | We accessed the system and changed the permissible flow range on the local host server. The new range is now set to 100. | We accessed the system and changed the permissible flow range on the local host server. The new range is now set to 100. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24128384162 | 2024-12-02 | 2024-12-05 | Cust 1285 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data mismatch on CPCB+LOCAL | We took access and change the scanner address and now Local and CPCB data is matching. | We took access and change the scanner address and now Local and CPCB data is matching. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24128704163 | 2024-12-03 | 2024-12-06 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB+MPCB server for - Vitthal Sugar Manufacturing Ltd. | We took access to the system and checked data is not coming till MODSCAN. We are closing this complaint as of now we will re-open once issue is resolved from hardware end. | We took access to the system and checked data is not coming till MODSCAN. We are closing this complaint as of now we will re-open once issue is resolved from hardware end. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24128714164 | 2024-12-03 | 2024-12-05 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on CPCB server for - Ajinkyatara Sahakari Sakhar Karkhana-Distillery | We took access to the system and Completed configuration now cameras are playing on CPCB server. | We took access to the system and Completed configuration now cameras are playing on CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24128724165 | 2024-12-03 | 2024-12-06 | Cust 617 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on MPCB server. | We took access to the system and added output configuration after that data was live on MPCB server. | We took access to the system and added output configuration after that data was live on MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24128824166 | 2024-12-03 | 2024-12-06 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Reinstallation and configuration software for Shree Halasiddhanath SSK Ltd | We took access to the system and completed the reinstallation and configuration of software and data is now on LOCAL+CPCB+KSPCB server. | We took access to the system and completed the reinstallation and configuration of software and data is now on LOCAL+CPCB+KSPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24128854167 | 2024-12-03 | 2024-12-06 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | data is not uploading to local, cpcb and mpcb | We have updated the Enviroconnect DAS version and data is now live on the Local, CPCB and MPCB servers. | We have updated the Enviroconnect DAS version and data is now live on the Local, CPCB and MPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24129164168 | 2024-12-04 | 2024-12-10 | Cust 1279 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | software re-installation | We have updated enviroconnect DAS version to 4.6 and completed Local, CPCB and MPCB connectivity. Now data is live on the same servers. | We have updated enviroconnect DAS version to 4.6 and completed Local, CPCB and MPCB connectivity. Now data is live on the same servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24129184169 | 2024-12-04 | 2024-12-05 | Cust 1244 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Re-installation of EnviroConnect software. | We accessed your system and successfully re-installed the EnviroConnect software. And completed CPCB + LOCAL + MPCB configuration. Now the data for Raw Spent Wash is now live on the CPCB server. However, we are facing an issue with the COM Port for Concentrated Spent Wash, as it is not being detected. So this issue is from the hardware side .so we are closing this complaint. | We accessed your system and successfully re-installed the EnviroConnect software. And completed CPCB + LOCAL + MPCB configuration. Now the data for Raw Spent Wash is now live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24129454170 | 2024-12-04 | 2024-12-05 | Cust 1305 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on CPCB + Local + MPCB servers | We accessed the system and noticed that the data is not communicating with the local system (as checked on MODSCAN). So this issue is from the hardware side .So we are closing this complaint. | We accessed the system and noticed that the data is not communicating with the local system (as checked on MODSCAN). So this issue is from the hardware side .So we are closing this complaint. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24129474171 | 2024-12-04 | 2024-12-06 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | software re-installation | We have re-installed the software with an updated version and completed Local, CPCB and MPCB connectivity. Data is live on the same servers. | We have re-installed the software with an updated version and completed Local, CPCB and MPCB connectivity. Data is live on the same servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24129504172 | 2024-12-04 | 2024-12-06 | Cust 1275 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera's not playing on CPCB server. | We took access and check streaming window was closed and there was upgraded version of firefox. Cameras are live now and told customer to downgrade version to 128.0.3. | We took access and check streaming window was closed and there was upgraded version of firefox. Cameras are live now and told customer to downgrade version to 128.0.3. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24129524173 | 2024-12-04 | 2024-12-06 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | cpcb data upload issue. | we took access to your system and now the hardware side issue is resolved for Stack 1_17.5 TPH Boiler I & II and data is live on CPCB server. | we took access to your system and now the hardware side issue is resolved for Stack 1_17.5 TPH Boiler I &amp;amp; II and data is live on CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24129684174 | 2024-12-05 | 2024-12-07 | Cust 1272 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | ETP Data is not live on the CPCB server. | We have shifted MPA connectivity from scanner to RS485 card and data is uploading on CPCB Server. | We have shifted MPA connectivity from scanner to RS485 card and data is uploading on CPCB Server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24129774175 | 2024-12-05 | 2024-12-06 | Cust 1296 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on CPCB server. | We are closing this complaint as of now we will re-open once we get access to the system. | We are closing this complaint as of now we will re-open once we get access to the system. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24129854176 | 2024-12-05 | 2024-12-05 | Cust 1255 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data upload issue. | Now the data is live on the CPCB + Local + MPCB servers. | Now the data is live on the CPCB + Local + MPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24130224177 | 2024-12-06 | 2024-12-12 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Re-installation of EnviroConnect Software. | We took access to the system and successfully completed the software installation and configuration and the data is live on CPCB + LOCAL + MPCB servers. | We took access to the system and successfully completed the software installation and configuration and the data is live on CPCB + LOCAL + MPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24130444178 | 2024-12-06 | 2024-12-09 | Cust 1352 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software installation and CPCB + MPCB connectivity needs to be done. | For now we are closing this compliant. | For now we are closing this compliant. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24130454179 | 2024-12-06 | 2024-12-12 | Cust 1289 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Local Configuration needs to be done. | WE TOOK ACCESS TO THE SYSTEM AND COMPLETED THE LOCAL SERVER CONFIGUARTION AND THE DATA IS LIVE ON LOCAL SERVER. | WE TOOK ACCESS TO THE SYSTEM AND COMPLETED THE LOCAL SERVER CONFIGUARTION AND THE DATA IS LIVE ON LOCAL SERVER. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24130654180 | 2024-12-07 | 2024-12-11 | Cust 1313 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | for installation of device | CX-2000 commissioning done. | CX-2000 commissioning done. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24130674181 | 2024-12-07 | 2024-12-10 | Cust 1259 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB DATA UPLOAD FOR SKODA AUTO Volkswagen India Private Limited SKODA AUTO Volkswagen India Private Limited (formerly known as Volkswagen India Private Limited | As of now we are closing this complaint and will reopen it once the hardware side issue is resolved. | As of now we are closing this complaint and will reopen it once the hardware side issue is resolved. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24131034182 | 2024-12-08 | 2024-12-11 | Cust 1308 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | CPU Display card for Multiparameter(AWA-CPU1) | MAINTENANCE AND DATA CONNECTIVITY | MAINTENANCE AND DATA CONNECTIVITY | completed the work | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24131594183 | 2024-12-09 | 2024-12-12 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on LOCAL+CPCB+MPCB server.- Vitthal Corporation Ltd. Sugar | We took access to the system and completed the installation and configuration and now data is live on LOCL+CPCB+MPCB server. | We took access to the system and completed the installation and configuration and now data is live on LOCL+CPCB+MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24131894184 | 2024-12-09 | 2024-12-10 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Flow value not uploading on CPCB server. | The data is also live on the CPCB server, and the flow values have been successfully uploaded. | The data is also live on the CPCB server, and the flow values have been successfully uploaded. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24131924185 | 2024-12-09 | 2024-12-12 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on MPCB Server | We took access to the system and completed the MPCB configuration and now data is live on MPCB server. | We took access to the system and completed the MPCB configuration and now data is live on MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24132174186 | 2024-12-10 | 2024-12-12 | Cust 1306 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | localhost page is not accessible. | We took access and found services were stopped due to which localhost was not accessible. | We took access and found services were stopped due to which localhost was not accessible. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24132184187 | 2024-12-10 | 2024-12-13 | Cust 1035 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Data mismatch issue | As discussed on the call, we accessed a Windows-based system, but the COM port was not detected, so we were unable to check the data on Modscan. For now, we are closing this complaint. We will reopen it once the COM port issue is resolved. | As discussed on the call, we accessed a Windows-based system, but the COM port was not detected, so we were unable to check the data on Modscan. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24132194188 | 2024-12-10 | 2024-12-16 | Cust 603 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | CPU with LCD of Multiparameter Display(1) | Repaired display mounted and analyzer AMC has to done | AMC has done display working properly | Needs to repair the display | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24132224189 | 2024-12-10 | 2024-12-10 | Cust 1035 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | localhost not opening | Localhost is now accessible after restarting the apache tomcat service. | Localhost is now accessible after restarting the apache tomcat service. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24132374190 | 2024-12-10 | 2024-12-13 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB + MPCB data upload issue., | we took access of the black box and made some changes from configuration end, and also the internet side issue is resolved from your end. Now data is live on CPCB + MPCB server and matching with actual readings same is verified by you also. | Now data is live on CPCB + MPCB server and matching with actual readings same is verified by you also. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24132384191 | 2024-12-10 | 2024-12-13 | Cust 1256 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | CPCB data mismatch issue. | we took access to the system and changed device ID for Stack_3 Incineration Boiler 45 TPH, now data is live on CPCB + MPCB server and matching with the actual readings. | we took access to the system and changed device ID for Stack_3 Incineration Boiler 45 TPH, now data is live on CPCB + MPCB server and matching with the actual readings. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24132494192 | 2024-12-11 | 2024-12-17 | Cust 1232 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | customer wants digital connectivity . | problem solved | problem solved | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24132504193 | 2024-12-11 | 2024-12-17 | Cust 731 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | sample pump not working | pump need to be replace with new pair | problem solved | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFPN24132514194 | 2024-12-11 | 2024-12-16 | Cust 809 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | MPCB data break issue. | we took access to the black box and updated the new patch, now data is continuously on the MPCB server | we took access to the black box and updated the new patch, now data is continuously on the MPCB server | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24132534195 | 2024-12-11 | 2024-12-12 | Cust 1289 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software installation for ETP. | We have successfully accessed the system and completed the installation of the EnviroConnect software for the ETP, including the Local server configuration and the data is live on local server | We have successfully accessed the system and completed the installation of the EnviroConnect software for the ETP, including the Local server configuration and the data is live on local server | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24132574196 | 2024-12-11 | 2024-12-14 | Cust 1293 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue. | We have not received the previous payment or the advance payment for this year. Therefore, we are closing this complaint. | We have not received the previous payment or the advance payment for this year. Therefore, we are closing this complaint. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24132584197 | 2024-12-11 | 2024-12-12 | Cust 1035 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue | We have accessed the system and observed that the ETP Flow data is not coming till the local system (as verified on Modscan) So this issue is from the hardware side .So we are closing this complaint. | We have accessed the system and observed that the ETP Flow data is not coming till the local system (as verified on Modscan) | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24132714198 | 2024-12-11 | 2024-12-13 | Cust 1352 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Installation and configuration of software. | We took access to the system and completed the installation and configuration of software and now data is live on LOCAL+CPCB+MPCB server. | We took access to the system and completed the installation and configuration of software and now data is live on LOCAL+CPCB+MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24132754199 | 2024-12-11 | 2024-12-17 | Cust 1353 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | inbuilt configuration problem | job done | problem solved | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24132774200 | 2024-12-11 | 2024-12-13 | Cust 1233 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | mpcb connectivity needs to be done. | we took access to your system and completed MPCB connectivity for Flow meter, now data is live on MPCB server. | we took access to your system and completed MPCB connectivity for Flow meter, now data is live on MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24133134201 | 2024-12-12 | 2024-12-13 | Cust 1291 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Localhost not accessible and data is not live on server. | We are closing this complaint and we will re-open once the internet issue is resolved from the customer side. | We are closing this complaint and we will re-open once the internet issue is resolved from the customer side. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24133144202 | 2024-12-12 | 2024-12-13 | Cust 1306 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Localhost not accessible. | We took access to the system and found services were stopped. Now localhost is accessible. | We took access to the system and found services were stopped. Now localhost is accessible. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24133164203 | 2024-12-12 | 2024-12-13 | Cust 1258 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Localhost page opening and camera not playing on CPCB Server. | We are closing this complaint and we will re-open once the internet issue is resolved from the customer side. | We are closing this complaint and we will re-open once the internet issue is resolved from the customer side. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24133224204 | 2024-12-12 | 2024-12-13 | Cust 1275 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on CPCB server and stack data is constant. | We took access to the system and found cameras were live on CPCB server and Stack data was also updating and was not constant. | We took access to the system and found cameras were live on CPCB server and Stack data was also updating and was not constant. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24133264205 | 2024-12-12 | 2024-12-13 | Cust 1232 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | FMBB configuration + CPCB + MPCB connectivity needs to be done. | we took access of the Black Box and completed Black Box configuration + CPCB + MPCB connectivity (From DATCon to FMBB), now data is live on both the servers. | we took access of the Black Box and completed Black Box configuration + CPCB + MPCB connectivity (From DATCon to FMBB), now data is live on both the servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24133274206 | 2024-12-12 | 2024-12-13 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | DATCON | DATCON(DATCON) | CPCB + MPCB data upload issue. | we have made some changes from configuration end (DATCon & Black Box), and now data is live on CPCB + MPCB server, the same is verified by you also. | we have made some changes from configuration end (DATCon &amp;amp; Black Box), and now data is live on CPCB + MPCB server, the same is verified by you also. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24133294207 | 2024-12-12 | 2024-12-13 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | DATCON | DATCON(DATCON) | cpcb + mpcb data upload issue. | we have made some changes from configuration end (DATCon & Black Box), and now data is live on CPCB + MPCB server, the same is verified by you also. | we have made some changes from configuration end (DATCon &amp;amp; Black Box), and now data is live on CPCB + MPCB server, the same is verified by you also. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24133594208 | 2024-12-13 | 2024-12-14 | Cust 1280 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue for stack. | We have accessed your system, and the data for Stack 1_(56TPH) Common Boiler I and II is now live on the CPCB + Local servers. | We have accessed the system, and the data for Stack 1_(56TPH) Common Boiler I and II is now live on the CPCB + Local servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24133694209 | 2024-12-13 | 2024-12-14 | Cust 1305 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Local host server is not accessible | We have successfully accessed your system, and the local host server is now accessible | We have successfully accessed your system, and the local host server is now accessible | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24133704210 | 2024-12-13 | 2024-12-14 | Cust 1305 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on CPCB + Local servers. | We took access of the system and the data for ETP 1 and ETP FLOW is live on both the CPCB and local servers. However, we have noticed that data is not being communicated with the local system for Stack 1 and Stack 2. We have checked the data using ModScan. So this issue is from the hardware side.so we are closing this complaint. | We took access of the system and the data for ETP 1 and ETP FLOW is live on both the CPCB and local server So this issue is from the hardware side.so we are closing this complaint. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24133724211 | 2024-12-13 | 2024-12-17 | Cust 1354 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | pH Transmitter Aquamon | Aquamon 2000 plus(AQ2000 plus) | Auto ph cleaning system commissioning | Auto ph cleaning system commissioning | completed the work | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24133754212 | 2024-12-13 | 2024-12-14 | Cust 1280 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on MPCB server for ETP_Outlet. | Additionally, as discussed during the call with the customer, the issue with uploading the ETP_Outlet data to the MPCB server has now been resolved. The data is now live on the MPCB server. | Additionally, as discussed during the call with the customer, the issue with uploading the ETP_Outlet data to the MPCB server has now been resolved. The data is now live on the MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24133784213 | 2024-12-13 | 2024-12-28 | Cust 1355 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB connectivity needs to be done. | CPCB connectivity is complete and data is live on the server. | CPCB connectivity is complete and data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24134004214 | 2024-12-14 | 2024-12-14 | Cust 1327 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | CPCB and MPCB data uploaded issue. | Now the ETP data for the Sugar unit is live on both the CPCB and MPCB servers. | Now the ETP data for the Sugar unit is live on both the CPCB and MPCB servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24134584215 | 2024-12-16 | 2024-12-17 | Cust 1302 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Flow data issue on MPCB server. | Now the data is live on the MPCB server. | Now the data is live on the MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24134674216 | 2024-12-16 | 2024-12-19 | Cust 1321 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on MPCB server. | We took access to the system and found apache services were stopped. Now data is live on CPCB+MPCB server after restarting services. | We took access to the system and found apache services were stopped. Now data is live on CPCB+MPCB server after restarting services. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24134684217 | 2024-12-16 | 2024-12-19 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Data is not uploading to CPCB and MPCB | data is reflected to the CPCB server continuously after replacing the converter. | data is reflected to the CPCB server continuously after replacing the converter. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24134704218 | 2024-12-16 | 2024-12-19 | Cust 528 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Customer requested for Nov moth report. | We have shared the Nov month data with customer. | We have shared the Nov month data with customer. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24134754219 | 2024-12-16 | 2024-12-19 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on MPCB Server - Shripati Sugar & Power Ltd. | We took access to the system and data is live on MPCB server. We restarted the services and for stack data is breaking checked on modscan | We took access to the system and data is live on MPCB server. We restarted the services and for stack data is breaking checked on modscan | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24135134220 | 2024-12-17 | 2024-12-18 | Cust 1233 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on CPCB + Local + MPCB Servers. | We accessed the system and the data for ETP Outlet & ETP Outlet_Flow is live on CPCB + Local + MPCB servers. However, we have noticed that data is not being communicated with the local system for STACK1_Boiler 110TPH. We have checked the data using ModScan.So this issue is from the hardware side . | We accessed the system and the data for ETP Outlet &amp;amp;amp;amp;amp; ETP Outlet_Flow is live on CPCB + Local + MPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24135324221 | 2024-12-17 | 2024-12-20 | Cust 1263 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | data not live on LOCAL+CPCB+MPCB | We took access to the system and completed the Enviroconnect software installation and configuration and now data is live on LOCAL+CPCB+MPCB server. | We took access to the system and completed the Enviroconnect software installation and configuration and now data is live on LOCAL+CPCB+MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24135474222 | 2024-12-17 | 2024-12-31 | Cust 1289 | PUNE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | no problem | no problem | completed the work | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24135664223 | 2024-12-18 | 2024-12-20 | Cust 1318 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Samarth SSK Unit 2 : MPCB data uploaded issue. | Now the data is live on the MPCB server. | Now the data is live on the MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24135674224 | 2024-12-18 | 2024-12-20 | Cust 1315 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software Re-installation . | As discussed on the call, the customer's PC is not working. Following our conversation, we have re-installed the EnviroConnect software on the new PC and completed the CPCB, LOCAL, and MPCB configurations. However, we have noticed that data is not communicating with the local system for ETP Outlet , Stack 1_Bagasse Fired Boiler & ETP _FLow. We checked the data using MODSCAN, So please check this issue from the hardware side | . Following our conversation, we have re-installed the EnviroConnect software on the new PC and completed the CPCB, LOCAL, and MPCB configurations. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24135684225 | 2024-12-18 | 2024-12-19 | Cust 809 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB + MPCB data upload issue. | we took access to your system and now the issue has been resolved and data is live on CPCB + MPCB server for Stack 1_Solid Fuel Fired Boiler. | we took access to your system and now the issue has been resolved and data is live on CPCB + MPCB server for Stack 1_Solid Fuel Fired Boiler. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24135694226 | 2024-12-18 | 2024-12-21 | Cust 1291 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data upload issue | Anydesk version | AnyDesk version issue | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24135704227 | 2024-12-18 | 2024-12-20 | Cust 1261 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB+MPCB data upload issue. | We took access to the system and found services were stopped and now data is live on LOCAL+CPCB+MPCB server. | We took access to the system and found services were stopped and now data is live on LOCAL+CPCB+MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24135714228 | 2024-12-18 | 2025-01-12 | Cust 1296 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on CPCB server for Ambedkar Sugar SSK distillery unit | Anydesk is access now there is hardware side issue from customer end once issue resolved we will reopen this complaint. | Anydesk is access now there is hardware side issue from customer end once issue resolved we will reopen this complaint. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24135724229 | 2024-12-18 | 2024-12-20 | Cust 1285 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | For stack data on CPCB+MPCB getting high values. | We took access to the system and checked data on modscan and same data is updating on CPCB server. Informed customer same. | We took access to the system and checked data on modscan and same data is updating on CPCB server. Informed customer same. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24135784230 | 2024-12-18 | 2024-12-20 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data upload issue | Data is live on the MPCB server. | Data is live on the MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24135834231 | 2024-12-18 | 2024-12-19 | Cust 731 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on CPCB + Local + MPCB servers for ETP Outlet Flow. | As discussed during the call , we have accessed the system, and the data is now live on the CPCB, Local, and MPCB servers for the ETP Outlet Flow. | As discussed during the call , we have accessed the system, and the data is now live on the CPCB, Local, and MPCB servers for the ETP Outlet Flow. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24135854232 | 2024-12-18 | 2024-12-20 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB and MPCB data upload issue | We have updated EnviroconnectDAS version and Data is now live on the CPCB and MPCB server and the same is verified. | We have updated EnviroconnectDAS version and Data is now live on the CPCB and MPCB server and the same is verified. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24135964233 | 2024-12-18 | 2024-12-26 | Cust 1356 | PUNE | Krohne - MIDC | CLOSED | PRODUCT FAILURE | Vortex Meter OPTISWIRL Primary | VFS 4000 - DN 100 / 4 inch OPTISWIRL Primary 1R(VFS 4000 DN100 1R) | Showing zero value | Necessary settings done and flow meter found working satisfactory. | Necessary settings done and flow meter found working satisfacctory. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24136664234 | 2024-12-20 | 2024-12-20 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Installation and configuration of software - Sonhira distillery | We took access to the system and completed the configuration and installation of software. | We took access to the system and completed the configuration and installation of software. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24137214235 | 2024-12-20 | 2024-12-23 | Cust 1278 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | ETP online monitoring System ISSUE | we took access of the system and it was observed that there was some loose connection with the USB converter and after reconnecting it properly the device communication was started. | we took access of the system and it was observed that there was some loose connection with the USB converter and after reconnecting it properly the device communication was started. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24138284236 | 2024-12-23 | 2024-12-24 | Cust 603 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB + KSPCB connectivity needs to be done. | as of now we are closing this complaint and will reopen it once we get internet details fro customers. | as of now we are closing this complaint and will reopen it once we get internet details fro customers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24138354237 | 2024-12-23 | 2024-12-26 | Cust 1275 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue. | Checked and data is not coming till modscan.We have informed customer. | Checked and data is not coming till modscan.We have informed customer. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24138524238 | 2024-12-23 | 2024-12-24 | Cust 1285 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Installation and configuration of software. | We took access to the system and completed the installation and configuration of software now data is live. | We took access to the system and completed the installation and configuration of software now data is live. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24138534239 | 2024-12-23 | 2024-12-24 | Cust 1357 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on MPCB server. | We are closing this complaint we will reopen once issue resolved from hardware end. | We are closing this complaint we will reopen once issue resolved from hardware end. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24138634240 | 2024-12-23 | 2024-12-24 | Cust 1297 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | reinstallation and configuration of software. | We took access to the system and completed the installation and configuration of software now data is live on CPCB server. | We took access to the system and completed the installation and configuration of software now data is live on CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24138644241 | 2024-12-23 | 2024-12-24 | Cust 1306 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB server. | We are closing this complaint we will reopen once issue resolved from hardware end. | We are closing this complaint we will reopen once issue resolved from hardware end. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24138664242 | 2024-12-23 | 2024-12-26 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation | We have re-installed the software with an upgraded version and data is now live on the Local, CPCB and MPCB server. | We have re-installed the software with an upgraded version and data is now live on the Local, CPCB and MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24138704243 | 2024-12-23 | 2024-12-26 | Cust 1358 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data uploading issue | As of now we are closing this complaint. We will reopen this once we get access to the system. | As of now we are closing this complaint. We will reopen this once we get access to the system. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24138954244 | 2024-12-24 | 2024-12-24 | Cust 991 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB+MPCB server. | We are closing this complaint we will reopen once issue resolved from hardware end. | We are closing this complaint we will reopen once issue resolved from hardware end. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24138974245 | 2024-12-24 | 2024-12-26 | Cust 1259 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | 24 VDC SMPS (1) | Commissioning of rs485 card and SMPS | Supply checked working properly | commission has been done working properly customer satisfied | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24138974245 | 2024-12-24 | 2024-12-26 | Cust 1259 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | RS485 card-FM of Multiaparameter (1) | Commissioning of rs485 card and SMPS | Data is online after commissioning | commission has been done working properly customer satisfied | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24139164246 | 2024-12-24 | 2024-12-26 | Cust 1035 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software re-installation | As of now we are closing this complaint and will reopen this once we get access to the system. | As of now we are closing this complaint and will reopen this once we get access to the system. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24139674247 | 2024-12-26 | 2024-12-28 | Cust 1337 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB server not working. | We took access to the system and made changes in output driver configuration and now data is live on MPCB server | We took access to the system and made changes in output driver configuration and now data is live on MPCB server | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24139694248 | 2024-12-26 | 2024-12-27 | Cust 1313 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Local Configuration needs to be done, | We have completed the software installation and now the data is live on the Local server. | We have completed the software installation and now the data is live on the Local server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24139754249 | 2024-12-26 | 2024-12-28 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | change configuration from scanner to direct rs485 for 110 tph | We have changed configuration from scanner to RS485 as per requirement and data is now live on the CPCB server. | We have changed configuration from scanner to RS485 as per requirement and data is now live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24139864250 | 2024-12-26 | 2024-12-28 | Cust 1297 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on MPCB server | We took access to the system and made changes in output driver configuration and now data is live on MPCB server | We took access to the system and made changes in output driver configuration and now data is live on MPCB server | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24139884251 | 2024-12-26 | 2025-01-27 | Cust 1355 | PUNE | Steam System Division | CLOSED | COMMISSIONING | ALL PRODUCT | FORBES MARSHALL STEAM SYSTEM DIVISION PRODUCTS(ALL PRODUCTS) | online connectivity of CPCB and KSPCB , commissioning of MP data not transmitting | on site for almost 2 month , MP commissioning completed along with all the CELRS commissioning activity & Instrumentation work | on site for almost 2 month , completed all the CELRS commissioning activity & Instrumentation work | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24139954252 | 2024-12-26 | 2024-12-28 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on MPCB server for Shripati Sugar & Power Ltd. | We took access to the system and found services were stopped we have started and now data is live on MPCB server. | We took access to the system and found services were stopped we have started and now data is live on MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24140134253 | 2024-12-26 | 2024-12-31 | Cust 1359 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | Showing wrong readings. | visit done. calibration activities carried out for pH and conductivity | visit done. calibration activities carried out for pH and conductivity | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24140234254 | 2024-12-27 | 2025-01-09 | Cust 1035 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | suyog enterprieses sant tukaram site issue for online connectivity | suyog enterprieses sant tukaram site issue for online connectivity | completed the work | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24140764255 | 2024-12-28 | 2024-12-28 | Cust 992 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data uploaded issue. | Now the data is live on the MPCB server | Now the data is live on the MPCB server | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24140774256 | 2024-12-28 | 2024-12-28 | Cust 1289 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data uploaded issue. | Completed the MPCB connectivity and data is live. | Completed the MPCB connectivity and data is live. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24140834257 | 2024-12-28 | 2024-12-30 | Cust 1264 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on CPCB + Local servers. | We have accessed the system, and the data for ETP 1 and Stack 2_130 TPH boiler is now live on both CPCB and Local servers. Regarding Stack 1_Common boiler 60- 90 TPH ,there is hardware issue. | We have accessed the system, and the data for ETP 1 and Stack 2_130 TPH boiler is now live on both CPCB and Local servers. Regarding Stack 1_Common boiler 60- 90 TPH ,there is hardware issue. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24140954258 | 2024-12-28 | 2024-12-30 | Cust 603 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Flow value showing zero on CPCB server. | The value for Spent wash to inceneration boiler (Flow) is now being uploaded to the CPCB server | The value for Spent wash to inceneration boiler (Flow) is now being uploaded to the CPCB server | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24142064259 | 2024-12-30 | 2025-01-02 | Cust 1347 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | EnviroConnect software re-installation and configuration needs to be done. | We have not been able to access the system due to an internet issue on the customer's end. As a result, we are closing this complaint for now. We will reopen it once the internet issue is resolved and we are able to access the system. | We have not been able to access the system due to an internet issue on the customer's end. As a result, we are closing this complaint for now. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24142084260 | 2024-12-30 | 2025-01-02 | Cust 1035 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB and MPCB data uploading issue | We have upgraded the Enviroconnect DAS version i.e. 4.6 and the data from ETP Outlet and ETP Outlet_Flow is live on the CPCB. Currently, we are not getting counts on the modscan from STACK 1_85TPH Boiler. hence will reopen this complaint once the issue is resolved from hardware side | We have upgraded the Enviroconnect DAS version i.e. 4.6 and the data from ETP Outlet and ETP Outlet_Flow is live on the CPCB.we are not getting counts on the modscan from STACK 1_85TPH Boiler. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24142104261 | 2024-12-30 | 2025-01-02 | Cust 1255 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera is not playing on CPCB server. | We took access to the system and now camera is live on CPCB server. | We took access to the system and now camera is live on CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24142154262 | 2024-12-30 | 2025-01-02 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on CPCB Server for Shri Dnyaneshwar Sahakari Sakhar Karkhana Ltd,DP | As of now we are closing this complaint. We will reopen once the camera issue is resolved from customer end. | As of now we are closing this complaint. We will reopen once the camera issue is resolved from customer end. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24142414263 | 2024-12-30 | 2025-01-02 | Cust 1297 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on LOCAL+CPCB server. | We took access to the system and now data is live on CPCB+LOCAL Server. | We took access to the system and now data is live on CPCB+LOCAL Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24143074264 | 2024-12-31 | 2024-12-31 | Cust 904 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC Visit | AMC Visit done | AMC Visit done | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFPN24143524265 | 2024-12-31 | 2025-01-03 | Cust 1272 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on CPCB + KSPCB servers. | Now the data is live on CPCB + KSPCB servers. | Now the data is live on CPCB + KSPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24143614266 | 2024-12-31 | 2025-01-07 | Cust 1242 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC | amc work complete | completed the work | NO | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24144184267 | 2024-12-31 | 2024-12-31 | Cust 1354 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | pH Transmitter Aquamon | Aquamon 2000 plus(AQ2000 plus) | COMMISSIONING OF THE AUTO pH, CLEAN MAX SYSTEM | COMMISSIONING OF THE AUTO pH, CLEAN MAX SYSTEM successfully | completed the work | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24144224268 | 2024-12-31 | 2025-01-02 | Cust 632 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data uploading issue. | We have made some changes from configuration end and updated Enviroconnect DAS version. Now data is live on the CPCB server. Also camera is now live. | We have made some changes from configuration end and updated Enviroconnect DAS version. Now data is live on the CPCB server. Also camera is now live. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24144604269 | 2024-12-31 | 2025-01-09 | Cust 936 | PUNE | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | for installation purpose | problem solved | problem solved | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24144634270 | 2024-12-31 | 2025-01-11 | Cust 1260 | PUNE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | for calibration job | for calibration job | completed the work | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24144654271 | 2024-12-31 | 2025-01-07 | Cust 1266 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | Xenon Lamp Power supply for Multiparameter(AWA-XLAMPPS) | xenon power supply card not working | xenon power supply card not working | completed the work | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24145204272 | 2025-01-02 | 2025-01-05 | Cust 1035 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data uploading issue to the Local, CPCB and MPCB servers | Data is not coming till the Modscan. As of now we are closing this complaint, we will reopen this once the issue is resolved from hardware side. | Data is not coming till the Modscan. As of now we are closing this complaint, we will reopen this once the issue is resolved from hardware side. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24145224273 | 2025-01-02 | 2025-01-02 | Cust 1307 | PUNE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Smartpro-con) | Calibration with conductivity simulator. | Calibration done. | Closed | NO | Paid | Process | N |  | Analytical - CHAKAN
AFPN24145224273 | 2025-01-02 | 2025-01-02 | Cust 1307 | PUNE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Smartpro-con) | Calibration with conductivity simulator. | Calibration done. | Closed | NO | Paid | Process | N |  | Analytical - CHAKAN
AFPN24145244274 | 2025-01-02 | 2025-01-02 | Cust 1307 | PUNE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Chlorine sensor | CHLORINE SENSOR(1) | Calibration of Chlorine Analyser | Calibration has been done. | Closed | NO | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN24145284275 | 2025-01-02 | 2025-01-03 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data upload issue. | we took access of the black box and observed that the internet is not coming til the FMBB, now the same issue has been resolved from your side and data is live on CPCB + MPCB server. | we took access of the black box and observed that the internet is not coming til the FMBB, now the same issue has been resolved from your side and data is live on CPCB + MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24145374276 | 2025-01-02 | 2025-01-09 | Cust 1313 | PUNE | Analytical - CHAKAN | CLOSED | ENGINEERING ISSUE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | splitter not working | Splitter found faulty | problem solved | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24145454277 | 2025-01-02 | 2025-01-03 | Cust 1357 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on MPCB server. | We took access to the system and now data is live on MPCB server. | We took access to the system and now data is live on MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24145794278 | 2025-01-03 | 2025-01-07 | Cust 1360 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB connectivity needs to be done. | As of now we are closing this complaint.We will reopen it once we get system access. | As of now we are closing this complaint.We will reopen it once we get system access. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24146114279 | 2025-01-03 | 2025-01-06 | Cust 517 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Display Board Data update issue. | we took access and made some changes in the configuration so now the Data is live on the Display Board. | we took access and made some changes in the configuration so now the Data is live on the Display Board. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24146244280 | 2025-01-04 | 2025-01-08 | Cust 1355 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | Mother board-PCX2 of Multiparameter Analyzer(1) | need to upgrade the software version so that both MP and Flow meter data can be transferred to CPCB | on site only , updated MP software version and also data is live on CPCB server for both MP , FLow meter and dust . | on site only , updated MP software version and also data is live on CPCB server for both MP , FLow meter and dust . | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24146824281 | 2025-01-06 | 2025-01-09 | Cust 464 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on CPCB server. | We took access to the system and Camera is playing on CPCB Server. | We took access to the system and Camera is playing on CPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24146844282 | 2025-01-06 | 2025-01-09 | Cust 991 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Hardware and software data mismatch. | We took access to the system and added scaling on CPCB+LOCAL and now data is matching on CPCB server. | We took access to the system and added scaling on CPCB+LOCAL and now data is matching on CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24146864283 | 2025-01-06 | 2025-01-09 | Cust 1320 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on MPCB server. | There was customer side issue. Data is live now on CPCB server and same confirmed by customer. | There was customer side issue. Data is live now on CPCB server and same confirmed by customer. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24146994284 | 2025-01-06 | 2025-01-09 | Cust 1035 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to CPCB server. | As of now we are closing this complaint. We will reopen this once the issue is resolved from the hardware side. | As of now we are closing this complaint. We will reopen this once the issue is resolved from the hardware side. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24147234285 | 2025-01-07 | 2025-01-10 | Cust 1308 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB+MPCB server. | We took access to the system and now data is Updating on CPCB+MPCB server. | We took access to the system and now data is Updating on CPCB+MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24147394286 | 2025-01-07 | 2025-01-10 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB Data uploading issue | Data from Stack1_Boiler(120TPH) is now live on the CPCB server. | Data from Stack1_Boiler(120TPH) is now live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24147414287 | 2025-01-07 | 2025-01-12 | Cust 1282 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC need to be done for device and connectivity | AMC has done, need to give data online | AMC has done | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24147444288 | 2025-01-07 | 2025-01-10 | Cust 1361 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data mismatch on CPCB+LOAL+KSPCB | We took access to the system and changed the scaling on LOCAL+KSPCB and now data is matching with the CPCB server. | We took access to the system and changed the scaling on LOCAL+KSPCB and now data is matching with the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24147524289 | 2025-01-07 | 2025-01-11 | Cust 1332 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4 trans-A4TW033000(1) | commissioning of the ph and conductivity | commissioning of the ph and conductivity | completed the work | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24147604290 | 2025-01-07 | 2025-01-12 | Cust 1362 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Guidelines for installation | Guidelines given material yet to install, Nothing is ready | Guidelines has given for installation | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24147614291 | 2025-01-08 | 2025-01-12 | Cust 1354 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | CLEANMAX system commissioning | Site attended for CLEANMAX and system performance checked and found to be satisfactory. | Site attended for CLEANMAX and system performance checked and found to be satisfactory. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24148034292 | 2025-01-08 | 2025-01-10 | Cust 1296 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Camera not playing on CPCB server. | As of now we are closing this complaint we will reopen once we receive reply from customer. | As of now we are closing this complaint we will reopen once we receive reply from customer. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24148124293 | 2025-01-08 | 2025-01-10 | Cust 1360 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB connectivity needs to be done. | We have completed CPCB connectivity but currently data is not coming till the PC, so for now we are closing this complaint. | We have completed CPCB connectivity but currently data is not coming till the PC, so for now we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24148564294 | 2025-01-09 | 2025-01-12 | Cust 1259 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to MPCB server. | As of now we are closing this complaint. We will reopen this complaint once we get access to your system. | As of now we are closing this complaint. We will reopen this complaint once we get access to your system. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24148614295 | 2025-01-09 | 2025-01-10 | Cust 603 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on CPCB server. | Camera is not playing on VLC / Google informed customer to check the IP from their end. | Camera is not playing on VLC / Google informed customer to check the IP from their end. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24148784296 | 2025-01-09 | 2025-01-12 | Cust 1035 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB and MPCB data upload issue. | As of now we are closing this complaint. We will reopen this complaint once the issue is resolved from the hardware side. | As of now we are closing this complaint. We will reopen this complaint once the issue is resolved from the hardware side. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24148804297 | 2025-01-09 | 2025-01-12 | Cust 1344 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data breaking issue. | We have updated the new patch in FMBB, and now data is matching with the actual reading for all parameters. | We have updated the new patch in FMBB, and now data is matching with the actual reading for all parameters. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24148824298 | 2025-01-09 | 2025-01-10 | Cust 1284 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on CPCB server for Stack 1_22 TPH Incineration Boiler. | After changing the device ID on the analyzer, the data for Stack 1_22 TPH Incineration Boiler started showing live on the CPCB server. However, we have observed that the data keeps breaking continuously on the server. So this issue is from the hardware side . so we are closing this compliant for now. | the data for Stack 1_22 TPH Incineration Boiler started showing live on the CPCB server. However, we have observed that the data keeps breaking continuously on the server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24149004299 | 2025-01-10 | 2025-01-10 | Cust 603 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data uploaded issue. | We took access of the system and observe that the Apache servers was not running and after restarting the services the data started updating. | We took access of the system and observe that the Apache servers was not running and after restarting the services the data started updating. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24149174300 | 2025-01-10 | 2025-01-13 | Cust 1355 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | KSPCB connectivity needs to be done. | We took access of the black box and completed KSPCB connectivity, now data is live on the KSPCB server for all parameters. | We took access of the black box and completed KSPCB connectivity, now data is live on the KSPCB server for all parameters. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24149194301 | 2025-01-10 | 2025-01-11 | Cust 1232 | PUNE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | TSS CALIBRATION WAS PENDING | TSS CALIBRATION WAS PENDING | analyser maintananace has been done &amp;amp; calibration pending | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFPN24149204302 | 2025-01-10 | 2025-01-10 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on CPCB server. | We took access to the system and downgraded the firefox version now both cameras are live on cpcb server. | We took access to the system and downgraded the firefox version now both cameras are live on cpcb server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24149254303 | 2025-01-10 | 2025-01-10 | Cust 1245 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB DATA UPLODE ISSUE. | Not found any issue as the Data was live. | Not found any issue as the Data was live. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24149874304 | 2025-01-13 | 2025-01-16 | Cust 1259 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | data upload issue on MPCB server | issue has been resolved and data is live on the MPCB server. | issue has been resolved and data is live on the MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24149954305 | 2025-01-13 | 2025-01-16 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Display Board configuration needs to be done. | Successfully completed Display board configuration, now display board files are genrating. | Successfully completed Display board configuration, now display board files are genrating. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24150194306 | 2025-01-13 | 2025-01-16 | Cust 1035 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to cpcb server | We have upgraded an Enviroconnect DAS version to the latest and now data is live on the Local, CPCB and MPCB servers. | We have upgraded an Enviroconnect DAS version to the latest and now data is live on the Local, CPCB and MPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24150254307 | 2025-01-13 | 2025-02-17 | Cust 1298 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Amc of the device | AMC done | Closed | NO | Periodic Maintenance (AMC) | Core | N |  | Codel Products - CHAKAN
AFPN24150644308 | 2025-01-14 | 2025-01-20 | Cust 1298 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Maintenance under AMC visit | AMC visit done. Necessary maintenance activities carried out for Mutlipara.,dust | AMC visit done. Necessary maintenance activities carried out for Mutlipara.,dust | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24151184309 | 2025-01-15 | 2025-01-17 | Cust 1363 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Need to shift MPA connectivity from scanner to RS485 card | We have shifted Stack connectivity from the scanner to the RS485 card and data is uploading on the CPCB Server. | We have shifted Stack connectivity from the scanner to the RS485 card and data is uploading on the CPCB Server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24151194310 | 2025-01-15 | 2025-01-17 | Cust 1251 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | MPCB data uploaded issue for Ro Outelt | We took access of the BB and found a time sync issue and now we have synchronised the BB time with thr real-time so now the Data is live on the server. | We took access of the BB and found a time sync issue and now we have synchronised the BB time with thr real-time so now the Data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24151634311 | 2025-01-16 | 2025-01-17 | Cust 1035 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Camera streaming is not live on the CPCB server. | Camera streaming is now live on the CPCB server. | Camera streaming is now live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24151714312 | 2025-01-16 | 2025-01-17 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data uploading issue on cpcb and mpcb server. | Now the data is live on the mpcb and cpcb server. | Now the data is live on the mpcb and cpcb server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24151784313 | 2025-01-16 | 2025-01-19 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data upload issue. | as of now we are closing thgis complaint. we will reopen this once we get access to the system, | as of now we are closing this complaint. we will reopen this once we get access to the system, | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24151794314 | 2025-01-16 | 2025-01-24 | Cust 1336 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | no problem device is in healthy condition | problem solved | problem not solved | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24151804315 | 2025-01-16 | 2025-01-24 | Cust 1313 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | cleaning sample error | fill sample tank error ,problem solved | problem solved no mom made | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFPN24152054316 | 2025-01-17 | 2025-01-20 | Cust 1270 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue | Now the data is live on the server. | now the data is live on the server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24152264317 | 2025-01-17 | 2025-01-20 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data mismatch issue. | As of now we are closing this complaint and will reopen this case once we get remote access to the FMBB from customers end. | As of now we are closing this complaint and will reopen this case once we get remote access to the FMBB from customers end. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24152284318 | 2025-01-17 | 2025-01-20 | Cust 1282 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software re-installation | We have re-installed the Enviroconnect software with an upgraded version and completed Local,CPCB & MPCB configuration and data is now live on the same server. Also upgraded the Enviroconnect DAS version to the latest. | We have re-installed the Enviroconnect software with an upgraded version and completed Local,CPCB &amp;amp;amp; MPCB configuration and data is now live on the same server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24152314319 | 2025-01-17 | 2025-01-20 | Cust 1289 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB DATA UPLOADED ISSUE | Now the Data is live on the MPCB server | Now the Data is live on the MPCB server | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24152624320 | 2025-01-18 | 2025-01-22 | Cust 1270 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | CPCB and MPCB data upload issue | We have completed Local, CPCB and MPCB configuration on a Windows based PC and the data is live on the same servers. Also we have upgraded the Enviroconnect DAS version. | We have completed Local, CPCB and MPCB configuration on a Windows based PC and the data is live on the same servers. Also we have upgraded the Enviroconnect DAS version. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24152844321 | 2025-01-19 | 2025-01-24 | Cust 1346 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | pH input card-FM of Multiparameter(1) | installation of sample motor and ph input card | problem solved | problem solved | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24153524322 | 2025-01-20 | 2025-01-23 | Cust 1035 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCb data upload issue. | Now the issue has been resolved and data is live on the CPCB server. | Now the issue has been resolved and data is live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24153604323 | 2025-01-20 | 2025-01-22 | Cust 1308 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Value upload issue for Stack 1_2X20TPH Boiler | we accessed the system and now the values for Stack 1_2X20TPH Boiler are now matching on the CPCB, Local & MPCB Servers . | we accessed the system and now the values for Stack 1_2X20TPH Boiler are now matching on the CPCB, Local &amp;amp; MPCB Servers . | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24153674324 | 2025-01-20 | 2025-01-25 | Cust 1248 | PUNE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | Xenon Lamp Power supply for Multiparameter(AWA-XLAMPPS) | Xenon lamp power module was found faulty, Tubers were found brittle, Acid fumes evidences were also found. | Will submit the offer for repair and for new analyser model CX2000 as well. | Closed | NO | Paid | Process | N |  | Analytical - CHAKAN
AFPN24154364325 | 2025-01-22 | 2025-01-24 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | data upload issue ON CPCB + Local + MPCB servers. | We took access to the system and the data is now live on the CPCB + Local + MPCB servers. | We took access to the system and the data is now live on the CPCB + Local + MPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24154584326 | 2025-01-22 | 2025-01-27 | Cust 1269 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Silica Analyser | Hach Silica 5500 series(Hach Silica 5500 series) | AWA SILICA ANLYSER CL-1000 FAT at ADAGE Automation | Analyser functionality checked and FAT done. | Analyser functionality checked and FAT done. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24154634327 | 2025-01-22 | 2025-01-24 | Cust 1035 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data uploading issue on CPCB server. | Data is now live on the CPCB and MPCB server. | Data is now live on the CPCB and MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24155154328 | 2025-01-23 | 2025-01-24 | Cust 1035 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue. | We accessed the system and observed that some software related files were missing from the PC. Hence we have re-installed the software with an upgraded version and completed Local, CPCB and MPCB connectivity. Also updated Enviroconnect DAS version. Now the data is reflecting to the same servers. | re-installed the software with an upgraded version and completed Local, CPCB and MPCB connectivity. Also updated Enviroconnect DAS version. Now the data is reflecting to the same servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24155174329 | 2025-01-23 | 2025-01-27 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue. | We have changed the configuration for Stack_70TPH from Scanner to RS485, now data is live on CPCB server | We have changed the configuration for Stack_70TPH from Scanner to RS485, now data is live on CPCB server | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24155294330 | 2025-01-23 | 2025-02-02 | Cust 1256 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioned the Xenon power supply | spare connected the Xenon power supply | completed the work | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24155314331 | 2025-01-23 | 2025-02-07 | Cust 803 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning of repaired analyser | Commissioning done and suggestions given | commission has been done working properly customer satisfied | NO | Paid | Core | N |  | Analytical - CHAKAN
AFPN24155364332 | 2025-01-24 | 2025-02-04 | Cust 1364 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | no problem found | work done | commission has been done working properly customer satisfied | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24155374333 | 2025-01-24 | 2025-02-04 | Cust 1362 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | visit for commissioning | work done | work completed | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24155484334 | 2025-01-24 | 2025-01-24 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB connectivity needs to be done for Dudhaganga Vedganga | We have completed the MPCB connectivity and the Data is live on the MPCB server. | We have completed the MPCB connectivity and the Data is live on the MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24155504335 | 2025-01-24 | 2025-01-24 | Cust 936 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Request for Company name change | As per the CPCB's approval, we have now updated the industry name on our server. | As per the CPCB's approval, we have now updated the industry name on our server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24155554336 | 2025-01-24 | 2025-01-25 | Cust 731 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera streaming issue on FM CPCB server. | We have accessed the system, and the issue has now been resolved. The camera streaming is live on the CPCB server. | We have accessed the system, and the issue has now been resolved. The camera streaming is live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24155564337 | 2025-01-24 | 2025-01-27 | Cust 1270 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Flow data not uploading on localhost server. | We have accessed the system, and the flow values are now being uploaded to the local server. | We have accessed the system, and the flow values are now being uploaded to the local server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24155724338 | 2025-01-24 | 2025-02-04 | Cust 1271 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | pH input card-FM of Multiparameter(1) | received confirmation from customer for the receipt of the spares on site | work done | customer satisfied with values and suggestions | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24155734339 | 2025-01-24 | 2025-01-25 | Cust 651 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue. | We have accessed the system, and the data is now live on both CPCB and local servers for the Distillery Outlet to Lagoon and Concentrated Spent Wash Inlet to Boiler_massflow. As discussed during our call, there is a hardware issue with Stack 1_ESP_35 TPH_Incineration_Boiler. | We have accessed the system, and the data is now live on both CPCB and local servers for the Distillery Outlet to Lagoon and Concentrated Spent Wash Inlet to Boiler_massflow. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24156674340 | 2025-01-26 | 2025-01-28 | Cust 1313 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4 trans-A4TW033000(1) | pH and DO dual channel aqua4trans | visit done. Necessary maintenance activities carried out for pH transmitter and electrode. | visit done. Necessary maintenance activities carried out for pH transmitter and electrode. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24157114341 | 2025-01-27 | 2025-03-12 | Cust 803 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Commissioning of blackbox to replace the local host | Visited the site and checked the issue of display. Need to lay wire between analyser and Gate display. | Visited the site and | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24157804342 | 2025-01-28 | 2025-01-30 | Cust 1270 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on MPCB server. | We have accessed the system, and the issue has been resolved. The data is now live on the MPCB server. | We have accessed the system, and the issue has been resolved. The data is now live on the MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24157924343 | 2025-01-28 | 2025-01-31 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | need to change configuration for stack from direct RS485 to scanner. | We have changed the configuration for stack from direct RS485 to scanner and data is now live on the CPCB server. | We have changed the configuration for stack from direct RS485 to scanner and data is now live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24158154344 | 2025-01-28 | 2025-01-31 | Cust 1035 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to CPCB server. | Data is now live on the CPCB and MPCB servers. | Data is now live on the CPCB and MPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24158304345 | 2025-01-28 | 2025-01-29 | Cust 904 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC visit no-2 | AMC visit done | Connected air as motive of 5.5 bar and issue resolved | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFPN24159034346 | 2025-01-29 | 2025-01-31 | Cust 1035 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to MPCB server. | Issue resolved from customer end and Sugar unit data is live in the MPCB portal. | Issue resolved from customer end and Sugar unit data is live in the MPCB portal. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24159224347 | 2025-01-29 | 2025-01-31 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to MPCB server. | Data is live on the MPCB server. | Data is live on the MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24159234348 | 2025-01-29 | 2025-02-01 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Localhost configuration and data not updating on MPCB portal - Shri Dudhganga Vedganga Sahakari Karkhana ltd | We took access to the system and completed the software configuration for LOCAL+MPCB. | We took access to the system and completed the software configuration for LOCAL+MPCB. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24159894349 | 2025-01-30 | 2025-01-31 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue | We have upgraded Enviroconnect DAS version to the latest. Data is live on the CPCB and MPCB server. | We have upgraded Enviroconnect DAS version to the latest. Data is live on the CPCB and MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24161324350 | 2025-01-31 | 2025-02-03 | Cust 1343 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MUVX Card of Multiparameter(1) | MUVX card issue | MUVX card issue offer shared | amc done | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFPN24162454351 | 2025-02-03 | 2025-02-06 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Configuration needs to be change from direct RS 485 to scanner | We have changed the configuration for STACK_2_BOILER_25TPH from direct RS 485 to scanner and data is now live on the Local, CPCB and MPCB servers. | We have changed the configuration for STACK_2_BOILER_25TPH from direct RS 485 to scanner and data is now live on the Local, CPCB and MPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24162474352 | 2025-02-03 | 2025-02-04 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Display Board configuration needs to be done. | We have completed Display board configuration, now JSON files are generated. | We have completed Display board configuration, now JSON files are generated. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24162994353 | 2025-02-03 | 2025-02-05 | Cust 1261 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Reinstallation and configuration of software. | We took access to the system and completed the installation and configuration od software and now data is live on CPCB+LOCAL+MPCB server. | We took access to the system and completed the installation and configuration od software and now data is live on CPCB+LOCAL+MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24163024354 | 2025-02-03 | 2025-02-05 | Cust 1306 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on CPCB server. | There was customer side issue which has been resolved and now camera is live on CPCB server. | There was customer side issue which has been resolved and now camera is live on CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24163234355 | 2025-02-04 | 2025-02-07 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB approval received need to add STACK under LOCAL+ MPCB => Loknete Marutrao Ghule Patil Dnyaneshwar Sahakari Sakhar Karkhana Ltd | We took access to the system and completed the LOCAL+MPCB configuration and now data is live on both the servers. | We took access to the system and completed the LOCAL+MPCB configuration and now data is live on both the servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24163244356 | 2025-02-04 | 2025-02-16 | Cust 1268 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | FM make current output card(1) | mA card needs to insatll | mA card installed data is live and analyzer is in healthy condition | mA card replaced Analyzer data is live and in healthy condition | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFPN24163294357 | 2025-02-04 | 2025-02-07 | Cust 1289 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB connectivity needs to be done. | We have completed CPCB connectivity and data is now live on the same server. Also we have updated Enviroconnect DAS version to the latest. | We have completed CPCB connectivity and data is now live on the same server. Also we have updated Enviroconnect DAS version to the latest. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24163444358 | 2025-02-04 | 2025-02-06 | Cust 936 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | MPCB connectivity needs to be done. | we have completed MPCB configuration, now data is live on MPCB server. | we have completed MPCB configuration, now data is live on MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24163904359 | 2025-02-05 | 2025-02-07 | Cust 1273 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB+MPCB server. | We took access to the system and now data is updating on CPCB server. | We took access to the system and now data is updating on CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24163964360 | 2025-02-05 | 2025-02-07 | Cust 1035 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB server for Yogeshwari Sugar Industries Pvt Ltd | We took access to the system and now data is updating on CPCB server. | We took access to the system and now data is updating on CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24164514361 | 2025-02-07 | 2025-02-16 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Issue with Xenon power supply | issue with the xenon card and photodetector. Analyzer needs to repair | needs to repair after the shut down | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFPN24164524362 | 2025-02-07 | 2025-02-16 | Cust 936 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Online connectivity of MPA | Data is online for multiparameter analyzer and flowmeter | Blackbox commissioned data is online | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFPN24164534363 | 2025-02-07 | 2025-02-07 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Display board ip address needs to be change | we took access of the system and changed the Display Board IP address from 192.168.0.102 to 192.168.1.103, also did some modifications in configuration now the JSON is successfully generated and also Flow data matching with the actual reading | now the JSON is successfully generated and also Flow data matching with the actual reading | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24164544364 | 2025-02-07 | 2025-02-10 | Cust 1289 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data breaking issue on MPCB server. | Data is now continuously reflected on the MPCB server. | Data is now continuously reflected on the MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24164554365 | 2025-02-07 | 2025-02-07 | Cust 936 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Data not updating on MPCB server. | We took access to the system and now data is updating on MPCB server. | We took access to the system and now data is updating on MPCB server. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFPN24164774366 | 2025-02-08 | 2025-02-16 | Cust 517 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC of the device | AMC has done checked the device need spare for pH | AMC has done | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24165484367 | 2025-02-10 | 2025-02-13 | Cust 1270 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to CPCB server. | Accessed system and not getting any counts on the modscan for ETP Flow2_ Lagoon Area. Now the issue has been resolved from hardware end and data is showing live on the CPCB server. | Accessed system and not getting any counts on the modscan for ETP Flow2_ Lagoon Area. Now the issue has been resolved from hardware end and data is showing live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24165514368 | 2025-02-10 | 2025-02-13 | Cust 1273 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Stack data need to configure on MPCB server. | Are are closing this complaint we will re-open once the system is accessible. | Are are closing this complaint we will re-open once the system is accessible. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24165634369 | 2025-02-10 | 2025-02-16 | Cust 1281 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | ph sensor faulty | need to replace pH ssensor | Sensor found faulty needs to replace pH sensor | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFPN24165744370 | 2025-02-10 | 2025-02-13 | Cust 1035 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Flow data is not updating on CPCB server for Yogeshwari Sugar Industries Pvt Ltd | We took access to the system and data is not coming till MODSCAN we will take access once issue is resolved from hardware end. | We took access to the system and data is not coming till MODSCAN we will take access once issue is resolved from hardware end. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24165784371 | 2025-02-10 | 2025-02-11 | Cust 618 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB+MPCB server. | We took access to the system and observed that there was time sync issue which has been resolved and now data is live on CPCB and MPCB server. | We took access to the system and observed that there was time sync issue which has been resolved and now data is live on CPCB and MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24166184372 | 2025-02-11 | 2025-02-24 | Cust 1338 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | 0 Stack data updating and cameras not playing on CPCB server. | As of now we are closing this complaint.We will re-open the complaint once current output checked from your end. | As of now we are closing this complaint.We will re-open the complaint once current output checked from your end. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24166474373 | 2025-02-12 | 2025-02-14 | Cust 936 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPC B connectivity needs to be done. | we took access of the black box and completed CPCB connectivity now data is live on CPCB server. | we took access of the black box and completed CPCB connectivity now data is live on CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24166544374 | 2025-02-12 | 2025-02-13 | Cust 1002 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on LOCAL+CPCB+MPCB server | We took access to the system and now data is live on LOCAL+CPCB+MPCB server. | We took access to the system and now data is live on LOCAL+CPCB+MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24167014375 | 2025-02-13 | 2025-02-17 | Cust 936 | PUNE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | PH input card found faulty | Work has been done. | work done | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24167424376 | 2025-02-13 | 2025-02-14 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on MPCB server. | Data is now live on the MPCB server after restarting the services. | Data is now live on the MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24167444377 | 2025-02-13 | 2025-02-16 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data upload isssue. | As of now, we are closing this complaint. We will reopen it once we get access to the system. | As of now, we are closing this complaint. We will reopen it once we get access to the system. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24167594378 | 2025-02-13 | 2025-03-04 | Cust 1289 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | Cleaning Pump(1) | Cleaning pump failure running analyser | Cleaning pump failure running analyser | completed the work | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24167764379 | 2025-02-14 | 2025-02-16 | Cust 872 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation + CPCB, Local and PPCB configuration | We have re-installed the software with an updated version and completed Local, CPCB and PPCB connectivity. Data is now live on the same servers. | We have re-installed the software with an updated version and completed Local, CPCB and PPCB connectivity. Data is now live on the same servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24167774380 | 2025-02-14 | 2025-02-14 | Cust 1273 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Configuration of STACK on MPCB server. | We took access to the system and completed the STACK configuration on MPCB server. Now data is live on both CPCB+MPCB server. | We took access to the system and completed the STACK configuration on MPCB server. Now data is live on both CPCB+MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24168544381 | 2025-02-17 | 2025-02-25 | Cust 811 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | sample tubes damaged | Checked the analyzer found ok but needs to change the tubings | Guidelines has given for installation | NO | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24168774382 | 2025-02-17 | 2025-03-04 | Cust 936 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | pH input card-FM of Multiparameter(1) | PH CARD FAILURE in new CX-2000 ANALYSER Sr.NO-2412089 | PH CARD FAILURE in new CX-2000 ANALYSER Sr.NO-2412089 | completed the work | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24168964383 | 2025-02-18 | 2025-02-18 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB + MPCB data upload issue. | We have changed the configuration for MPL (from Scanner to RS485) and Flow meter (from RS485 to Scanner), now data is live on CPCB + MPCB server. | We have changed the configuration for MPL (from Scanner to RS485) and Flow meter (from RS485 to Scanner), now data is live on CPCB + MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24169164384 | 2025-02-18 | 2025-02-21 | Cust 1035 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | local and MPCB data upload issue | Now the issue has been resolved and data is live on the Local & MPCB server. | Now the issue has been resolved and data is live on the Local &amp;amp; MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24169184385 | 2025-02-18 | 2025-02-21 | Cust 1251 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | CPCB data uploded issue. | We took access of the system today. During the access, we observed that there was a loose wiring connection. This issue has now been resolved by our site engineer, and the Flow data has started updating to the server. The Camera streaming is also live on the CPCB server. | During the access, we observed that there was a loose wiring connection. This issue has now been resolved by our site engineer, and the Flow data has started updating to the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24169204386 | 2025-02-18 | 2025-02-21 | Cust 795 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data uploded issue. | Not found any issue as the data was live. | Not found any issue as the data was live. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24169614387 | 2025-02-19 | 2025-02-22 | Cust 1365 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCP+MPCB data upload issue. | As of now we are closing this complaint and we will re-open this complaint once the system is accessible. | As of now we are closing this complaint and we will re-open this complaint once the system is accessible. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24169644388 | 2025-02-19 | 2025-02-21 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on CPCB server. | We took access to the system and found camera streaming window was closed .Now camera is live on CPCB server. | We took access to the system and found camera streaming window was closed .Now camera is live on CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24169664389 | 2025-02-19 | 2025-02-21 | Cust 1273 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Flow data not updating on MPCB server. | We took access to the system and found flow was not correctly configured on output configuration after changing data is live on MPCB server. | We took access to the system and found flow was not correctly configured on output configuration after changing data is live on MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24169704390 | 2025-02-19 | 2025-02-22 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | STACK data not updating for - Loknete Marutrao Ghule Patil Dnyaneshwar Sahakari Sakhar Karkhana [Formerly known as Shri Dnyaneshwar Sahakari Sakhar Karkhana Ltd.], DP | As of now we are closing this complaint and we will re-open this complaint once the system is accessible. | As of now we are closing this complaint and we will re-open this complaint once the system is accessible. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24169714391 | 2025-02-19 | 2025-02-21 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB & MPCB data upload issue | Data is now live on the CPCB and MPCB server. | Data is now live on the CPCB and MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24170094392 | 2025-02-20 | 2025-02-21 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | multiparameter analyzer data uploading issue on cpcb server. | We have made changes as per below mail and data is now live on the CPCB server. | We have made changes as per below mail and data is now live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24170244393 | 2025-02-20 | 2025-02-21 | Cust 795 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | BOD value mismatch on Hardware and CPCB server. | We have checked from our end and found the value has automatically returned within limits. There might be an issue from hardware side | We have checked from our end and found the value has automatically returned within limits. There might be an issue from hardware side | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24170654394 | 2025-02-21 | 2025-02-21 | Cust 1318 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to MPCB server. | Issue has been resolved and data is now live on the MPCB server. | Issue has been resolved and data is now live on the MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24171034395 | 2025-02-24 | 2025-03-04 | Cust 1355 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | Mother board-PCX2 of Multiparameter Analyzer(1) | MP cleaning pump head assembly damaged | comp logged for RF , material received , visit will be planned as soon as received confirmation from customer. | comp logged for RF , material received , visit will be planned as soon as received confirmation from customer. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFPN24171044396 | 2025-02-24 | 2025-03-11 | Cust 1344 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | Peristaltic sampling pump motor for Multiparameter(AWA-PSPM) | Sample pump not working properly , big noise in the pump | REF : OUR RF OA NO.: 5525800776 DT: 18-FEB-2025 OUR RF INV. NO : FC/RF552511691 DT. 18-FEB-25 | complaint logged for the RF purpose. REF : OUR RF OA NO.: 5525800776 DT: 18-FEB-2025 OUR RF INV. NO : FC/RF552511691 DT. 18-FEB-25 | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFPN24171504397 | 2025-02-24 | 2025-02-27 | Cust 991 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB server. | We took access to the system and now data is live on CPCB+MPCB server. | We took access to the system and now data is live on CPCB+MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24171694398 | 2025-02-24 | 2025-02-27 | Cust 1366 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB+MPCB server. | As of now we are closing this complaint. We will re-open this issue once the PC is accessible. | As of now we are closing this complaint. We will re-open this issue once the PC is accessible. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24172024399 | 2025-02-25 | 2025-02-28 | Cust 603 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data uploded issue. | The issue has been resolved and now the Data is live on the MPCB server. | The issue has been resolved and now the Data is live on the MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24172074400 | 2025-02-25 | 2025-03-12 | Cust 1289 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Facing issues on both CPCB & MPCB portal data Updating | We have updated the software version but from 8th March there plant stopped for shutdown maintenance, now internet and converter cable removed. hence as of now we are closing this complaint. | We have updated the software version but from 8th March there plant stopped for shutdown maintenance, now internet and converter cable removed. hence as of now we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24172134401 | 2025-02-25 | 2025-02-28 | Cust 1275 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on CPCB server. | As of now we are closing this complaint.We will re-open the complaint once current output checked from your end. | As of now we are closing this complaint.We will re-open the complaint once current output checked from your end. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24172144402 | 2025-02-25 | 2025-02-28 | Cust 1306 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB+MPCB portal | As of now we are closing this complaint. We will re-open the hardware issue is resolved from your end. | As of now we are closing this complaint. We will re-open the hardware issue is resolved from your end. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24172154403 | 2025-02-25 | 2025-02-25 | Cust 1344 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | Peristaltic sampling pump motor for Multiparameter(AWA-PSPM) | Sampling pump is not working | Need to repair or replace the sampling pump with the new one | required repair | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24172174404 | 2025-02-25 | 2025-02-25 | Cust 1289 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | Peristaltic cleaning pump head for Multiparameter(AWA-PCPH) | Replaced the new pump head by faulty one | Analyzer Pump head replaced with the new one due to its in warranty | Commissioned the repaired DDu and replaced the new pump head for the Analyzer | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24172174404 | 2025-02-25 | 2025-02-25 | Cust 1289 | PUNE | Codel Products - CHAKAN | CLOSED | AMC | DCEM21XX | DDU DCEM21XX(DDU DCEM21XX) | Replaced the new pump head by faulty one | DDU repaired and commissioned and checked the values found satisfactory | Commissioned the repaired DDu and replaced the new pump head for the Analyzer | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24172304405 | 2025-02-25 | 2025-03-06 | Cust 936 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | PH Sensor for Multiparameter(AWA-PH-SENS) | pH sensor failed withing 2 month of commissioning , not getting calibrate with buffer solution | comp logged for RF purpose. material receievd . | comp logged for RF purpose. material receievd . | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFPN24172324406 | 2025-02-25 | 2025-03-11 | Cust 1344 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | Peristaltic cleaning pump motor for Multiparameter(AWA-PCPM) | Sample pump motor not working , making sound | RF No 138255627 Approval Required : FW: ETP system not working-Reg comp logged for the RF purpose, | RF No 138255627 Approval Required : FW: ETP system not working-Reg comp logged for the RF purpose, | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFPN24172534407 | 2025-02-26 | 2025-02-27 | Cust 1367 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB+MPCB server and reports are not generating. | We took access to the system and CPCB+MPCB server and reports are also regenerating and same shared with customer. | We took access to the system and CPCB+MPCB server and reports are also regenerating and same shared with customer. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24172584408 | 2025-02-26 | 2025-02-28 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software re-installation | We have re-installed the software with an updated version and completed Local, CPCB and MPCB connectivity. Data is now live on the same servers. | We have re-installed the software with an updated version and completed Local, CPCB and MPCB connectivity. Data is now live on the same servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24172684409 | 2025-02-26 | 2025-02-28 | Cust 795 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data mismatch issue. | We have checked and observed that the CPCB readings are matching with the actual readings of the analyzer and the same is verified by the customer also. | We have checked and observed that the CPCB readings are matching with the actual readings of the analyzer and the same is verified by the customer also. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24172704410 | 2025-02-26 | 2025-02-28 | Cust 1239 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB server. | We took access to the system and now data is live on CPCB server. | We took access to the system and now data is live on CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24173824411 | 2025-02-27 | 2025-02-28 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to CPCB and MPCB servers. | As of now we are closing this complaint. We will reopen this once we get access to the system. | As of now we are closing this complaint. We will reopen this once we get access to the system. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24173934412 | 2025-02-27 | 2025-02-28 | Cust 1035 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Flow meter data is not uploading to CPCB server. | Issue has been resolved and the flow meter data is now live on the CPCB server | Issue has been resolved and the flow meter data is now live on the CPCB server | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24174684413 | 2025-02-28 | 2025-03-11 | Cust 1314 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | Peristaltic cleaning pump head for Multiparameter(AWA-PCPH) | cleaning pump head not working | Jigar visited the site and issue resolved succesfully | Jigar visited the site and issue resolved succesfully | NO | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24176444414 | 2025-03-03 | 2025-03-12 | Cust 1301 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Troubleshooting | Calibration has done all the values are ok | calibration has done | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24176614415 | 2025-03-03 | 2025-03-06 | Cust 1306 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB server. | We took access to the system and changed the salve ID after which data is live on CPCB server. | We took access to the system and changed the salve ID after which data is live on CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24176634416 | 2025-03-03 | 2025-03-06 | Cust 1002 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB server. | There was hardware side issue from customer side which has been resolved and data is live on CPCB Server. | There was hardware side issue from customer side which has been resolved and data is live on CPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24176784417 | 2025-03-03 | 2025-03-12 | Cust 1292 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC visit for etp and Stack online monitoring system | AMC visit done. Necessary maintenance activities carried out for Mutlipara. and dust | AMC visit done. Necessary maintenance activities carried out for Mutlipara. and dust | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24177214418 | 2025-03-04 | 2025-03-12 | Cust 1368 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning of new Analyzer | Commissioning has been done | commission has been done working properly customer satisfied | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24177274419 | 2025-03-04 | 2025-03-07 | Cust 1268 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | KSPCB data uploaded issue. | not found any issue as the Data was live on the server. | not found any issue as the Data was live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24177524420 | 2025-03-05 | 2025-03-13 | Cust 1270 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Electrode | Forbes Marshall Cell Model 8311, K=0.1(FM -8311) | CONDUCTIVITY SENSOR COMMISSIONING WITH DPI -201 INDICATOR | COMMISSIONING WITH DPI -201 INDICATOR was done | completed the work | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24177694421 | 2025-03-05 | 2025-03-08 | Cust 1244 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | COD data is going High. | We took access to the system and observed COD modscan and software readings are matching. There is some hardware side issue. | We took access to the system and observed COD modscan and software readings are matching. There is some hardware side issue. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24177744422 | 2025-03-05 | 2025-03-13 | Cust 1344 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | Peristaltic sampling pump motor for Multiparameter(AWA-PSPM) | Replacement of motor with the new one | Replacement of motor with the new one | completed the work | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24177794423 | 2025-03-05 | 2025-03-16 | Cust 1313 | PUNE | Analytical - CHAKAN | CLOSED | ENGINEERING ISSUE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Splitter not working | work is done | work has been done | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24178124424 | 2025-03-06 | 2025-03-11 | Cust 1355 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | Peristaltic cleaning pump head for Multiparameter(AWA-PCPH) | receive spares and confirmation from customer to plan visit to attend the problem of the MP analyser , pump not working | visited the site problem resolved successfully .mail sent to customer is attached | visited the site problem resolved successfully .mail sent to customer is attached | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24178134425 | 2025-03-06 | 2025-03-11 | Cust 936 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | PH Sensor for Multiparameter(AWA-PH-SENS) | received spares from division for the pH of MP , need to plan visit for replacement and rectification of analyser issues | visited the site issue resolved successfully , mail sent to customer is attached | visited the site issue resolved successfully , mail sent to customer is attached | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24178254426 | 2025-03-06 | 2025-03-16 | Cust 1369 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Start off condition analyzer | work has been done | problem has been solved. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFPN24178314427 | 2025-03-06 | 2025-03-09 | Cust 803 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Installation and reconfiguration of software. | As of now we are closing this complaint we will re-open once we get full access to the system. | As of now we are closing this complaint we will re-open once we get full access to the system. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24178344428 | 2025-03-06 | 2025-03-17 | Cust 803 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Amc of analyser and online connectivity issue | visit done under EUT. Necessary maintenance activities carried out for Mutlipara., | customer satisfied with values and suggestions | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24179464429 | 2025-03-10 | 2025-03-13 | Cust 803 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Configuration of software. | We took access to the system and completed the software installation and configuration and now data is live on LOCAL+CPCB+MPCB server. | We took access to the system and completed the software installation and configuration and now data is live on LOCAL+CPCB+MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24179494430 | 2025-03-10 | 2025-03-13 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data mismatch issue | As of now we are closing this complaint. We will reopen this once we get access to the system | As of now we are closing this complaint. We will reopen this once we get access to the system | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24180134431 | 2025-03-10 | 2025-03-13 | Cust 1318 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation | Data is now live on the CPCB and MPCB servers. | Data is now live on the CPCB and MPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24180564432 | 2025-03-11 | 2025-03-13 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading on MPCB server. | As of now we are closing this complaint. We will reopen this once we get access to the system | As of now we are closing this complaint. We will reopen this once we get access to the system | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24180704433 | 2025-03-11 | 2025-03-14 | Cust 1370 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on MPCB server. | We took access to the system and now data is updating on MPCB server. | We took access to the system and now data is updating on MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24180834434 | 2025-03-11 | 2025-03-14 | Cust 1232 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Display Board configuration needs to be done through AIP. | We have share the Token details to the Aaxis nano team and they have configured the same hence we are closing this complaint. | We have share the Token details to the Aaxis nano team and they have configured the same hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24180854435 | 2025-03-11 | 2025-03-14 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to CPCB and MPCB servers. | Issue has been resolved and data is now live on the CPCB, Local and MPCB servers. | Issue has been resolved and data is now live on the CPCB, Local and MPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24181194436 | 2025-03-12 | 2025-03-14 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to CPCB and MPCB servers. | As of now we are closing this complaint. We will reopen this once we get access to the system. | As of now we are closing this complaint. We will reopen this once we get access to the system. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24181424437 | 2025-03-12 | 2025-03-13 | Cust 1251 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | DPI-Digital process indicator(1) | commissioning of 2 DO sensor | completed | completed the work | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24181424437 | 2025-03-12 | 2025-03-13 | Cust 1251 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | DPI-Digital process indicator(1) | commisssiopning | commisssiopning | completed the work | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24181454438 | 2025-03-12 | 2025-03-17 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data upload issue : waran SSK - DISTILLERY | Customer taken up the issue with MPCB team for correct CSV file and will reopen it once we get connect CSV file. | Customer taken up the issue with MPCB team for correct CSV file and will reopen it once we get connect CSV file. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24181694439 | 2025-03-13 | 2025-03-14 | Cust 1035 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software re-installation | We have reinstalled the software with an updated version and completed Local, CPCB and MPCB server connectivity. Data is now live on the same servers. | We have reinstalled the software with an updated version and completed Local, CPCB and MPCB server connectivity. Data is now live on the same servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24181984440 | 2025-03-14 | 2025-03-17 | Cust 1362 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB connectivity needs to be done. | We have completed the CPCB connectivity for the ETP device and Data is now live on the CPCB server. | We have completed the CPCB connectivity for the ETP device and Data is now live on the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24182064441 | 2025-03-14 | 2025-03-17 | Cust 1243 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | data mismatch issue. | Yesterday we took access of the black box and observed that the 0 readings are coming till the FMBB, so for now we are closing this complaint. | Yesterday we took access of the black box and observed that the 0 readings are coming till the FMBB, so for now we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24182324442 | 2025-03-16 | 2025-03-28 | Cust 517 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | PH Sensor for Multiparameter(AWA-PH-SENS) | ph sensor faulty | work done | work completed | NO | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24182674443 | 2025-03-17 | 2025-03-25 | Cust 1243 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | pH card burnout | Visit done. Necessary maintenance activities carried out for Multipara. Analyser | Visit done. Necessary maintenance activities carried out for Multipara. Analyser | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24183294444 | 2025-03-18 | 2025-03-24 | Cust 1371 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Electrode | Ph Electrode(FM1104(4C)) | PH ELECTRODE IS NOT RESPONDE TO buffer solution need to replace electrode | PH ELECTRODE IS NOT RESPONDE TO buffer solution need to replace electrode | completed the work | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24183424445 | 2025-03-18 | 2025-03-21 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Configuration for stack needs to be change. | We have made changes from configuration end and data is now live on the CPCB server. | We have made changes from configuration end and data is now live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24184334446 | 2025-03-20 | 2025-03-20 | Cust 1243 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | MPCB Data mismatch issue. | today we took access of the black box, and changed the Slave ID for MPL in controller after that issue has been resolved and data is matching with the actual readings. | today we took access of the black box, and changed the Slave ID for MPL in controller after that issue has been resolved and data is matching with the actual readings. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24184394447 | 2025-03-20 | 2025-03-25 | Cust 1269 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | FAT at customer place | Visit done for FAT for HC analyser, pH and conductivity. | Visit done for FAT for HC analyser, pH and conductivity. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24184454448 | 2025-03-20 | 2025-03-21 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | data is not uploading to CPCB and MPCB | Issue has been resolved and data is now live on the Local, CPCB and MPCB servers. | Issue has been resolved and data is now live on the Local, CPCB and MPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24184464449 | 2025-03-20 | 2025-03-25 | Cust 1260 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC visit | AMC visit done. Necessary maintenance activities carried out for Mutlipara Analyser. | AMC visit done. Necessary maintenance activities carried out for Mutlipara.Analyser. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24185594450 | 2025-03-24 | 2025-03-27 | Cust 1275 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not streaming on CPCB server. | As of now we are closing tis complaint we will reopen once the rtsp LINK is accessible. | As of now we are closing tis complaint we will reopen once the rtsp LINK is accessible. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24185694451 | 2025-03-24 | 2025-03-27 | Cust 1372 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on MPCB server. | As of now we are closing this complaint we will reopen the complaint once the system is accessible. | As of now we are closing this complaint we will reopen the complaint once the system is accessible. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24185874452 | 2025-03-24 | 2025-03-27 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data upload issue | Data is live on the MPCB server. | Data is live on the MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24186184453 | 2025-03-24 | 2025-03-29 | Cust 1308 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | pH card replacement | AMC DONE | completed the work | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24186374454 | 2025-03-25 | 2025-03-28 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to Local, CPCB and MPCB servers. | Data is now live on the Local, CPCB and MPCB servers. | Data is now live on the Local, CPCB and MPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24186854455 | 2025-03-26 | 2025-04-07 | Cust 809 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC of the Analyzer | AMC of the Analyzer | completed the work | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24186864456 | 2025-03-26 | 2025-04-07 | Cust 803 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC of the Analyzer | AMC of the Analyzer | completed the work | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24186864456 | 2025-03-26 | 2025-04-07 | Cust 803 | PUNE | Codel Products - CHAKAN | CLOSED | AMC | DCEM21XX | Transceiver DCEM21XX without cable(Transceiver DCEM21XX without cable) | AMC of the Analyzer | AMC of the Analyzer | completed the work | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24187324457 | 2025-03-27 | 2025-03-29 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB + MPCB + display board configuration needs to be done. | Software got crashed, so we have completed reinstallation of software + Local + CPCB + MPCB + Display board connectivity now data is live on all the servers. | Software got crashed, so we have completed reinstallation of software + Local + CPCB + MPCB + Display board connectivity now data is live on all the servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24187964458 | 2025-03-27 | 2025-03-28 | Cust 803 | PUNE | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | AQUADO Analyzer | Visiferm DO sensor(1) | Location needs to be changed. | Suggested to change the location and install it in either flowthrough chamber or on fermentation tank. | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN24188104459 | 2025-03-28 | 2025-03-31 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to CPCB and MPCB server. | As of now we are closing this complaint. We will reopen this once we get access to the system. | As of now we are closing this complaint. We will reopen this once we get access to the system. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24188154460 | 2025-03-28 | 2025-03-29 | Cust 1308 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB+MPCB server. | We took access to the system and observed there was hardware side issue now the issue has been resolved and data is live on LOCAL+CPCB+MPCB server. | We took access to the system and observed there was hardware side issue now the issue has been resolved and data is live on LOCAL+CPCB+MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24188204461 | 2025-03-28 | 2025-04-08 | Cust 1267 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | device not working | replaced spare and connectivity has been done. | completed the work | NO | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24188454462 | 2025-03-29 | 2025-04-04 | Cust 1373 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | AQUADO Analyzer | DO2 Arc Sensors(DO2 Arc Sensors) | ORP sensor and transmitter are not functioning Site address:- SPL infusion Pvt Ltd. Sr No 198, Hindustan Antibiotics Ltd, Bombay Pune road, Pimpri, Pune. | ORP sensor and transmitter are not functioning Site address:- SPL infusion Pvt Ltd. Sr No 198, Hindustan Antibiotics Ltd, Bombay Pune road, Pimpri, Pune. | completed the work | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN24188544463 | 2025-03-29 | 2025-03-29 | Cust 617 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Localhost not accessible and data not updating on MPCB server. | We took access to the system and observed that the services were stopped after restarting the services now data is live on LOCAL+MPCB server. | We took access to the system and observed that the services were stopped after restarting the services now data is live on LOCAL+MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN24188684464 | 2025-03-29 | 2025-04-11 | Cust 1242 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC | AMC visit done. Necessary maintenance carried out for Multiparameter. | AMC visit done. Necessary maintenance carried out for Multiparameter. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN24188714465 | 2025-03-29 | 2025-04-11 | Cust 517 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB AND MPCB DATA UPLODED ISSUE | We took access and found that the software data base was crashed hence we have reinstall the software and complete the connectivity and resolve the issue now the data is live. | We took access and found that the software data base was crashed hence we have reinstall the software and complete the connectivity and resolve the issue now the data is live. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN24190294466 | 2025-03-31 | 2025-04-03 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to local and MPCB server. | Data is now live on the Local & MPCB servers. Also we have updated an Enviroconnect DAS version. | Data is now live on the Local &amp;amp; MPCB servers. Also we have updated an Enviroconnect DAS version. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN25000054467 | 2025-04-01 | 2025-04-04 | Cust 1267 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | reinstallation and configuration of software. | We have installed and configured the software and for flow data is not coming till modscan so as of now we are closing this complaint. | We have installed and configured the software and for flow data is not coming till modscan so as of now we are closing this complaint. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN25000494468 | 2025-04-02 | 2025-04-04 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Data mismatch issue on CPCB server. | Issue has been resolved and Now the data is live on the CPCB server | Issue has been resolved and Now the data is live on the CPCB server | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN25000794469 | 2025-04-03 | 2025-04-04 | Cust 1035 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | data is not uploading to MPCB server | plant under shutdown | plant goes under shutdown | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN25001184470 | 2025-04-04 | 2025-04-07 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to CPCB and MPCB server. | As of now, we are closing this complaint. We will reopen this once the issue is resolved from the hardware side. | As of now, we are closing this complaint. We will reopen this once the issue is resolved from the hardware side. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN25001444471 | 2025-04-05 | 2025-04-06 | Cust 1350 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Transmitter Aquacon | Aquacon 2000(AQCN 2000 p) | Commissioning of Auacon2000 Transmitter | Commissioning has been done | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN25001444471 | 2025-04-05 | 2025-04-06 | Cust 1350 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Electrode | Forbes Marshall Cond. Cell Model 8310, K=0.01(FM 8310-1.5) | Commissioning of Conductivity cell | Commissioning has been done | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN25001444471 | 2025-04-05 | 2025-04-06 | Cust 1350 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Electrode | Liq-Glass PG(HT238 515) | Commissioning of PH sensor | Commissioning has been done | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN25001444471 | 2025-04-05 | 2025-04-06 | Cust 1350 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | pH Transmitter Aquamon | Ph Transmitter Aquamon 2000(PH2000) | Commissioning of PH transmitter Aquamon 2000 | Commissioning has been done | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN25001464472 | 2025-04-05 | 2025-04-11 | Cust 1350 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Electrode | ORP Electrode- Polilyte RX 120(1) | ORP sensor has been broken during commission . HT238523 PROCESS CONTROL INSTRUMENT SPARE,ORP ELECTRODE MODEL: EASYCONTROL ORP 120 TAG No. : IBR SERIAL NO. : IBR CERT. NO. : PRODUCT SERIAL NO : 20996 HSN/SAC : 90330000 | RF raised and successfully closed. | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN25001994473 | 2025-04-07 | 2025-04-09 | Cust 1308 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on MPCB server. | Issue from customer end which has been resolved and confirmed by site engineer over call Nikhil so we are closing this complaint. | Issue from customer end which has been resolved and confirmed by site engineer over call Nikhil so we are closing this complaint. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN25002334474 | 2025-04-07 | 2025-04-14 | Cust 1371 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Electrode | FM1104(1) | REPLACED PH ELECRODE | Suraj visited the site new pH sensor installed and commissioned successfully | Suraj visited the site new pH sensor installed and commissioned successfully | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFPN25003304475 | 2025-04-09 | 2025-04-09 | Cust 1236 | PUNE | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | Conductivity Electrode | Forbes Marshall Cell Model 8312, K=1(FM -8312) | Conductivity sensor testing 8312 | Checked and found RTD has been failed. | Closed | NO | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN25003304475 | 2025-04-09 | 2025-04-09 | Cust 1236 | PUNE | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Checking of Smartpro8966 wall mounted | Checked and electronics found ok but need to send at factory for calibration and components as well. | Closed | NO | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN25003304475 | 2025-04-09 | 2025-04-09 | Cust 1236 | PUNE | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Checking of Smartpro8966 wall mounted | Checked and electronics found ok but need to send at factory for calibration and components as well. | Closed | NO | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN25003304475 | 2025-04-09 | 2025-04-09 | Cust 1236 | PUNE | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Checking of Smartpro8966 wall mounted | Checked and electronics found ok but need to send at factory for calibration and components as well. | Closed | NO | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN25003304475 | 2025-04-09 | 2025-04-09 | Cust 1236 | PUNE | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | Conductivity transmitter testing and checkin | Checked and electronics found ok but need to send at factory for calibration and components as well. | Closed | NO | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN25003304475 | 2025-04-09 | 2025-04-09 | Cust 1236 | PUNE | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | Conductivity transmitter testing and checkin | Checked and electronics found ok but need to send at factory for calibration and components as well. | Closed | NO | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN25003304475 | 2025-04-09 | 2025-04-09 | Cust 1236 | PUNE | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | Conductivity transmitter testing and checkin | Checked and electronics found ok but need to send at factory for calibration and components as well. | Closed | NO | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN25003304475 | 2025-04-09 | 2025-04-09 | Cust 1236 | PUNE | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | Ph Tranmitter Smartpro | ORP transmitter(1) | calibration and testing is necessary | Checked and electronics found ok but need to send at factory for calibration and components as well. | Closed | NO | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN25003324476 | 2025-04-09 | 2025-04-11 | Cust 1295 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Flowmeter data is not live on CPCB+LOCAL server. | We took access to the system and updated the configuration from RS 485 to scanner and now data is live on CPCB+LOCAL server. | We took access to the system and updated the configuration from RS 485 to scanner and now data is live on CPCB+LOCAL server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN25003444477 | 2025-04-09 | 2025-04-12 | Cust 1275 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on CPCB server. | We are closing this complaint as of now as we are not getting full access to the system. | We are closing this complaint as of now as we are not getting full access to the system. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN25003604478 | 2025-04-09 | 2025-04-15 | Cust 936 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | Xenon Lamp Power supply for Multiparameter(AWA-XLAMPPS) | received mail from customer to attend the problem of the MP CX2000, and also received new Xenon power supply board. | visited the site new Xenon power supply board repaced, | visited the site new Xenon power supply board repaced, | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN25004324479 | 2025-04-11 | 2025-04-16 | Cust 1294 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Display blank issue | AMC visit done. Necessary maintenance carried out for Multiparameter. | AMC visit done. Necessary maintenance carried out for Multiparameter. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN25004444480 | 2025-04-11 | 2025-04-12 | Cust 640 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | New Flowmeter data not updating on CPCB+RTDMS portal | We took access to the system and completed the new flowmeter configuration after which data is live on CPCB+RTDMS server. | We took access to the system and completed the new flowmeter configuration after which data is live on CPCB+RTDMS server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN25005174481 | 2025-04-14 | 2025-04-21 | Cust 1313 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | TSS values not showing | work has been done | work has been done | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN25005324482 | 2025-04-14 | 2025-04-17 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Data not updating on MPCB server. | We took access to the system and now data is updating and live on MPCB server. | We took access to the system and now data is updating and live on MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN25005484483 | 2025-04-15 | 2025-04-17 | Cust 1260 | PUNE | Analytical - CHAKAN | CLOSED | DIGITAL CONEECTIVITY | SCADA Software | AIPL Software for CPCB(AIPL1) | Need to update customer details on Digital Dashboard. | We have updated the customer details on dashboard and informed same to the customer. | We have updated the customer details on dashboard and informed same to the customer. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN25005894484 | 2025-04-16 | 2025-04-17 | Cust 651 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software re-installation + CPCB and KSPCB connectivity | We have re-installed the software and completed Local, CPCB and KSPCB connectivity. Now the data is reflecting to the same servers. | We have re-installed the software and completed Local, CPCB and KSPCB connectivity. Now the data is reflecting to the same servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN25005934485 | 2025-04-16 | 2025-04-25 | Cust 1374 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Chlorine Analyzer-MAC 100 | MAC 100(1) | SR NO - 25A4TT090-W015000, 41712, 68644953M, Cl2 Analyzer Model : Chlorine Sensor:- FMSNX10 / pH Sensor:- POLILYTE PRO 120, Transmitter:- AQUA4TRANS 0 to 10 ppm, 4 to 20 mA With PH Compensate Make : Forbes Marshall NOS 1 1 423 | Invoice Notification of complaint no. FPN2500593 for HYDROPURE SYSTEMS PVT LTD. against Invoice No. :FC/RF552610106 Dated : 24-APR-25 , new sensor dispatched | visited the soite issue addressed need to repalce sensor . | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN25005944486 | 2025-04-16 | 2025-04-16 | Cust 1350 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | ORP transmitter(1) | Commissioning of ORP sensor | Commissioning has been done. | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN25005994487 | 2025-04-16 | 2025-04-19 | Cust 617 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on LOCAL+CPCB+MPCB server. | We took access to your system and now data is live on LOCAL+CPCB+MPCB server. | We took access to your system and now data is live on LOCAL+CPCB+MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN25006044488 | 2025-04-16 | 2025-04-18 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to CPCB, Local and MPCB server. | Now the data is live on the MPCB server. Also completed connectivity for Concentrated_Spentwash_flow and Distillery_Raw_SpentWash_Outlet_Flow. | Now the data is live on the MPCB server. Also completed connectivity for Concentrated_Spentwash_flow and Distillery_Raw_SpentWash_Outlet_Flow. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN25007494489 | 2025-04-21 | 2025-04-24 | Cust 1308 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | STACK Data not uploading on LOCAL+CPCB+MPCB server. | We took access to the system and now data is live on LOCAL+CPCB+MPCB server. | We took access to the system and now data is live on LOCAL+CPCB+MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN25007594490 | 2025-04-21 | 2025-04-28 | Cust 692 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | mA Output card for Multiparameter(AWA-4-20mA) | to check the healthiness of the MP and suggest for the spares. | visited the site checked healthyness of the meter and spare offer and AMC offer submited | visited the site checked healthyness of the meter and | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN25009564491 | 2025-04-25 | 2025-04-30 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | The daily Totalizer showing higher values than usual | Data is live on the server. We have checked on modscan and observed that the Totalizer readings are high. | Data is live on the server. We have checked on modscan and observed that the Totalizer readings are high. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN25009594492 | 2025-04-25 | 2025-04-28 | Cust 712 | PUNE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | AQUADO Analyzer | AQUAMAX DOT(1) | Calibration verification in Air | Calibration reports are attached. | Closed | NO | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN25009594492 | 2025-04-25 | 2025-04-28 | Cust 712 | PUNE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | AQUADO Analyzer | AQUAMAX DOT(1) | Calibration verification in Air | Calibration reports are attached. | Closed | NO | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN25011804493 | 2025-04-29 | 2025-05-04 | Cust 1302 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Plant shut down. | Calibration has done. Device working properly | AMC has done | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN25011974494 | 2025-04-29 | 2025-05-02 | Cust 1338 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on CPCB server. | Lagoon camera is live on CPCB server and for inlet camera there is issue from customer end. | Lagoon camera is live on CPCB server and for inlet camera there is issue from customer end. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN25012044495 | 2025-04-29 | 2025-05-19 | Cust 1375 | PUNE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | pH Transmitter Aquamon | Ph Transmitter Aquamon 2000(PH2000) | Material is installed at site but customer is facing problem in calibrating the pH meter and getting wrong reading. Site Address: HONOUR LAB LIMITED-UNIT-3B PLOT NO: A-88, MIDC, KURKUMBH, DAUND (TALUKA), PUNE (DIST), MAHARASTRA (ST), MAHARASHTRA 413802 Contact Person : Mr Srinivas Mobile No. : 8530219398 | Kartik visited the site pH issue resolved , conductivity sensor is damaged, customer to replacement. | Kartik visited the site pH issue resolved , conductivity sensor is damaged, customer to replacement. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN25015744496 | 2025-05-05 | 2025-05-06 | Cust 1284 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to MPCB server. | Data is now showing live on the MPCB server. | Data is now showing live on the MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN25016804497 | 2025-05-07 | 2025-05-09 | Cust 1275 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Inlet camera not playing on CPCB server. | We took access to the system and now both cameras is live on CPCB server. | We took access to the system and now both cameras is live on CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN25016944498 | 2025-05-07 | 2025-05-13 | Cust 517 | PUNE | Codel Products - CHAKAN | CLOSED | AMC | GCEM 40 Series New Gas Analyser | GCEM 40xx series DDU Only(GCEM 40xx DDU 4) | Amc of the device | Calibration carried out values are ok. Deice is in healthy condition. | calibration has done | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFPN25017034499 | 2025-05-07 | 2025-05-08 | Cust 712 | PUNE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | AQUADO Analyzer | Visiwater DO sensor(1) | calibration of sensor . | Air calibration has been performed on tank with ARC air software on laptop & Modbus convertor. But customer want to verify the same by doing zero in known concentration. So suggested to procure sodium bisulphate powder to verify zero. | Closed | NO | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN25017034499 | 2025-05-07 | 2025-05-08 | Cust 712 | PUNE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | AQUADO Analyzer | Visiwater DO sensor(1) | calibration of sensor. | Air calibration has been performed on tank with ARC air software on laptop & Modbus convertor. But customer want to verify the same by doing zero in known concentration. So suggested to procure sodium bisulphate powder to verify zero. | Closed | NO | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN25017354500 | 2025-05-08 | 2025-05-12 | Cust 1376 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Turbidity Meter | Optex turbidity sensor(1) | installation and commissioning of the Turbidity Analsyer | had detail discussion for the installation of the Turbidity chamber and sensor , powered on and commissioned, mail sent to customer is attached | had detail discussion for the installation of the Turbidity chamber and sensor , powered on and commissioned, mail sent to customer is attached | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFPN25018684501 | 2025-05-12 | 2025-05-13 | Cust 1039 | PUNE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Transmitter Aquacon | Conductivity Electrode Cable(CON-ELE-2000) | Conductivity values are Non-responding to the particular values. | Conductivity sensor found faulty need to replace | Conductivity sensor found faulty | NO | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN25018684501 | 2025-05-12 | 2025-05-13 | Cust 1039 | PUNE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | pH Transmitter Aquamon | Ph Transmitter Aquamon 2000(PH-Trans Aqu-2000) | pH values fluctuation. | Calibration has done found values ok | Conductivity sensor found faulty | NO | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN25021404502 | 2025-05-19 | 2025-05-21 | Cust 640 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data mismatch issue of Flow meter. | The data mismatch issue has been resolved and correct readings are reflecting on the CPCB server. | The data mismatch issue has been resolved and correct readings are reflecting on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN25021634503 | 2025-05-19 | 2025-05-23 | Cust 1243 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | found issue in blackbox | Blackbox found faulty | Blackbox found faulty | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFPN25021654504 | 2025-05-19 | 2025-05-23 | Cust 795 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | ph sensor faulty | Work done | Work done | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFPN25021734505 | 2025-05-19 | 2025-05-22 | Cust 1338 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Inlet and lagoon cameras not playing on CPCB server. | We took access to your system and now both Cameras are live on CPCB server. | We took access to your system and now both Cameras are live on CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN25022034506 | 2025-05-20 | 2025-05-21 | Cust 795 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to MPCB server. | data is live on the MPCB server. | data is live on the MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN25022874507 | 2025-05-21 | 2025-05-28 | Cust 571 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | ORP transmitter(1) | ORP not working | Need ORP not working | completed the work | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN25024404508 | 2025-05-25 | 2025-06-06 | Cust 1342 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | DO Transmitters | DO Sensor Cable(DO Sensor Cable) | received call from custoemr and tridev to to attend the problem of the Do transmitter , display getting OFF frequently. | Issue with the power supply need to change the FIT | Display getting off frequently | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFPN25024864509 | 2025-05-26 | 2025-05-29 | Cust 731 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not playing on CPCB software. | As of now we are closing this complaint. We will reopen this once the system is accessible. | As of now we are closing this complaint. We will reopen this once the system is accessible. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN25024904510 | 2025-05-26 | 2025-05-29 | Cust 1261 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB server. | As of now we are closing this complaint. We will reopen this once the system is accessible. | As of now we are closing this complaint. We will reopen this once the system is accessible. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN25025244511 | 2025-05-26 | 2025-05-30 | Cust 1301 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | received mail from Amol to visit the site for multiparameter inspection and healthiness checking | Analyser working fine. | Analyser working fine. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFPN25025434512 | 2025-05-27 | 2025-06-02 | Cust 803 | PUNE | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | AQUADO Analyzer | Visiferm DO sensor(1) | Online VC with Customer for doubts and queries | Suggested to increase the flow. And another option would be need to modify installation. | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN25026604513 | 2025-05-28 | 2025-05-31 | Cust 623 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue. | Not found any issue from the software side as the Data starts updating automatically on CPCB server. | Not found any issue from the software side as the Data starts updating automatically on CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN25026714514 | 2025-05-28 | 2025-06-02 | Cust 571 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | ORP transmitter(1) | ORP transmitter not in working | Reset and calibrated the meter | completed the work | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN25026954515 | 2025-05-29 | 2025-05-30 | Cust 1318 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data uploading issue at MPCB server. | As of now we are closing this complaint. We will reopen this once we get access to the system. | As of now we are closing this complaint. We will reopen this once we get access to the system. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN25027524516 | 2025-05-30 | 2025-05-31 | Cust 1377 | PUNE | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | General Panel | Panel Air Conditioning unit(AC Unit1) | received mail from customer to plan visit for installation guidelines of the hardness analyzer | visited the site all the material checked as per packing list, discussed requirement for erection of the hard ness analyser , customer will revert after completing the erection work. | visited the site all the material checked as per packing list, discussed requirement for erection of the hard ness analyser , customer will revert after completing the erection work. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN25028754517 | 2025-05-31 | 2025-06-04 | Cust 663 | PUNE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration of analyzer with standard solutions and DM water and Calibration of pH sensor | Multi-parameter analyser and flowmeter working OK. | Multi-parameter analyser and flowmeter working OK. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFPN25028754517 | 2025-05-31 | 2025-06-04 | Cust 663 | PUNE | Krohne - MIDC | CLOSED | CALIBRATION JOB | Magflow Meter Converter IFC series | Electronics Assembly of IFC 100(IFC 100Ele) | Calibration of Flowmeter using MS-01 Simulator | Multi-parameter analyser and flowmeter working OK. | Multi-parameter analyser and flowmeter working OK. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFPN25028754517 | 2025-05-31 | 2025-06-04 | Cust 663 | PUNE | Krohne - MIDC | CLOSED | CALIBRATION JOB | Magflow Meter Converter IFC series | Electronics Assembly of IFC 100(IFC 100Ele) | Calibration of flowmeter with MS-01 Simulator | Multi-parameter analyser and flowmeter working OK. | Multi-parameter analyser and flowmeter working OK. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFPN25029904518 | 2025-06-02 | 2025-06-05 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | data upload issue to MPCB server. | As of now we are closing this complaint. We will reopen this once the issue is resolved from the customer side. | As of now we are closing this complaint. We will reopen this once the issue is resolved from the customer side. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN25030034519 | 2025-06-02 | 2025-06-04 | Cust 663 | PUNE | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | no problem found | Analyser working Ok. | Analyser working Ok. | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFPN25030864520 | 2025-06-04 | 2025-06-06 | Cust 1276 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | MPCB site down issue. | As per your latest email, we understand that your industry process has been stopped since 13th March 2025, hence we request them to add the necessary workflow on the MPCB server to avoid these alerts. | As per your latest email, we understand that your industry process has been stopped since 13th March 2025, hence we request them to add the necessary workflow on the MPCB server to avoid these alerts. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN25030884521 | 2025-06-04 | 2025-06-06 | Cust 803 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data upload issue. | we took access of the system and now the time stamp issue has been resolved and data is live on the MPCB server. | we took access of the system and now the time stamp issue has been resolved and data is live on the MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN25031334522 | 2025-06-05 | 2025-06-09 | Cust 1378 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Electrode | Polilyte HT 120(HT238 431) | Commissioning of PH sensor with Aua4trans | Commissioned. | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN25031334522 | 2025-06-05 | 2025-06-09 | Cust 1378 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4 trans-A4TW033000(1) | Commissioning of transmitter | Commission done. | Closed | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFPN25031614523 | 2025-06-05 | 2025-06-10 | Cust 1379 | PUNE | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | AQUAMAX DOT(1) | Display get off repeatedly. | kartik visited the site DOT indicator replaced successfully , faulty carried back . | kartik visited the site DOT indicator repalced. succesfully | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN25032034524 | 2025-06-06 | 2025-06-10 | Cust 564 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Conductivity Electrode | Forbes Marshall Cell Model 8311, K=0.1(FM -8311) | received call from Amol Bhamare to attend the problem fo the conductivity transmitter not showing correct reading | Amit visited the site conductivity issue resolved successfully | Amit visited the site conductivity issue resolved successfully | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN25032074525 | 2025-06-06 | 2025-06-06 | Cust 1355 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | KSPCB station reports are not generating. | We taken up this issue with the KSPCB team and same has been resolved now. | We taken up this issue with the KSPCB team and same has been resolved now. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN25032084526 | 2025-06-06 | 2025-06-09 | Cust 1238 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to MPCB server. | As of now we are closing this complaint, we will reopen the same once the issue is resolved from the customer end. | As of now we are closing this complaint, we will reopen the same once the issue is resolved from the customer end. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN25032654527 | 2025-06-09 | 2025-06-13 | Cust 882 | PUNE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Electrode | Easycontrol 120(HT238 522) | pH not showing correct reading , | pH not showing correct reading ,Calibration done | completed the work | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFPN25033304528 | 2025-06-10 | 2025-06-13 | Cust 517 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB and MPCB data upload issue. | We took access of the system and observed that the software service was not in the running mode; hence, after restarting the same, the Data starts updating | We took access of the system and observed that the software service was not in the running mode; hence, after restarting the same, the Data starts updating | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN25034094529 | 2025-06-11 | 2025-06-13 | Cust 617 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | DATA not updating on CPCB server. | We took access to the system and now data is updating on CPCB server. | We took access to the system and now data is updating on CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN25034394530 | 2025-06-12 | 2025-06-16 | Cust 1239 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC | work has been done | work has been done | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFPN25034404531 | 2025-06-12 | 2025-06-16 | Cust 1260 | PUNE | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | mail from customer | customer check wrong data on our server, we inform him correct method to check live data as well as history data. | work has been done issue has been resolved | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFPN25034454532 | 2025-06-12 | 2025-06-16 | Cust 731 | PUNE | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not updating on CPCB server. | There is issue from customer end regarding as the Camera IP is not playing on google and VLC. | There is issue from customer end regarding as the Camera IP is not playing on google and VLC. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFPN25035884533 | 2025-06-16 |  | Cust 1243 | PUNE | Analytical - CHAKAN | ASSIGNED | DIGITAL CONEECTIVITY | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | online connectivity problem |  |  | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFPN25036444534 | 2025-06-17 |  | Cust 1239 | PUNE | Analytical - CHAKAN | ASSIGNED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Data Availability issue on MPCB server. |  |  | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFPN25036464535 | 2025-06-17 |  | Cust 1242 | PUNE | Analytical - CHAKAN | ASSIGNED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | calibration visit |  |  | NO | No Trip | Core | N |  | Analytical - CHAKAN
AFPP24155704536 | 2025-01-24 | 2025-01-30 | Cust 1380 | PHILIPINES | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Aqua4Trans | Aqua4Trans(1) | We have supplied 10 Nos of Aqua tran transmitter 2 nos have problem with the display request you to send 2 New Units | Material to be dispatch against RF 138255434 | Material to be dispatch against RF 138255434 | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFPP24155704536 | 2025-01-24 | 2025-01-30 | Cust 1380 | PHILIPINES | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Aqua4Trans | Aqua4Trans(1) | We have supplied 10 Nos of Aqua tran transmitter 2 nos have problem with the display request you to send 2 New Units | Material to be dispatch against RF 138255434 | Material to be dispatch against RF 138255434 | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFSB25011524537 | 2025-04-29 | 2025-04-30 | Cust 1381 | S ARABIA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | OIL ANALYSER | MUVX card(1) | Mux card and Ao card failure. health checks for ESS system. | Found Mux card and AO card failure. need to give offer for the spare | Found Mux card and AO card failure. need to give offer for the spare | NO | No Trip | Intops | N |  | Analytical - CHAKAN
AFSB25011544538 | 2025-04-29 | 2025-04-30 | Cust 1381 | S ARABIA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | ESS health check and OIL in water analyzer problems check. | Found Mux card and AO card failure. neoprene tube failure,,, need to give offer for the spare | Found Mux card and AO card failure. neoprene tube failure,,, need to give offer for the spare | NO | No Trip | Intops | N |  | Analytical - CHAKAN
AFSL23063034539 | 2023-07-27 | 2023-07-31 | Cust 1382 | SRI-LANKA | Analytical - CHAKAN | CLOSED | APPLICATION ISSUE | pH Transmitter Aquamon | Aquamon 2000 plus(AQ2000 plus) | PH meter not working need to visit plant to check it. | visited with ruwan and given an offer to customer | visited with ruwan and given an offer to customer | NO | No Trip | Intops | N |  | Analytical - CHAKAN
AFSL23083674540 | 2023-09-04 | 2023-09-08 | Cust 1383 | SRI-LANKA | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | MULTIPARAMETER ANALYSER | Water Detector for Multiparameter(AWA-WD) | Analyser Installation guide line | Installation guideline job Completed | Installation guideline job Completed | NO | No Trip | Intops | N |  | Analytical - CHAKAN
AFSL23086344541 | 2023-09-11 | 2023-09-15 | Cust 1383 | SRI-LANKA | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | LCD of Multiparameter Display(1) | Analyser Commissioning | Analyses Installation Guide line job Completed | Analyses Installation Guide line job Completed | NO | No Trip | Intops | N |  | Analytical - CHAKAN
AFSL23107734542 | 2023-10-23 | 2023-10-29 | Cust 1383 | SRI-LANKA | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MIL card of Multiparameter(1) | Analyser Commissioning | Site visit Completed | Site visit Completed | NO | No Trip | Intops | N |  | Analytical - CHAKAN
AFSL23125544543 | 2023-11-27 | 2023-12-03 | Cust 1383 | SRI-LANKA | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | 12 V POWER SUPPLY CARD for Multiparameter(AWA-PS1) | Commissioning | Commissioning job Completed | Commissioning job Completed | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFSL23172254544 | 2024-02-19 | 2024-02-25 | Cust 1384 | SRI-LANKA | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(Ph-Trans-Smartpro) | PH TX Commissioning | PH Controller Commissioning job Completed | PH Controller Commissioning job Completed | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFSL23178724545 | 2024-02-28 | 2024-03-03 | Cust 1383 | SRI-LANKA | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | MULTIPARAMETER ANALYSER | Sample Pump(1) | Analysser Installation guide line | Installation guide line job Completed | Installation guide line job Completed | NO | No Trip | Intops | N |  | Analytical - CHAKAN
AFSL23181864546 | 2024-03-04 | 2024-05-16 | Cust 1384 | SRI-LANKA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Aqua4Trans | Display card of Aqua4Trans(1) | Aqua4 TX Sensor faulty | Material dispatched against RF OA 5524601995, Fed Ex 7758 5461 3598 | Material dispatched against RF OA 5524601995, Fed Ex 7758 5461 3598 | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFSL24000684547 | 2024-04-02 | 2024-04-08 | Cust 1385 | SRI-LANKA | Analytical - CHAKAN | CLOSED | APPLICATION ISSUE | DO Transmitters | optex series(optex series) | Customer is complaining Some time display showing error need to visit site and check it | Sensor having an issue, needs to give offer for it. | Sensor having an issue, needs to give offer for it. | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFSL24017244548 | 2024-05-06 | 2024-05-09 | Cust 1384 | SRI-LANKA | Analytical - CHAKAN | CLOSED | APPLICATION ISSUE | Conductivity Tranmitter Polymetron | input card(IN-card-9125) | Conductivity Meter Repair | Repair job Completed | Repair job Completed | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFSL24085554549 | 2024-09-09 | 2024-09-16 | Cust 1386 | SRI-LANKA | Analytical - CHAKAN | CLOSED | APPLICATION ISSUE | pH Transmitter Aquamon | Ph Electrode Cable(PH-cab-2000) | Customer is complaining PH reading in incorrect | job completed | job completed | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFSL24171144550 | 2025-02-24 | 2025-03-04 | Cust 1387 | SRI-LANKA | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | MULTIPARAMETER ANALYSER | Current output card-FM(1) | Analyzer Spares Checking | Job Completed | Job Completed | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFSL25000044551 | 2025-04-01 | 2025-04-05 | Cust 1385 | SRI-LANKA | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | Commissioning new sensor for DO Transmitter | job completed | job completed | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFSL25016564552 | 2025-05-06 | 2025-05-13 | Cust 1388 | SRI-LANKA | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | PH TX Installation Guide line | Installation guideline job Completed | Installation guideline job Completed | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFSL25035544553 | 2025-06-16 |  | Cust 1388 | SRI-LANKA | Analytical - CHAKAN | ASSIGNED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(Ph-Trans-Smartpro) | PH Tx Commissioning |  |  | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFSP24053974554 | 2024-07-15 | 2024-07-15 | Cust 460 | SINGAPORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | CPU with LCD of Multiparameter Display(1) | The memory gets reset when the power is turned on. Priya has checked this unit, because the cell was not seated properly. Even after putting it right, the memory is getting reset. | Closed against RF OA 112460469 | Closed against RF OA 112460469 | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFSP24063414555 | 2024-07-30 | 2024-08-06 | Cust 460 | SINGAPORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | CPU Display card for Multiparameter(AWA-CPU1) | The received RF against 554240041 for EA330 has an issue, 'Display shows MPHDO in all the slots' So, kindly send another replacement against the faulty material. As we are not able to use this material. RF no. | Material dispatch against RF OA 112560232 | Material dispatch against RF OA 112560232 | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFSP24066684556 | 2024-08-05 | 2024-08-12 | Cust 460 | SINGAPORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | Xenon Lamp Power supply for Multiparameter(AWA-XLAMPPS) | Failure of Xenon power supply. 1. Not getting 1000V DC 2. Xenon power supply not trigerring | RF OA 112560242 booked , material dispatch against the same | RF OA 112560242 booked , material dispatch against the same | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFSP24066684556 | 2024-08-05 | 2024-08-12 | Cust 460 | SINGAPORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | Xenon Lamp Power supply for Multiparameter(AWA-XLAMPPS) | Failure of Xenon power supply. 1. Not getting 1000V DC 2. Xenon power supply not trigerring | RF OA 112560242 booked , material dispatch against the same | RF OA 112560242 booked , material dispatch against the same | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFSP24066684556 | 2024-08-05 | 2024-08-12 | Cust 460 | SINGAPORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | Xenon Lamp Power supply for Multiparameter(AWA-XLAMPPS) | Failure of Xenon power supply. 1. Not getting 1000V DC 2. Xenon power supply not trigerring | RF OA 112560242 booked , material dispatch against the same | RF OA 112560242 booked , material dispatch against the same | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFSP24066684556 | 2024-08-05 | 2024-08-12 | Cust 460 | SINGAPORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | Xenon Lamp Power supply for Multiparameter(AWA-XLAMPPS) | Failure of Xenon power supply. 1. Not getting 1000V DC 2. Xenon power supply not trigerring | RF OA 112560242 booked , material dispatch against the same | RF OA 112560242 booked , material dispatch against the same | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFSP24066684556 | 2024-08-05 | 2024-08-12 | Cust 460 | SINGAPORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | Xenon Lamp Power supply for Multiparameter(AWA-XLAMPPS) | Failure of Xenon power supply. 1. Not getting 1000V DC 2. Xenon power supply not trigerring | RF OA 112560242 booked , material dispatch against the same | RF OA 112560242 booked , material dispatch against the same | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFSP24066684556 | 2024-08-05 | 2024-08-12 | Cust 460 | SINGAPORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | Xenon Lamp Power supply for Multiparameter(AWA-XLAMPPS) | Failure of Xenon power supply. 1. Not getting 1000V DC 2. Xenon power supply not trigerring | RF OA 112560242 booked , material dispatch against the same | RF OA 112560242 booked , material dispatch against the same | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFSP24066684556 | 2024-08-05 | 2024-08-12 | Cust 460 | SINGAPORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | Xenon Lamp Power supply for Multiparameter(AWA-XLAMPPS) | Failure of Xenon power supply. 1. Not getting 1000V DC 2. Xenon power supply not trigerring | RF OA 112560242 booked , material dispatch against the same | RF OA 112560242 booked , material dispatch against the same | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFSP24066684556 | 2024-08-05 | 2024-08-12 | Cust 460 | SINGAPORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | Xenon Lamp Power supply for Multiparameter(AWA-XLAMPPS) | Failure of Xenon power supply. 1. Not getting 1000V DC 2. Xenon power supply not trigerring | RF OA 112560242 booked , material dispatch against the same | RF OA 112560242 booked , material dispatch against the same | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFSP24066684556 | 2024-08-05 | 2024-08-12 | Cust 460 | SINGAPORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | Xenon Lamp Power supply for Multiparameter(AWA-XLAMPPS) | Failure of Xenon power supply. 1. Not getting 1000V DC 2. Xenon power supply not trigerring | RF OA 112560242 booked , material dispatch against the same | RF OA 112560242 booked , material dispatch against the same | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFSP24067914557 | 2024-08-07 | 2024-12-03 | Cust 460 | SINGAPORE | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | CPU Display card for Multiparameter(AWA-CPU1) | Display blank Only backlit is on | Closed against 112560241 | Closed against 112560241 | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFSP24132734558 | 2024-12-11 |  | Cust 460 | SINGAPORE | Analytical - CHAKAN | ASSIGNED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | Xenon Lamp Power supply for Multiparameter(AWA-XLAMPPS) | Flash Faulty |  |  | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFSP24132734558 | 2024-12-11 |  | Cust 460 | SINGAPORE | Analytical - CHAKAN | ASSIGNED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | Xenon Lamp Power supply for Multiparameter(AWA-XLAMPPS) | Flash Faulty |  |  | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFSP24132734558 | 2024-12-11 |  | Cust 460 | SINGAPORE | Analytical - CHAKAN | ASSIGNED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | Xenon Lamp Power supply for Multiparameter(AWA-XLAMPPS) | Flash Faulty |  |  | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFSP24132734558 | 2024-12-11 |  | Cust 460 | SINGAPORE | Analytical - CHAKAN | ASSIGNED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | Xenon Lamp Power supply for Multiparameter(AWA-XLAMPPS) | Flash Faulty |  |  | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFSP24132734558 | 2024-12-11 |  | Cust 460 | SINGAPORE | Analytical - CHAKAN | ASSIGNED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | Xenon Lamp Power supply for Multiparameter(AWA-XLAMPPS) | Flash Faulty. |  |  | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFSP24132734558 | 2024-12-11 |  | Cust 460 | SINGAPORE | Analytical - CHAKAN | ASSIGNED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | PH2 detector(1) | PD2 count is below 100 |  |  | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFSP24132734558 | 2024-12-11 |  | Cust 460 | SINGAPORE | Analytical - CHAKAN | ASSIGNED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | PH1 detector(1) | The count dropping below 100 is considered very low |  |  | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFSP24132734558 | 2024-12-11 |  | Cust 460 | SINGAPORE | Analytical - CHAKAN | ASSIGNED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | PH1 detector(1) | The count dropping below 100 is considered very low |  |  | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFSP24132734558 | 2024-12-11 |  | Cust 460 | SINGAPORE | Analytical - CHAKAN | ASSIGNED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | PH2 detector(1) | The count dropping below 100 is considered very low |  |  | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFSP24132734558 | 2024-12-11 |  | Cust 460 | SINGAPORE | Analytical - CHAKAN | ASSIGNED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | PH2 detector(1) | The count dropping below 100 is considered very low |  |  | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFSP24132734558 | 2024-12-11 |  | Cust 460 | SINGAPORE | Analytical - CHAKAN | ASSIGNED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | Xenon Lamp Power supply for Multiparameter(AWA-XLAMPPS) | Zero Voltage |  |  | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFSP24132734558 | 2024-12-11 |  | Cust 460 | SINGAPORE | Analytical - CHAKAN | ASSIGNED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | Xenon Lamp Power supply for Multiparameter(AWA-XLAMPPS) | Zero Voltage |  |  | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFSP24132734558 | 2024-12-11 |  | Cust 460 | SINGAPORE | Analytical - CHAKAN | ASSIGNED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | Xenon Lamp Power supply for Multiparameter(AWA-XLAMPPS) | Zero Voltage |  |  | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFSP24132734558 | 2024-12-11 |  | Cust 460 | SINGAPORE | Analytical - CHAKAN | ASSIGNED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | Peristaltic sampling pump motor for Multiparameter(AWA-PSPM) | Sampling pump motor not working. |  |  | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFSP24132734558 | 2024-12-11 |  | Cust 460 | SINGAPORE | Analytical - CHAKAN | ASSIGNED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | Peristaltic sampling pump motor for Multiparameter(AWA-PSPM) | The motor is not operating |  |  | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFSP24132734558 | 2024-12-11 |  | Cust 460 | SINGAPORE | Analytical - CHAKAN | ASSIGNED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | Peristaltic sampling pump motor for Multiparameter(AWA-PSPM) | The motor is not operating |  |  | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFSP24132734558 | 2024-12-11 |  | Cust 460 | SINGAPORE | Analytical - CHAKAN | ASSIGNED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | Peristaltic sampling pump motor for Multiparameter(AWA-PSPM) | The motor is not operating |  |  | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFSP24132734558 | 2024-12-11 |  | Cust 460 | SINGAPORE | Analytical - CHAKAN | ASSIGNED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | Peristaltic sampling pump motor for Multiparameter(AWA-PSPM) | The motor is not operating |  |  | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFSP24132734558 | 2024-12-11 |  | Cust 460 | SINGAPORE | Analytical - CHAKAN | ASSIGNED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | Peristaltic sampling pump motor for Multiparameter(AWA-PSPM) | The motor is not operating |  |  | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFSP24132734558 | 2024-12-11 |  | Cust 460 | SINGAPORE | Analytical - CHAKAN | ASSIGNED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | Peristaltic sampling pump motor for Multiparameter(AWA-PSPM) | The motor is not operating |  |  | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFSP24132734558 | 2024-12-11 |  | Cust 460 | SINGAPORE | Analytical - CHAKAN | ASSIGNED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | Peristaltic sampling pump motor for Multiparameter(AWA-PSPM) | The motor is not operating |  |  | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFSP24132734558 | 2024-12-11 |  | Cust 460 | SINGAPORE | Analytical - CHAKAN | ASSIGNED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | Peristaltic sampling pump motor for Multiparameter(AWA-PSPM) | The motor is not operating |  |  | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFSP24132734558 | 2024-12-11 |  | Cust 460 | SINGAPORE | Analytical - CHAKAN | ASSIGNED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | Peristaltic sampling pump motor for Multiparameter(AWA-PSPM) | The motor is not operating |  |  | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFSR23033374559 | 2023-06-01 | 2023-06-05 | Cust 1389 | SURAT | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB and GPCB connectivity need to be done. | CPCB connectivity is completed and GPCB is pending as server is not accessible. | CPCB connectivity is completed and GPCB is pending as server is not accessible. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFSR23033774560 | 2023-06-02 | 2023-06-02 | Cust 1155 | SURAT | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | GPCB connectivity needs to be done. | Black Box configuration and CPCB and GPCB connectivity have been completed and data is live on the server. | Data is live on both servers hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFSR23033894561 | 2023-06-02 | 2023-06-04 | Cust 1155 | SURAT | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | For commissioning the FMBB & GPCB connectivity. | GPCB Configuration done successfully. | GPCB Configuration done successfully. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR23038424562 | 2023-06-12 | 2023-06-13 | Cust 1390 | SURAT | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | For calibration. | Analyzer working fine. | Analyzer working fine. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFSR23038434563 | 2023-06-12 | 2023-06-13 | Cust 1390 | SURAT | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | PH1 Signal is low | Analyzer working fine. | Analyzer working fine. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFSR23039284564 | 2023-06-13 | 2023-06-19 | Cust 1389 | SURAT | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC visit | work done | Work done | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFSR23039904565 | 2023-06-14 | 2023-06-15 | Cust 1155 | SURAT | Analytical - CHAKAN | CLOSED | COMMISSIONING | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | for commissioning | FMBB Commissioned successfully. | FMBB Commissioned successfully. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR23052254566 | 2023-07-07 | 2023-07-08 | Cust 1390 | SURAT | Analytical - CHAKAN | CLOSED | COMMISSIONING | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | For commissioning of FMBB | Commissioning successfully. | Commissioning successfully. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR23052264567 | 2023-07-07 | 2023-07-07 | Cust 1390 | SURAT | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB + GPCB configuration needs to be done. | as of now we are closing this complaint. | as of now we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFSR23055074568 | 2023-07-12 | 2023-07-14 | Cust 1389 | SURAT | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | For adding the TSS channel. | TSS Channel configure successfully. | TSS Channel configure successfully. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR23055094569 | 2023-07-12 | 2023-07-14 | Cust 1389 | SURAT | Analytical - CHAKAN | CLOSED | COMMISSIONING | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | For Commissioning the FMBB | Commissioning successfully. | Commissioning successfully. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR23056174570 | 2023-07-14 | 2023-07-15 | Cust 1155 | SURAT | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Flow meter data not transfer on GPCB server. | Output cable found faulty. | Output cable found faulty. | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR23059294571 | 2023-07-20 | 2023-07-21 | Cust 1390 | SURAT | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB and GPCB connectivity needs to be done. | we have completed the CPCB and GPCB connectivity through FMBB for Micas unit-1 and data is live on the CPCB server | Data is live hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFSR23061614572 | 2023-07-25 | 2023-07-28 | Cust 1390 | SURAT | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB + GPCB data upload issue. | now data is live on CPCB + GPCB server, hence we are closing this complaint. | now data is live on CPCB + GPCB server, hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFSR23067274573 | 2023-08-02 | 2023-08-07 | Cust 1390 | SURAT | Analytical - CHAKAN | CLOSED | DIGITAL CONEECTIVITY | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | For commisssioning | FMBB port found faulty | FMBB port found faulty | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR23074074574 | 2023-08-17 | 2023-08-21 | Cust 1391 | SURAT | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4Trans(1) | R& D work installation | Aqua 2trans, 1 transmitter not working | Work done | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR23074844575 | 2023-08-18 | 2023-08-19 | Cust 1392 | SURAT | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | For checking & calibration. | Analyzer working fine | Analyzer working fine. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFSR23080524576 | 2023-08-28 | 2023-08-30 | Cust 1391 | SURAT | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | For connectivity of GPCB & CPCB | device commission successfully. | device commission successfully. | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFSR23082314577 | 2023-08-31 | 2023-09-02 | Cust 1155 | SURAT | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | For GPCB & CPCB Connectivity. | Connectivity done. | Connectivity done. | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFSR23083174578 | 2023-09-02 | 2023-09-04 | Cust 1145 | SURAT | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | For installing new device | device commission successfully. | device commission successfully. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR23088654579 | 2023-09-14 | 2023-09-16 | Cust 1155 | SURAT | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | DATCON | DATCON(DATCON) | For GPCB & CPCB connectivity. | Datcon send to pune for repairing | Datcon send to pune for repairing | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR23089224580 | 2023-09-16 | 2023-09-18 | Cust 1155 | SURAT | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | For GPCB & CPCB connectivity. | IP not given by customer so connectivity pending. | IP not given by customer so connectivity pending. | NO | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR23089534581 | 2023-09-16 | 2023-09-18 | Cust 1393 | SURAT | Steam System Division | CLOSED | PRO-ACTIVE VISIT | CCDS | CCDS - TDS CONTROLLER - SMARTPRO(SMARTPRO - TDS) | REPLACEMENT OF FAULTY TDS SENSOR AND PWR CARD | REPLACEMENT DONE AND FOUND SYSTEM WORKING OK. | VISITED THE SITE AND RESOLVED THE ISSUE. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFSR23099414582 | 2023-10-04 | 2023-11-27 | Cust 1392 | SURAT | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Electrode | FM1106(1) | Electrode not working 2 nos. at the time of commissioning | Replacement material has been booked through COLLECT FROM FACTORY vide LR No - cff LR Dt - 20-OCT-2023 | Replacement material has been booked through COLLECT FROM FACTORY vide LR No - cff LR Dt - 20-OCT-2023 | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFSR23102904583 | 2023-10-11 | 2023-10-13 | Cust 1390 | SURAT | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB and GPCB connectivity issue. | The issue has been resolved and now the data is live on the server. | The issue has been resolved and now the data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFSR23102974584 | 2023-10-11 | 2023-10-12 | Cust 1145 | SURAT | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Black Box Commissioning | we have completed the Black Box configuration + CPCB and GPCB connectivity and now the data is live on both sites. | Data is live hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFSR23104744585 | 2023-10-16 | 2023-10-20 | Cust 1145 | SURAT | Analytical - CHAKAN | CLOSED | COMMISSIONING | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | For FMBB configuration and GPCB & CPCB connectivity | FMBB configuration and GPCB & CPCB connectivity | FMBB configuration and GPCB & CPCB connectivity | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFSR23106384586 | 2023-10-18 | 2023-10-20 | Cust 1153 | SURAT | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB and GPCB data uploaded issue. | we took access of the BB and updated it to the latest version 1.9, Data is live on the CPCB + GPCB Server. | we took access of the BB and updated it to the latest version 1.9, Data is live on the CPCB + GPCB Server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFSR23106884587 | 2023-10-19 | 2023-10-25 | Cust 1392 | SURAT | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | CPU Display card for Multiparameter(AWA-CPU1) | Display not getting power on after giving the separate 12V supply. | Display found faulty. | Display found faulty. | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFSR23107304588 | 2023-10-20 | 2023-10-25 | Cust 1155 | SURAT | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | For checking and power on the analyser | PH2 filter found faulty. | PH2 filter found faulty. | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFSR23107314589 | 2023-10-20 | 2023-10-25 | Cust 1155 | SURAT | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | For checking the multiparameter and power on the analyzer | Analyzer working OK. | Analyzer working OK. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFSR23107324590 | 2023-10-20 | 2023-10-25 | Cust 1394 | SURAT | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | For checking and power on the analyser | Analyzer working OK. | Analyzer working OK. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFSR23107384591 | 2023-10-20 | 2023-10-25 | Cust 1155 | SURAT | Analytical - CHAKAN | CLOSED | AMC | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | For FMBB configuration, GPCB & CPCB Connectivity | Connectivity done. | Connectivity done. | NO | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR23107424592 | 2023-10-20 | 2023-10-25 | Cust 1155 | SURAT | Analytical - CHAKAN | CLOSED | AMC | DATCON | DATCON(DATCON) | DATCON (GPCB & CPCB connectivity of spm) | Connectivity done. | Rf done for Display. | NO | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR23107434593 | 2023-10-20 | 2023-10-20 | Cust 1145 | SURAT | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB and GPCB data uploaded issue | Now the data is live on the CPCB + GPCB server please check once from your end. | Now the data is live on the CPCB + GPCB server please check once from your end. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFSR23108664594 | 2023-10-25 | 2023-10-28 | Cust 1392 | SURAT | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | To Change the display of analyzer and start analyser | To change the display of analyser and start analyser | FMBB configuration and GPCB & CPCB connectivity | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFSR23110494595 | 2023-10-28 | 2023-10-28 | Cust 1390 | SURAT | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB and GPCB data uploaded issue. | we have completed the CPCB and GPCB connectivity for unit 2 through FMBB and now the data is live on both sites. | we have completed the CPCB and GPCB connectivity for unit 2 through FMBB and now the data is live on both sites. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFSR23111874596 | 2023-10-30 | 2023-11-05 | Cust 1390 | SURAT | Analytical - CHAKAN | CLOSED | COMMISSIONING | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | M/S FMPL Engineer visited the site for commissioning of FM Black Box | connectivity done | connectivity done | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR23113294597 | 2023-10-31 | 2023-11-02 | Cust 1145 | SURAT | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB and GPCB data uploaded issue. | we took access of the system and made some changes in the configuration so now the data is live on the server. | we took access of the system and made some changes in the configuration so now the data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFSR23116034598 | 2023-11-05 | 2023-11-05 | Cust 1155 | SURAT | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | 12 V POWER SUPPLY CARD for Multiparameter(AWA-PS1) | power supply card found not ok | issue is pending to be resolved | issue is pending to be resolved | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR23116544599 | 2023-11-06 | 2023-12-12 | Cust 1395 | SURAT | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB and GPCB data uploaded issue. | Data is live on the CPCB and For GPCB we are checking this issue iwth GPCB so as of now we are closing this complaint. | Data is live on the CPCB and For GPCB we are checking this issue iwth GPCB so as of now we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFSR23116994600 | 2023-11-07 | 2023-11-10 | Cust 1153 | SURAT | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB and GPCB data uploaded issue. | Issue has been resolved and data is live on the server. | Issue has been resolved and data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFSR23117144601 | 2023-11-07 | 2023-11-09 | Cust 1155 | SURAT | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB AND GPCB data uploded issue. | the issue has been resolved and data is live on both sites for both the units. | the issue has been resolved and data is live on both sites for both the units. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFSR23117274602 | 2023-11-07 | 2023-11-13 | Cust 1155 | SURAT | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | ph filter is faulty ,ph electrode faulty | calibration done | calibration done | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR23117934603 | 2023-11-08 | 2023-11-13 | Cust 1392 | SURAT | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | CPU Display card for Multiparameter(AWA-CPU1) | problem solve | configuration and alignment done ok | troubleshooting solve | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR23117944604 | 2023-11-08 | 2023-11-13 | Cust 1155 | SURAT | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | configuration and connectivity done ok | all done | configuration and connectivity done ok | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFSR23120744605 | 2023-11-17 | 2023-11-18 | Cust 1155 | SURAT | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | Xenon Lamp Power supply for Multiparameter(AWA-XLAMPPS) | CHANGE POWER SUPPLY CARD | change card and done with configuration | change card and done with configuration | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFSR23120784606 | 2023-11-17 | 2023-11-18 | Cust 1155 | SURAT | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | CPU with LCD of Multiparameter Display(1) | DISPLAY SHOWING KEYPAD FAILED | issue is pending | issue is pending to be resolved | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFSR23120794607 | 2023-11-17 | 2023-11-18 | Cust 1155 | SURAT | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | CALIBRATION DONE | calibration and connectivity done | configuration and connectivity done ok | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFSR23123494608 | 2023-11-23 | 2023-11-23 | Cust 1394 | SURAT | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Black Box configuration | we completed the BB configuration and now the Stack data is live on both sites. | we completed the BB configuration and now the Stack data is live on both sites. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFSR23124654609 | 2023-11-25 | 2023-11-27 | Cust 1394 | SURAT | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB and GPCB connectivity needs to be done. | Now ETP data is live on both the sites. | Now ETP data is live on both the sites. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFSR23124824610 | 2023-11-26 | 2023-11-28 | Cust 1391 | SURAT | Analytical - CHAKAN | CLOSED | DIGITAL CONEECTIVITY | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CONNECTIVIY NOT DONE 1 | Connectivity not done due to scanner configuration remain from customer side | connectivity not done | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR23127114611 | 2023-11-28 | 2023-12-01 | Cust 1389 | SURAT | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB and GPCB data uploaded issue. | Data is now live on the server. | Data is now live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFSR23130044612 | 2023-12-01 | 2023-12-04 | Cust 1390 | SURAT | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Please to be check GPCB& CPCB communication for Online Monitoring System of Micas Organics Ltd. Vapi. Unit-I | Data is live hence we are closing this complaint. | Data is live hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFSR23133184613 | 2023-12-08 | 2023-12-11 | Cust 1155 | SURAT | Analytical - CHAKAN | CLOSED | DIGITAL CONEECTIVITY | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | connectivity not done | found EA 370 card faulty and connectivity not done. | FMBB configuration and GPCB & CPCB connectivity not done | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFSR23133414614 | 2023-12-09 | 2023-12-11 | Cust 1155 | SURAT | Analytical - CHAKAN | CLOSED | DIGITAL CONEECTIVITY | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | connectivity not done due to network issue | Display Problem in multiparameter ,req customer to change display | FMBB configuration and GPCB & CPCB connectivity not done | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFSR23133424615 | 2023-12-09 | 2023-12-11 | Cust 1155 | SURAT | Analytical - CHAKAN | CLOSED | DIGITAL CONEECTIVITY | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | connectivity done | FMBB Connectivity done. | FMBB configuration and GPCB & CPCB connectivity done | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFSR23134604616 | 2023-12-12 | 2023-12-15 | Cust 1155 | SURAT | Analytical - CHAKAN | CLOSED | AMC | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | data check on modscan found ok from ddu ,but from cable data not found | FMBB Connectivity not done | FMBB configuration and GPCB & CPCB connectivity not done | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFSR23135124617 | 2023-12-13 | 2023-12-15 | Cust 1155 | SURAT | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | PH1 detector(1) | ph1 signal not come | request customer to change ph1 filter | PH! signal not come | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR23136434618 | 2023-12-15 | 2023-12-15 | Cust 1155 | SURAT | Analytical - CHAKAN | CLOSED | COMMISSIONING | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | FMBB Connectivity done and multiparameter display also change | due to not proper supply amc of spm is not done. | FMBB configuration and GPCB & CPCB connectivity done | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFSR23139594619 | 2023-12-21 | 2023-12-23 | Cust 1155 | SURAT | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | display problem | due to fault in cable connectivtiy data not come req coustmer to change cable from stack to blackbox. change the display now parameter running ok | due to fault in cable connectivtiy data not come req coustmer to change cable from stack to blackbox. change the display now parameter running okj | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFSR23139854620 | 2023-12-21 | 2023-12-27 | Cust 1396 | SURAT | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC visit | Checked the healthiness of the analyzer and found MPHDO card faulty, Recommended for spare. | Checked the healthiness of the analyzer and found MPHDO card faulty, Recommended for spare. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFSR23141504621 | 2023-12-25 | 2023-12-26 | Cust 1155 | SURAT | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | PH2 Filter(1) | change ph2 filter | change ph2 filter now reading is ok | change ph2 filter | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFSR23143794622 | 2023-12-28 | 2023-12-29 | Cust 1155 | SURAT | Analytical - CHAKAN | CLOSED | DIGITAL CONEECTIVITY | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | DATA NOT LIVE ON GPCB | Data is live now | NOW data is live | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFSR23143794622 | 2023-12-28 | 2023-12-29 | Cust 1155 | SURAT | Analytical - CHAKAN | CLOSED | DIGITAL CONEECTIVITY | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | DATA NOT LIVE ON GPCB | now data is luve | NOW data is live | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFSR23143814623 | 2023-12-28 | 2023-12-30 | Cust 1397 | SURAT | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | GPCB data uploaded issue. | Data is live on the GPCB server for both the units. | Data is live on the GPCB server for both the units. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFSR23147784624 | 2024-01-03 | 2024-01-09 | Cust 1392 | SURAT | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC visit against PO: 4700269790 | Checked the healthiness of the analyzer and carried out the calibration, found working ok. | Checked the healthiness of the analyzer and carried out the calibration, found working ok. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFSR23150274625 | 2024-01-08 | 2024-03-25 | Cust 610 | SURAT | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Electrode Cable(Ph-Cab) | Issue : As per MOM from Mr.Nayan Rana, Surat branch. Below new pH analyser is faulty and needs to be replaced. | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210091281 LR Dt - 09-MAR-2024 | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210091281 LR Dt - 09-MAR-2024 | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFSR23151484626 | 2024-01-10 | 2024-01-12 | Cust 1153 | SURAT | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera configuration needs to be done. | we took access to your system and did camera configuration, and now the camera is streaming live on CPCB server for Bio Compost & Spent Wash. | we took access to your system and did camera configuration, and now the camera is streaming live on CPCB server for Bio Compost &amp;amp; Spent Wash. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFSR23156374627 | 2024-01-19 | 2024-01-19 | Cust 1397 | SURAT | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera configuration needs to be done. | We took the access to the system and did Camera configuration for Spent wash Outlet, and now camera streaming live on CPCB server. | We took the access to the system and did Camera configuration for Spent wash Outlet, and now camera streaming live on CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFSR23156594628 | 2024-01-20 | 2024-01-22 | Cust 1155 | SURAT | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | DATA NOT UPLOAD ON GPCB/CPCB | NOW DATA IS LIVE | now data is live | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR23158414629 | 2024-01-23 | 2024-01-29 | Cust 1155 | SURAT | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | data not come ,and tp change flow disk | problem solve | problem solve | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR23163594630 | 2024-01-31 | 2024-02-02 | Cust 1155 | SURAT | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | BUBBLE COMMING IN FLOW CELL THATS WHY READING COMING IN ZER0 | TROUBLESHOOTING DONE | TROUBLESHOOTING DONE | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR23166884631 | 2024-02-07 | 2024-02-09 | Cust 1396 | SURAT | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | pH input card-FM of Multiparameter(1) | visited site for new ph card and ph electrode sensor | open the box of both found physically okay, mount it done wiring ,done calibration all found okay now | open the box of both found physically okay, mount it done wiring ,done calibration all found okay now | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR23176084632 | 2024-02-25 | 2024-02-28 | Cust 1145 | SURAT | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | commissioning | commissioning done ok | commissioning done ok | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR23180134633 | 2024-02-29 | 2024-03-05 | Cust 1390 | SURAT | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | Water Detector for Multiparameter(AWA-WD) | AMC VISIT | AMC DONE FOUND XENON LAMP ASSEMBLY BROKEN DUE TO THAT TSS READING COMING ZERO SO REQ CUSTOMER TO ORDER NEW ASSEMBLY | AMC DONE FOUND XENON LAMP ASSEMBLY BROKEN DUE TO THAT TSS READING COMING ZERO SO REQ CUSTOMER TO ORDER NEW ASSEMBLY | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFSR23180144634 | 2024-02-29 | 2024-03-05 | Cust 1390 | SURAT | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC VISIIT | done amc | AMC DONE | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFSR23182564635 | 2024-03-06 | 2024-03-09 | Cust 1398 | SURAT | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | FOR INSTALLATION GUIDELINES | gave installation guidelines | gave installation guidelines | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR23182574636 | 2024-03-06 | 2024-03-09 | Cust 1155 | SURAT | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | DATA ISSUE | now data is live | now data is live | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR23184114637 | 2024-03-11 | 2024-03-14 | Cust 1155 | SURAT | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC Visit | AMC DONE | AMC DONE | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFSR23185594638 | 2024-03-14 | 2024-03-20 | Cust 1145 | SURAT | Analytical - CHAKAN | CLOSED | DIGITAL CONEECTIVITY | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | for connectivity | connectivity done | done connectivity | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR23185614639 | 2024-03-14 | 2024-03-17 | Cust 1155 | SURAT | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | DISPLAY RESET AND FOR AMC | FOUND DISPLAY RESET ,DONE CALIBRATION ,DONE CONFIGURATION DONE AMC | DONE AMC | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFSR23188094640 | 2024-03-18 | 2024-03-20 | Cust 1392 | SURAT | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | VISIT FOR AMC | AMC DONE | AMC DONE | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFSR23190194641 | 2024-03-20 | 2024-03-26 | Cust 1399 | SURAT | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | some works are pending | now all pending work done by m/s fmpl online work ,cod readinfg and error | now all pending work done by m/s fmpl online work ,cod readinfg and error | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR24001294642 | 2024-04-03 | 2024-04-04 | Cust 1391 | SURAT | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Aqua4Trans | Aqua4Trans(1) | Visited site for to check transmitter and sensor | check transmitter found ok but pH electrode found faulty | check transmitter found ok but pH electrode found faulty | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR24001294642 | 2024-04-03 | 2024-04-04 | Cust 1391 | SURAT | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Aqua4Trans | Aqua4Trans(1) | Visited site for to check transmitter and sensor | check transmitter found ok but pH electrode found faulty so req customer to purchase new ph electrode | check transmitter found ok but pH electrode found faulty | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR24006624643 | 2024-04-15 | 2024-04-22 | Cust 1400 | SURAT | Steam System Division | CLOSED | QUALITY PROBLEM | PRV - MEASURETROL | FM MEASURETROL(SIZE 080NB) | P1 AND P2 TRANSMITTER FAILURE ISSUE. | The spare PCB/ELE were sent to Surat Office by Abhijeet Deshpande. The failed PCB/ELE to be sent to Kasarwadi for RCA. | The spare PCB/ELE were sent to Surat Office by Abhijeet Deshpande. The failed PCB/ELE to be sent to Kasarwadi for RCA. | YES | No Trip | Process | N | Product / equipment failure due to which customers plant or part of plant is stopped which leaads to production loss. | Analytical - CHAKAN
AFSR24006624643 | 2024-04-15 | 2024-04-22 | Cust 1400 | SURAT | Steam System Division | CLOSED | QUALITY PROBLEM | PRV - MEASURETROL | FM MEASURETROL(SIZE 080NB) | P1 AND P2 TRANSMITTER FAILURE ISSUE. | The spare PCB/ELE were sent to Surat Office by Abhijeet Deshpande. The failed PCB/ELE to be sent to Kasarwadi for RCA. | The spare PCB/ELE were sent to Surat Office by Abhijeet Deshpande. The failed PCB/ELE to be sent to Kasarwadi for RCA. | YES | No Trip | Process | N | Product / equipment failure due to which customers plant or part of plant is stopped which leaads to production loss. | Analytical - CHAKAN
AFSR24006624643 | 2024-04-15 | 2024-04-22 | Cust 1400 | SURAT | Steam System Division | CLOSED | QUALITY PROBLEM | PRV - MEASURETROL | FM MEASURETROL(SIZE 080NB) | P1 AND P2 TRANSMITTER FAILURE ISSUE. | The spare PCB/ELE were sent to Surat Office by Abhijeet Deshpande. The failed PCB/ELE to be sent to Kasarwadi for RCA. | The spare PCB/ELE were sent to Surat Office by Abhijeet Deshpande. The failed PCB/ELE to be sent to Kasarwadi for RCA. | YES | No Trip | Process | N | Product / equipment failure due to which customers plant or part of plant is stopped which leaads to production loss. | Analytical - CHAKAN
AFSR24006624643 | 2024-04-15 | 2024-04-22 | Cust 1400 | SURAT | Steam System Division | CLOSED | QUALITY PROBLEM | PRV - MEASURETROL | FM MEASURETROL(SIZE 080NB) | P1 AND P2 TRANSMITTER FAILURE ISSUE. | The spare PCB/ELE were sent to Surat Office by Abhijeet Deshpande. The failed PCB/ELE to be sent to Kasarwadi for RCA. | The spare PCB/ELE were sent to Surat Office by Abhijeet Deshpande. The failed PCB/ELE to be sent to Kasarwadi for RCA. | YES | No Trip | Process | N | Product / equipment failure due to which customers plant or part of plant is stopped which leaads to production loss. | Analytical - CHAKAN
AFSR24007454644 | 2024-04-17 | 2024-04-22 | Cust 1401 | SURAT | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | we have supplied pH transmitters Smartpro 8966 2 Nos., 1 nos is not getting power ON. | Checked and found that it seems power supply related issue, hence suggested to send the faulty unit to pune for detail checking. | Checked and found that it seems power supply related issue, hence suggested to send the faulty unit to pune for detail checking. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR24013834645 | 2024-04-29 | 2024-07-01 | Cust 1402 | SURAT | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Electrode | Polilyte HT 120(HT238 431) | Customer purchased the new electrode in Feb 2024, pH electrode is not responding with any buffer solution. | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210100846 LR Dt - 15-MAY-2024 | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210100846 LR Dt - 15-MAY-2024 | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFSR24016204646 | 2024-05-02 | 2024-05-06 | Cust 1389 | SURAT | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | visit for amc | amc done | AMC DONE | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFSR24022554647 | 2024-05-19 | 2024-05-27 | Cust 610 | SURAT | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Ph system is not working at AMNS Surat site. Avinash Malik SSD Goa Ion Exchange (India) Ltd. A: Plot No. L-48 & L-49, Verna Industrial Area, Phase II, Verna, Salcete, Goa -403 722, India. B: 0832 6693000 D: 09403355955 | check RO common feed inlet, RO-!A permeapes and RO -2B Permeapes , Carried out calibration with buffer and pH simulator found working ok found RO common (RO_1A&B cf (s/n - 23SIRBD021) electrode sensor faulty due to dry condition | check RO common feed inlet, RO-!A permeapes and RO -2B Permeapes , Carried out calibration with buffer and pH simulator found working ok | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR24025704648 | 2024-05-24 | 2024-05-28 | Cust 1145 | SURAT | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | data is not live and to check analyzer | DATA IS LIVE AND ALSO AMC DONE | DATA IS LIVEAND AMC DONE | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFSR24028814649 | 2024-05-29 | 2024-05-31 | Cust 1391 | SURAT | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Data mismatch issue on CPCB server. | Issue has been resolved from the customer end. | Issue has been resolved from the customer end. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFSR24031514650 | 2024-06-01 | 2024-06-03 | Cust 1155 | SURAT | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | VISIT SITE FOR TO CHANGE THE DISPLAY AND FOR DAILY ROUTINE | done configuration of display clean the flowcell by acid and dm water but found EA270 card faulty (found 240 vdc) | Configuration done but found EA370 card faulty | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR24032404651 | 2024-06-04 | 2024-07-16 | Cust 1390 | SURAT | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB and GPCB data uploaded issue. | FMBB Checked at Our Pune works and carried out the necessary exercise, Now working fine. | FMBB Checked at Our Pune works and carried out the necessary exercise, Now working fine. | YES | Non Warranty (Free) | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFSR24035754652 | 2024-06-11 | 2024-06-13 | Cust 845 | SURAT | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | AQUAMAX DOT(1) | Commissioning visit required for 4 nos. at SMC | Commissioning has been done and working satisfactorily. | Commissioning has been done and working satisfactorily. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR24035754652 | 2024-06-11 | 2024-06-13 | Cust 845 | SURAT | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | AQUAMAX DOT(1) | Commissioning visit required for 4 nos. at SMC | Commissioning has been done and working satisfactorily. | Commissioning has been done and working satisfactorily. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR24035754652 | 2024-06-11 | 2024-06-13 | Cust 845 | SURAT | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | AQUAMAX DOT(1) | Commissioning visit required for 4 nos. at SMC | Commissioning has been done and working satisfactorily. | Commissioning has been done and working satisfactorily. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR24035754652 | 2024-06-11 | 2024-06-13 | Cust 845 | SURAT | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | AQUAMAX DOT(1) | Commissioning visit required for 4 nos. at SMC | ODOCAP is missing, Hence suggested customer for procurement. | Commissioning has been done and working satisfactorily. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR24036874653 | 2024-06-13 | 2024-06-14 | Cust 1145 | SURAT | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data upload issue. | as of now we are closing this complaint and will reopen it once we get PDC for Multiparameter and BB access for flow . | as of now we are closing this complaint and will reopen it once we get PDC for Multiparameter and BB access for flow . | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFSR24037964654 | 2024-06-17 | 2024-06-20 | Cust 1155 | SURAT | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB connectivity needs to be done. | We took access of the black box, and completed CPCB connectivity for CETP OUTLET (through RS485), and now data is live on CPCB + GPCB server. | We took access of the black box, and completed CPCB connectivity for CETP OUTLET (through RS485), and now data is live on CPCB + GPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFSR24038224655 | 2024-06-17 | 2024-06-18 | Cust 845 | SURAT | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4 trans-A4TW033000(1) | commisioing | Commissioning is already done on 11th June and duplicate complaint logged by the Pune division. | Commissioning is already done on 11th June and duplicate complaint logged by the Pune division. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR24039664656 | 2024-06-19 | 2024-06-24 | Cust 1155 | SURAT | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | VISIT SITE FOR TO INSTALLED SCANNER AND FOR REMAINING CONNECTIVITY | install scanner and done with connectivity | Install scanner and done with connectivity | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR24041754657 | 2024-06-24 | 2024-06-25 | Cust 1145 | SURAT | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB + GPCB data upload issue. | we took access of the black box and completed CPCB + GPCB connectivity for ETP Outlet to CETP (Flow). And now data is live on CPCB + GPCB server | we took access of the black box and completed CPCB + GPCB connectivity for ETP Outlet to CETP (Flow). And now data is live on CPCB + GPCB server | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFSR24044764658 | 2024-06-29 | 2024-07-03 | Cust 1392 | SURAT | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | visit site for amc | AMC DONE | AMC DONE | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFSR24050964659 | 2024-07-09 | 2024-07-10 | Cust 1145 | SURAT | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB + GPCB data upload issue. | we took access of the BB and made some modifications, now data is live on the CPCB + GPCB server. | we took access of the BB and made some modifications, now data is live on the CPCB + GPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFSR24051594660 | 2024-07-10 | 2024-07-12 | Cust 1145 | SURAT | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | GPCB connectivity needs to be done. | we took access of the black box and completed GPCB connectivity for MPL device now data is live on the GPCB server. | we took access of the black box and completed GPCB connectivity for MPL device now data is live on the GPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFSR24051964661 | 2024-07-11 | 2024-10-29 | Cust 1403 | SURAT | Swas - CHAKAN | CLOSED | PRODUCT FAILURE | CCDS | CCDS - CONDENSATE CONTAMINATION DETECTION SYSTEM - PH+CONDUCTIVITY BASE(CCDS-PC) | FOUND SEAL LEKAGE OF OIL AND CONDENSATE IN ALL THE PUMP PROVIDED FOR CCDS SYSTEM. | Pump 1 to 3 gasket already replaced. Pump 4 gasket sent through shree maruti transport 24033200291666 | Pump 1 to 3 gasket already replaced. Pump 4 gasket sent through shree maruti transport 24033200291666 | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFSR24051964661 | 2024-07-11 | 2024-10-29 | Cust 1403 | SURAT | Swas - CHAKAN | CLOSED | PRODUCT FAILURE | CCDS | CCDS - CONDENSATE CONTAMINATION DETECTION SYSTEM - PH+CONDUCTIVITY BASE(CCDS-PC) | FOUND SEAL LEKAGE OF OIL AND CONDENSATE IN ALL THE PUMP PROVIDED FOR CCDS SYSTEM. | Pump 1 to 3 gasket already replaced. Pump 4 gasket sent through shree maruti transport 24033200291666 | Pump 1 to 3 gasket already replaced. Pump 4 gasket sent through shree maruti transport 24033200291666 | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFSR24051964661 | 2024-07-11 | 2024-10-29 | Cust 1403 | SURAT | Swas - CHAKAN | CLOSED | PRODUCT FAILURE | CCDS | CCDS - CONDENSATE CONTAMINATION DETECTION SYSTEM - PH+CONDUCTIVITY BASE(CCDS-PC) | FOUND SEAL LEKAGE OF OIL AND CONDENSATE IN ALL THE PUMP PROVIDED FOR CCDS SYSTEM. | Pump 1 to 3 gasket already replaced. Pump 4 gasket sent through shree maruti transport 24033200291666 | Pump 1 to 3 gasket already replaced. Pump 4 gasket sent through shree maruti transport 24033200291666 | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFSR24051964661 | 2024-07-11 | 2024-10-29 | Cust 1403 | SURAT | Swas - CHAKAN | CLOSED | PRODUCT FAILURE | CCDS | CCDS - CONDENSATE CONTAMINATION DETECTION SYSTEM - PH+CONDUCTIVITY BASE(CCDS-PC) | FOUND SEAL LEKAGE OF OIL AND CONDENSATE IN ALL THE PUMP PROVIDED FOR CCDS SYSTEM. | Pump 1 to 3 gasket already replaced. Pump 4 gasket sent through shree maruti transport 24033200291666 | Pump 1 to 3 gasket already replaced. Pump 4 gasket sent through shree maruti transport 24033200291666 | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFSR24053984662 | 2024-07-15 | 2024-07-19 | Cust 487 | SURAT | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera Streaming issue on the CPCB server. | The camera streaming is showing live on the CPCB server. | The camera streaming is showing live on the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFSR24054964663 | 2024-07-17 | 2024-07-22 | Cust 1389 | SURAT | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 2nd AMC visit against PO 5600016985 | AMC DONE | AMC DONE | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFSR24055564664 | 2024-07-18 | 2024-07-19 | Cust 1392 | SURAT | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data uploaded issue. | now the data is live on the server. | now the data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFSR24055604665 | 2024-07-18 | 2024-07-20 | Cust 1391 | SURAT | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data uploaded issue. | As of now we are closing this complaint and will reopen this case once we receive remote access of your system from the customer end. | As of now we are closing this complaint and will reopen this case once we receive remote access of your system from the customer end. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFSR24055904666 | 2024-07-18 | 2024-07-24 | Cust 1396 | SURAT | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | VISIT SITE FOR AMC | AMC DONE | AMC DONE | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFSR24057164667 | 2024-07-22 | 2024-07-24 | Cust 1155 | SURAT | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | data found close on server | NOW DATA IS LIVE | now data is live | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR24058714668 | 2024-07-24 | 2024-07-26 | Cust 1145 | SURAT | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB + GPCB data upload issue. | today we took access of the black box, and now the internet side issue is resolved and the internet is coming till the cpcb servar. | today we took access of the black box, and now the internet side issue is resolved and the internet is coming till the cpcb servar. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFSR24065134669 | 2024-08-01 | 2024-08-05 | Cust 610 | SURAT | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Smartpro-con) | Visit site for conducvity transmitter not working properly | Visited the site and end user is not able to start the plant hence work is not completed. Hence visit will be planned after OEM confirmation. | Visited the site and end user is not able to start the plant hence work is not completed. Hence visit will be planned after OEM confirmation. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR24067364670 | 2024-08-06 | 2024-08-12 | Cust 1389 | SURAT | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | PH1 filter(1) | visit site for tss issue | ph 1 filter glass chamge by the cousyomer spare now reading found ok | TROUBLESHOOTING DONE | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFSR24068594671 | 2024-08-08 | 2024-08-12 | Cust 610 | SURAT | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | The pending instruments are as follows: pH - 7 units and CIT - 6 units. needs visit for troubleshooting | Checked and detailed mentioned in MOM. | Checked and detailed mentioned in MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR24068594671 | 2024-08-08 | 2024-08-12 | Cust 610 | SURAT | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | The pending instruments are as follows: pH - 7 units and CIT - 6 units. needs to visit for troubleshooting. | Checked and detailed mentioned in MOM. | Checked and detailed mentioned in MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR24073554672 | 2024-08-20 | 2024-08-23 | Cust 1391 | SURAT | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | ETP CPCB data uploded issue. | As of now, we are closing this complaint and will reopen it again once we receive remote access of the system from coustomer end. | As of now, we are closing this complaint and will reopen it again once we receive remote access of the system from coustomer end. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFSR24076844673 | 2024-08-26 | 2024-09-02 | Cust 1390 | SURAT | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | visit site for amc | amc done | AMC DONE | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFSR24076864674 | 2024-08-26 | 2024-08-29 | Cust 1400 | SURAT | Steam System Division | CLOSED | QUALITY PROBLEM | PRV - MEASURETROL | FM MEASURETROL(SIZE 050NB) | in measurtrol controlling is ok but is showing red light indication continuous. also showing error " PILOT VALVE DRIVER CKT FAILURE" | RF OA Number 5525103892 - SPARE-050MTROL-Q1EGS41MI = Qty03Nos | RF OA Number 5525103892 - SPARE-050MTROL-Q1EGS41MI = Qty03Nos | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFSR24076924675 | 2024-08-26 | 2024-08-29 | Cust 1400 | SURAT | Steam System Division | CLOSED | QUALITY PROBLEM | PRV - MEASURETROL | FM MEASURETROL(SIZE 050NB) | modbus data not found. ( when Bluetooth is connected modbus data is found ok) | RF OA Number 5525103892 - SPARE-050MTROL-Q1EGS41MI = Qty03Nos | RF OA Number 5525103892 - SPARE-050MTROL-Q1EGS41MI = Qty03Nos | YES | No Trip | Process | Y | Product / equipment failure due to which customers plant or part of plant is stopped which leaads to production loss. | Analytical - CHAKAN
AFSR24087114676 | 2024-09-11 | 2024-12-05 | Cust 1390 | SURAT | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Replaced Blackbox configuration need to be done. | data was started uploading on the CPCB server. | data was started uploading on the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFSR24087134677 | 2024-09-11 | 2024-09-16 | Cust 1404 | SURAT | Analytical - CHAKAN | CLOSED | COMMISSIONING | Chlorine sensor | CHLORINE SENSOR(1) | Commissioning visit required for Silica, Turbidity and Chlorine analyzer. | Visited the site and checked the installation of the analyzer, detailed mentioned in the MOM. | Visited the site and checked the installation of the analyzer, detailed mentioned in the MOM. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR24089064678 | 2024-09-16 | 2024-09-22 | Cust 610 | SURAT | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | visit site for calibration of 7nos of pH electrode | calibration done found transmitter and sensor working ok | calibration done found transmitter and sensor working ok | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR24089064678 | 2024-09-16 | 2024-09-22 | Cust 610 | SURAT | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | visit site for calibration of 7nos of pH electrode | calibration done found transmitter and sensor working ok | calibration done found transmitter and sensor working ok | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR24089064678 | 2024-09-16 | 2024-09-22 | Cust 610 | SURAT | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | visit site for calibration of 7nos of pH electrode | calibration done found transmitter and sensor working ok | calibration done found transmitter and sensor working ok | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR24089064678 | 2024-09-16 | 2024-09-22 | Cust 610 | SURAT | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | visit site for calibration of 7nos of pH electrode | calibration done found transmitter and sensor working ok | calibration done found transmitter and sensor working ok | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR24089064678 | 2024-09-16 | 2024-09-22 | Cust 610 | SURAT | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | visit site for calibration of 7nos of pH electrode | check and found ok | calibration done found transmitter and sensor working ok | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR24089064678 | 2024-09-16 | 2024-09-22 | Cust 610 | SURAT | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | visit site for calibration of 7nos of pH electrode | calibration done found transmitter and sensor working ok | calibration done found transmitter and sensor working ok | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR24089064678 | 2024-09-16 | 2024-09-22 | Cust 610 | SURAT | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | visit site for calibration of 7nos of pH electrode | calibration done found transmitter and sensor working ok | calibration done found transmitter and sensor working ok | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR24092274679 | 2024-09-23 | 2024-09-30 | Cust 1155 | SURAT | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | MULTIPARAMETER ANALYSER | Xenon Lamp Power supply for Multiparameter(AWA-XLAMPPS) | visit site for xenon lamp supply troubleshooting and for installation guidelines for dust analyzer | troubleshooting done | TROUBLESHOOTING DONE | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR24095034680 | 2024-09-28 | 2024-09-30 | Cust 1392 | SURAT | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB and GPCB data uploaded issue. | Now the data is live on the CPCB and GPCB server. | Now the data is live on the CPCB and GPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFSR24096464681 | 2024-09-30 | 2024-10-05 | Cust 1392 | SURAT | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | VISIT SITE FOR AMC | AMC Done | Amc done | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFSR24101494682 | 2024-10-10 | 2024-10-11 | Cust 1397 | SURAT | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Data not updating on CPCB Server. | We took access to the system and now data is live on CPCB Server. | We took access to the system and now data is live on CPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFSR24101844683 | 2024-10-11 | 2024-10-17 | Cust 1155 | SURAT | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | visit site for power on the analyzer | power on the analyzer and also amc done | power on the analyzer and also amc done | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFSR24101864684 | 2024-10-11 | 2024-10-17 | Cust 1399 | SURAT | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | by mail | visit site check pd1 and pd2 value found lopw clean the flowcell now after that check pd1 and pd2 value 28000 | TROUBLESHOOTING DONE | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR24104414685 | 2024-10-17 | 2024-10-24 | Cust 1155 | SURAT | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | visit site for power on the analyzer | Carried out AMC work and now analyzer working ok. | Carried out AMC work and now analyzer working ok. | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFSR24106064686 | 2024-10-20 | 2024-10-27 | Cust 1404 | SURAT | Analytical - CHAKAN | CLOSED | COMMISSIONING | Silica Analyser | Silica 9210(9210) | Commissioning visit of AWA Silica analyzer. | Silica and Chlorine analyzer commissioning has been. Chlorine sample is not available hence not taken in line. | Silica and Chlorine analyzer commissioning has been. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR24108224687 | 2024-10-24 | 2024-10-27 | Cust 480 | SURAT | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | Ph Tranmitter Smartpro | ORP transmitter(1) | Kindly arrange your depute engineer at Atul valsad site for Calibration of the ORP transmitter and sensor. Our site contact person name : Mr. Saravanan Mno: 9618236633 | Checked the performance and found working satisfactory. Verified the readings with the mV simulators and standard solutions. | Checked the performance and found working satisfactory. Verified the readings with the mV simulators and standard solutions. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR24108224687 | 2024-10-24 | 2024-10-27 | Cust 480 | SURAT | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | Ph Tranmitter Smartpro | ORP transmitter(1) | Kindly arrange your depute engineer at Atul valsad site for Calibration of the ORP transmitter and sensor. Our site contact person name : Mr. Saravanan Mno: 9618236633 | Checked the performance and found working satisfactory. Verified the readings with the mV simulators and standard solutions. | Checked the performance and found working satisfactory. Verified the readings with the mV simulators and standard solutions. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR24108224687 | 2024-10-24 | 2024-10-27 | Cust 480 | SURAT | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | Ph Tranmitter Smartpro | ORP transmitter(1) | Kindly arrange your depute engineer at Atul valsad site for Calibration of the ORP transmitter and sensor. Our site contact person name : Mr. Saravanan Mno: 9618236633 | Checked the performance and found working satisfactory. Verified the readings with the mV simulators and standard solutions. | Checked the performance and found working satisfactory. Verified the readings with the mV simulators and standard solutions. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR24108224687 | 2024-10-24 | 2024-10-27 | Cust 480 | SURAT | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | Ph Tranmitter Smartpro | ORP transmitter(1) | Kindly arrange your depute engineer at Atul valsad site for Calibration of the ORP transmitter and sensor. Our site contact person name : Mr. Saravanan Mno: 9618236633 | Checked the performance and found working satisfactory. Verified the readings with the mV simulators and standard solutions. | Checked the performance and found working satisfactory. Verified the readings with the mV simulators and standard solutions. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR24110434688 | 2024-10-29 | 2024-10-30 | Cust 1391 | SURAT | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | PH East 3 boiler not reflecting live data on CPCB RTDMS portal. | Commercially not clear | Commercially not clear | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFSR24110914689 | 2024-10-30 | 2024-11-04 | Cust 578 | SURAT | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | TOC ANALYSER | FM 101 TOC Analyser(FM 101 TOC) | Required visit for checking healthiness of TOC analyzer. | Checked healthiness at LAB condition and found working ok. | Checked healthiness at LAB condition and found working ok. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR24121194690 | 2024-11-21 | 2024-11-26 | Cust 1155 | SURAT | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | pH input card-FM of Multiparameter(1) | Troubleshooting visit | Checked and seems like issue is with motherboard or Display board, hence suggested for spare. | Checked and seems like issue is with motherboard or Display board, hence suggested for spare. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFSR24128904691 | 2024-12-03 | 2024-12-06 | Cust 1155 | SURAT | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | DATCON | DATCON(DATCON) | CPCB/GPCB data upload issue. | as of now we closing this complaint and will reopen it once the site engineer visits the site. | as of now we closing this complaint and will reopen it once the site engineer visits the site. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFSR24129104692 | 2024-12-04 | 2024-12-06 | Cust 1155 | SURAT | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Visit for troubleshooting of Multiparameter analyzer. | Visited the site and installed the Main power supply board PCX-2, Found working Ok and data transmitted to CPCB/GPCB server. | Visited the site and installed the Main power supply board PCX-2, Found working Ok and data transmitted to CPCB/GPCB server. | NO | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR24129444693 | 2024-12-04 | 2024-12-06 | Cust 1405 | SURAT | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | PH1 detector(1) | AMC visit of multiparameter | Replaced the Ph1 detector and calibration has been done with DM water, found ok. | Replaced the Ph1 detector and calibration has been done with DM water, found ok. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFSR24129494694 | 2024-12-04 | 2024-12-06 | Cust 1155 | SURAT | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB + GPCB data upload issue. | today we took access of the black box and now the hardware side issue is resolved for ETP Outlet and data is live on CPCB + GPCB server. | today we took access of the black box and now the hardware side issue is resolved for ETP Outlet and data is live on CPCB + GPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFSR24130424695 | 2024-12-06 | 2024-12-10 | Cust 1145 | SURAT | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Visit for calibration. | Calibration has been done and found working ok. | Calibration has been done and found working ok. | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFSR24130434696 | 2024-12-06 | 2024-12-10 | Cust 1390 | SURAT | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Visit required for Installation of FMBB and connectivity. | Installation of the FMBB has been done and data was live through sim card. | Installation of the FMBB has been done and data was live through sim card. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFSR24130514697 | 2024-12-06 | 2024-12-09 | Cust 1390 | SURAT | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | internet configuration needs to be done. | Now data is live on CPCB server for GPCB in customers network URL is block and will reopen it once it allow. | Now data is live on CPCB server for GPCB in customers network URL is block and will reopen it once it allow. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFSR24132404698 | 2024-12-10 | 2024-12-13 | Cust 1155 | SURAT | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC & connectivity visit. | Calibration of total 2 nos of SPM and 1 no of MPA, Found ok. | Calibration of total 2 nos of SPM and 1 no of MPA, Found ok. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFSR24135064699 | 2024-12-17 | 2024-12-22 | Cust 1389 | SURAT | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 4th AMC visit | Checked and calibrated the analyzer. | Checked healthiness of the analyzer and carried out the calibration, found working ok. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFSR24136874700 | 2024-12-20 | 2024-12-22 | Cust 1155 | SURAT | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration visit of CX2000 | Calibration has been done and found working ok. | Checked healthiness of the analyzer and carried out the calibration, found working ok. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR24138944701 | 2024-12-24 | 2024-12-26 | Cust 1155 | SURAT | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | GPCB connectivity needs to be done. | we took access of the black box and successfully completed GPCB connectivity, now data is live on the GPCB server for both the parameters and matching with the actual readings. | we took access of the black box and successfully completed GPCB connectivity, now data is live on the GPCB server for both the parameters and matching with the actual readings. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFSR24140124702 | 2024-12-26 | 2025-02-28 | Cust 1145 | SURAT | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | Falcon Scanner for Multiparameter(AWA-SCAN) | modbus output channel hanged randomly, needs to replaced the scanner. | Replacement Material dispatched on 27th Feb 25. | Replacement Material dispatched on 27th Feb 25. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFSR24140594703 | 2024-12-27 | 2025-01-31 | Cust 1400 | SURAT | Steam System Division | CLOSED | QUALITY PROBLEM | PRV - MEASURETROL | FM MEASURETROL(SIZE 050NB) | we check Mtrol on site and found Mtrol not saw accuiat flow under 5Bar presser. | Mr. Jaykumar visited the site and have upgraded the new logic software at site on 09th Jan 2025. The ME MeasureTrol is under observation. | Mr. Jaykumar visited the site and have upgraded the new logic software at site on 09th Jan 2025. The ME MeasureTrol is under observation. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFSR24141094704 | 2024-12-28 | 2024-12-30 | Cust 1155 | SURAT | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | gpcb data upload issue. | We took access of the black box and successfully updated the new Blackbox patch, now real time data is live on CPCB + GPCB server. | We took access of the black box and successfully updated the new Blackbox patch, now real time data is live on CPCB + GPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFSR24143464705 | 2024-12-31 | 2025-01-03 | Cust 1155 | SURAT | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB + GPCB data mismatch issue. | we took access of the black box and successfully updated the new Blackbox patch, now real time data is live on CPCB + GPCB server. | we took access of the black box and successfully updated the new Blackbox patch, now real time data is live on CPCB + GPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFSR24145784706 | 2025-01-03 | 2025-01-06 | Cust 1155 | SURAT | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera connectivity needs to be done. | We are closing this complaint because we have not received the system access. | We are closing this complaint because we have not received the system access. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFSR24145914707 | 2025-01-03 | 2025-01-06 | Cust 1392 | SURAT | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC work | all AMC work activities has been done on site | all AMC work activities has been done on site by Mr. Lakhan Shende | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFSR24148344708 | 2025-01-08 | 2025-01-11 | Cust 1155 | SURAT | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | CPU Display card for Multiparameter(AWA-CPU1) | MPA had a blinking display issue. also AMC of 21XX (2 Qty) having S/R nos. 23023 RCU & 24001 RCU. | Customer has MPA display blank issue. So, Mr. Lakhan Shende visited the site and found no earthing supplied to the panel. Due to this, the CPU of the motherboard got corrupted. Also, carried out AMC of 2 DCEM successfully | Mr. Lakhan Shende visited the site. found CPU faulty due to improper earthing. Need to replace the motherboard. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFSR24154184709 | 2025-01-21 | 2025-01-25 | Cust 1396 | SURAT | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC of MPA | AMC work has been completed. | Mr. Lakhan shende visited the site and carried out AMC activities successfully. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFSR24154374710 | 2025-01-22 | 2025-01-24 | Cust 487 | SURAT | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data upload issue. | we took access of the black box and updated the new patch, now real time is uploading on CPCB + GPCB server. | we took access of the black box and updated the new patch, now real time is uploading on CPCB + GPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFSR24154394711 | 2025-01-22 | 2025-01-25 | Cust 1391 | SURAT | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | pH reading mismatch issue. | We took access of the system and checked data on modscan and whatever readings are coming on the Modscan same readings are uploaded to the server. ( I have attached a screenshot for your reference) Hence this issue needs to be checked from the hardware side. | This issue needs to be checked from the hardware side. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFSR24159414712 | 2025-01-29 | 2025-01-31 | Cust 1155 | SURAT | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | mA Output card for Multiparameter(AWA-4-20mA) | TSS and pH data were not live on the GPCB and CPCB servers. 4-20 mA output module faulty | replaced the 4-20 mA output module as old found faulty. RS485 cable found broken. resolved the same. Now the analyzer is working proper. | Mr. Lakhan Shende replaced the 4-20 mA output module as old found faulty. RS485 cable found broken. resolved the same. Now analyzer is working proper. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR24162244713 | 2025-02-01 | 2025-02-05 | Cust 1155 | SURAT | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | MULTIPARAMETER ANALYSER | CPU Display card for Multiparameter(AWA-CPU1) | Due to improper earthing and direct power supply without UPS power supply. CPU diplay board was corrupted. After repairing, need engineer to install and configuration the same | installed and configured repaired display with CPU motherboard. working OK | MR. Lakhan Shende visited the site; installed and configured repaired display with CPU motherboard. working OK | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR24163054714 | 2025-02-03 | 2025-02-07 | Cust 610 | SURAT | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Electrode | Ph Electrode(FM1104(4C)) | Service visit to check pH, Cond and ORP analyzers. | A total of 11 nos of conductivity, 7 nos of pH, 2 nos of ORP, and 1 no turbidity analyzers have been checked and calibrated. Details mentioned in attached MOM copy | M/s FMPL engineers visited the site 3 days as per service order for calibration of pH, conductivity, ORP and turbidity analyzers. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFSR24167454715 | 2025-02-13 | 2025-02-16 | Cust 1155 | SURAT | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC visit. | PH detector 1 reading low due to inlet sample line pipe tubing found decayed. flowcell contaminated. Need to replace the tubes. | Mr. Lakhan Shende visited the site, found PH detector 1 reading low due to inlet sample line pipe tubing decayed. flowcell contaminated. Need to replace the tubes. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR24170784716 | 2025-02-22 | 2025-02-22 | Cust 1155 | SURAT | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | pH input card-FM of Multiparameter(1) | pH card is not detecting, showing 'wrong' symbol while configuring. | checked and found display module faulty. | Checked and found display module faulty. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFSR24173144717 | 2025-02-26 | 2025-03-01 | Cust 1406 | SURAT | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | AQUADO Analyzer | AQUAMAX DOT(1) | The transmitter display is not turning ON. Also 5 nos of IFC100 is not turning ON. | Checked and found issue in internal circuitry, need to send pune for repairing. | Checked and found issue in internal circuitry, need to send pune for repairing. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR24173144717 | 2025-02-26 | 2025-03-01 | Cust 1406 | SURAT | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | AQUADO Analyzer | AQUAMAX DOT(1) | The transmitter display is not turning ON. Also 5 nos of IFC100 is not turning ON. | Checked and found issue in internal circuitry, need to send pune for repairing. | Checked and found issue in internal circuitry, need to send pune for repairing. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR24177974718 | 2025-03-05 | 2025-03-08 | Cust 1389 | SURAT | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC visit | AMC has been successfully carried out | AMC has been successfully carried out | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFSR24182124719 | 2025-03-14 | 2025-03-18 | Cust 1155 | SURAT | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | CPU Display card for Multiparameter(AWA-CPU1) | MPA display blinking issue | resolved the issue | Mr. Lakhan Shende visited the site and resolved the issue | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFSR24182794720 | 2025-03-17 | 2025-03-20 | Cust 610 | SURAT | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Commissioning visit of 2 ORP sensors, 1 ph sensor and calibration of Turbidity analyzer. | Commissioning has been done and found working ok. | Commissioning has been done and found working ok. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFSR24183224721 | 2025-03-18 | 2025-03-24 | Cust 1145 | SURAT | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC of MPA and commissioning of Falcon mini scanner (S/R: U20824251377) | AMC has been done and calibrated, found ok. | AMC has been done and calibrated, found ok. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFSR24186314722 | 2025-03-25 | 2025-03-27 | Cust 1390 | SURAT | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | GPCB data upload issue. | Now the internet side issue is resolved and data is live on GPCB server. | Now the internet side issue is resolved and data is live on GPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFSR24186794723 | 2025-03-26 | 2025-03-28 | Cust 1392 | SURAT | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC of MPA | AMC of MPA has been carried out successfully | AMC has been carried out successfully | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFSR24186824724 | 2025-03-26 | 2025-04-01 | Cust 1155 | SURAT | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Display faulty. modus data for TSS is not showing | Checked and calibrated the analyzer. | Checked healthiness of the analyzer and carried out the calibration, found working ok. | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFSR25000634725 | 2025-04-02 | 2025-04-07 | Cust 610 | SURAT | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | Ph Tranmitter Smartpro | ORP transmitter(1) | Visit for checking of newly supplied ORP sensor 2 nos | Checked with the standards found working ok and in found difference in process sample, Decided to send both the analyzer along with process sample to Pune for further testing. | Checked with the standards found working ok and in found difference in process sample, Decided to send both the analyzer along with process sample to Pune for further testing. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR25002714726 | 2025-04-08 | 2025-04-10 | Cust 1407 | SURAT | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera connectivity to CPCB server. | We have completed camera configuration. Now the camera is streaming live on the CPCB server. | We have completed camera configuration. Now the camera is streaming live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFSR25004374727 | 2025-04-11 | 2025-05-27 | Cust 1406 | SURAT | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | AQUADO Analyzer | AQUAMAX DOT(1) | 1 nos of AQUAMAX DOT Transmitter not getting ON after one month of installation. | Analyser installed at site and working Ok. | Analyser installed at site and working Ok. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFSR25006094728 | 2025-04-16 | 2025-04-21 | Cust 573 | SURAT | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | AQUAMAX DOT(1) | Installation / Erection | Discussed with the customer and visit will be planned in the week of 28th April. Hence complaint will be reopen in that week. | Discussed with the customer and visit will be planned in the week of 28th April. Hence complaint will be reopen in that week. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR25007204729 | 2025-04-19 | 2025-04-25 | Cust 1404 | SURAT | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4 trans-A4TW033000(1) | Commissioning of multiple products including Special analyzers like calcium hardness and fluoride analyzers from Thermo Fisher, turbidity, pH, conductivity and ORP analyzers. also multiple numbers of Krohne magnetic flowmeters. | Commissioning has been done and will be visit for special analyzer in next week. | Commissioning has been done and will be visit for special analyzer in next week. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR25014224730 | 2025-05-01 | 2025-05-02 | Cust 573 | SURAT | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4 trans-A4TW033000(1) | Commissioning of multiple analytical products included pH+ORP, Conductivity, pH and DO | Commissioning of multiple analytical products included pH+ORP, Conductivity and pH has been done successfully. DO sensor found faulty sent for furter testing to Pune factory from FMPL surat office. | Mr. Lakhan Shende visited the site, Commissioning of multiple analytical products included pH+ORP, Conductivity and pH has been done successfully. DO sensor found faulty. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR25015064731 | 2025-05-03 | 2025-05-06 | Cust 1404 | SURAT | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Multiple pH, Conductivty and ORP sensors commisioning | Visited and attended the site for commissioning of pH, cond, ORP and Turbidity. | Visited and attended the site for commissioning of pH, cond, ORP and Turbidity. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR25016254732 | 2025-05-06 | 2025-05-06 | Cust 1408 | SURAT | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Chlorine Analyzer-MAC 100 | MAC 100(1) | Site visit | Wrong complaint logged by delhi branch. | Wrong complaint logged by Delhi branch. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR25018764733 | 2025-05-12 | 2025-05-15 | Cust 575 | SURAT | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | ORP transmitter(1) | Need support for ORP transmitter installed at ONGC, Mehsana. | Support provided remotely and resolved the wiring connection. found working ok. | Support provided remotely and resolved the wiring connection. found working ok. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR25021934734 | 2025-05-19 | 2025-05-20 | Cust 1145 | SURAT | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC of MPA | AMC work has been carried out successfully. pH electrode found faulty due to no maintainance. | Mr. Lakhan Shende visited the site and AMC work has been carried out successfully. pH electrode found faulty due to no maintainance. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFSR25021944735 | 2025-05-19 | 2025-05-20 | Cust 1145 | SURAT | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC of MPA | AMC work has been carried out successfully. pH electrode found faulty due to no maintainance. | Mr. Lakhan Shende visited the site and AMC work has been carried out successfully. pH electrode found faulty due to no maintenance. | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFSR25027114736 | 2025-05-29 | 2025-06-01 | Cust 1405 | SURAT | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data upload issue. | we took access of the black box and now the time sync issue has been resolved and data is live on the CPCB server. | we took access of the black box and now the time sync issue has been resolved and data is live on the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFSR25028624737 | 2025-05-31 | 2025-06-02 | Cust 1402 | SURAT | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | ORP transmitter(1) | Transmitter not turning power ON | found instrument electronics damaged due to high power/power fluctuations. | Mr. Lakhan Shende visited the site and found instrument electronics damaged due to high power/power fluctuations. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR25028624737 | 2025-05-31 | 2025-06-02 | Cust 1402 | SURAT | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | ORP transmitter(1) | Transmitter not turning power ON | found instrument electronics damaged due to high power/power fluctuations. | Mr. Lakhan Shende visited the site and found instrument electronics damaged due to high power/power fluctuations. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFSR25032974738 | 2025-06-10 | 2025-06-11 | Cust 1389 | SURAT | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | mA Output card for Multiparameter(AWA-4-20mA) | TSS mA not coming out | Needed to reset the analyzer as TSS not showing proper reading in standard as well as sample solutions. Gradually turns to nearer to zero. Also mA configuration for TSS has been done successfully. | AMC work and troubleshooting of mA output & TSS reading has been caried out successfully. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFSR25035784739 | 2025-06-16 |  | Cust 1407 | SURAT | Analytical - CHAKAN | ASSIGNED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera streaming issue at CPCB server. |  |  | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFTH23073554740 | 2023-08-16 | 2023-08-21 | Cust 1409 | THAILAND | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | AQUADO Analyzer | DO2 Arc Sensors(DO2 Arc Sensors) | replacement new DO sensor and calibration with standard | replacement and calibration done. | replacement and calibration done | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFTH23073564741 | 2023-08-16 | 2023-08-21 | Cust 1409 | THAILAND | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Electrode | Ph Electrode(FM1104(4C)) | Replacement new ph sensor and calibration with buffer | Replacement and calibration done. | Replacement and calibration done. | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFTH23181614742 | 2024-03-04 | 2024-03-11 | Cust 584 | THAILAND | Steam System Division | CLOSED | PRODUCT FAILURE | TRAP - MONITORING SYSTEM | FMTMS(FMTMS) | END_USER_Lam Soon FMTMS led light show not correct, need to replace to customer | Material dispatch against RF OA 5524601957 | Material dispatch against RF OA 5524601957 | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFTK23086684743 | 2023-09-11 | 2023-09-13 | Cust 1410 | TURKEY | Steam System Division | CLOSED | PRODUCT DESIGN FAILURE | TRAP - MONITORING SYSTEM | FMTMS(FMTMS) | Cracking issue and performance issue, need to be replaced as per request from Maxval | Material to be dsipatched against RF OA 5524600897 | Material to be dsipatched against RF OA 5524600897 | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFTK23093794744 | 2023-09-26 | 2023-09-27 | Cust 1410 | TURKEY | Analytical - CHAKAN | CLOSED | PRODUCT DESIGN FAILURE | Conductivity Electrode | Forbes Marshall Cell Model 8311, K=0.1(FM -8311) | is advising to install a PP make Flow through Chamber (FGC2011705) in order to solve the problem. | Customer to place order, quote provided. Processing RF will be much more expensive and time consuming | Customer to place order, quote provided. Processing RF will be much more expensive and time consuming | NO | No Trip | Intops | N |  | Analytical - CHAKAN
AFTK24054354745 | 2024-07-16 | 2024-07-16 | Cust 1410 | TURKEY | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Conductivity Electrode | Forbes Marshall Cell Model 8311, K=0.1(FM -8311) | Site - Danone, Turkey. The plastic sheath in which the FM8311 transmitter of the Aquacon SmartPro 8967 system, which we commissioned in September 2023, fits inside cracked and as a result, water leakage occurred. | Damage is physical due to over tightening, can not be provided under warrant and communicated to Maxval | Damage is physical due to over tightening, can not be provided under warrant and communicated to Maxval | YES | No Trip | Intops | N |  | Analytical - CHAKAN
AFTK24136294746 | 2024-12-19 | 2024-12-24 | Cust 1410 | TURKEY | Steam System Division | CLOSED | PRODUCT FAILURE | TRAP - MONITORING SYSTEM | FMTMS(FMTMS) | FMTMS, which was commissioned in August 2023. 58 products were put into operation. However, 16 months have passed and it has been observed that all FMTMS installed in the system are not working. I took the measurements and saw the 0 VDC value of the batteries in the measurements.I am sharing some images with you. In order to test it, I have shared with you the video where I replaced it with a full battery and observed that it worked. Additionally, there are cracks in the white plastic bottom covers of the equipment. | RF 138245161 under draft, awaiting further inputs from maxval to process RF commercially. To be considered for quality review | RF 138245161 under draft, awaiting further inputs from maxval to process RF commercially. To be considered for quality review | NO | No Trip | Intops | N |  | Analytical - CHAKAN
AFVD23033444747 | 2023-06-01 | 2023-06-05 | Cust 1411 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 4th AMC visit for Multiparameter Analyzer | visited site. Found issue in display. Asked customer to send display to Pune factory. | visited site. Found issue in display. Asked customer to send display to Pune factory. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD23035124748 | 2023-06-05 | 2023-06-08 | Cust 1412 | BARODA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration of Multiparameter Analyzer | Pankaj visited site. Maintenance and calibration done. | Pankaj visited site. Maintenance and calibration done. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFVD23035144749 | 2023-06-05 | 2023-06-09 | Cust 578 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 1st AMC visit of Multiparameter Analyzer at Saffron Unit | Hitendra Parmar visited site and AMC done. | Hitendra Parmar visited site and AMC done. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD23035644750 | 2023-06-06 | 2023-06-13 | Cust 1413 | BARODA | Krohne - MIDC | CLOSED | COMMISSIONING | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Commissioning of FMBB for GPCB connectivity unit 2 | Commissioning done. | Commissioning done. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD23035664751 | 2023-06-06 | 2023-06-13 | Cust 1414 | BARODA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | Issue: DO Meter is not showing correct reading | Kinjal Chauhan visiting site on 14th June | Kinjal Chauhan visiting site on 14th June | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD23036754752 | 2023-06-08 | 2023-06-14 | Cust 1411 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC visit for Multiparameter Analyzer at API1 | Uday Rajgor visited site. AMC visit done. | Uday Rajgor visited site. AMC visit done. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD23036774753 | 2023-06-08 | 2023-06-14 | Cust 1054 | BARODA | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Pro-active visit for Multiparameter Analyzer | Visited site and explained benifits of CX2000 analyzer. | Visited site and explained benifits of CX2000 analyzer. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD23036784754 | 2023-06-08 | 2023-06-14 | Cust 578 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 1st AMC visit of 4 Nos COD Analyzers at Ruby division | Pankaj visited site. AMC visit done. | Pankaj visited site. AMC visit done. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD23040034755 | 2023-06-14 | 2023-06-19 | Cust 795 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 1st AMC visit of Multiparameter Analyzer | AMC visit done. | AMC visit done. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD23040584756 | 2023-06-15 | 2023-06-22 | Cust 1415 | BARODA | Analytical - CHAKAN | CLOSED | PRODUCT UPGRADE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | End user: Nissan Bharat Rasayan, Bharuch. PH Meter not working properly. Contact: Shadab: 9730445519 | Visited site. Found pH sensors not responding. Observed high pressure in sample line. | Visited site. Found pH sensors not responding. Observed high pressure in sample line. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD23040594757 | 2023-06-15 | 2023-06-19 | Cust 1054 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC visit of Multiparameter Analyzer at Rau plant | Kinjal Chauhan visited site and AMC done. | Kinjal Chauhan visited site and AMC done. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD23040604758 | 2023-06-15 | 2023-06-19 | Cust 1416 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 1st AMC visit of Multiparameter Analyzer | 1st AMC visit done. Uday visited site. | 1st AMC visit done. Uday visited site. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD23041284759 | 2023-06-18 | 2023-06-25 | Cust 1411 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC visit at API3 Karakhadi | Uday visited site. AMC visit done. | Uday visited site. AMC visit done. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD23044094760 | 2023-06-22 | 2023-06-24 | Cust 1054 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB connectivity needs to be done. | he CPCB connectivity has been completed for the addition of Flow and data is live on the server. | the CPCB connectivity has been completed for the addition of Flow and data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVD23044494761 | 2023-06-23 | 2023-06-29 | Cust 492 | BARODA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Electrode | FM1104(1) | . pH transmitter are not working due to sensor issue | As per attached mail communication asked customer to replace pH electrode as pH electrode is consumable. | As per attached mail communication asked customer to replace pH electrode as pH electrode is consumable. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD23046744762 | 2023-06-27 | 2023-07-03 | Cust 1417 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 1st AMC visit of Multiparameter Analyzer | Uday Rajgor visited site. AMC done. | Uday Rajgor visited site. AMC done. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD23046764763 | 2023-06-27 | 2023-07-03 | Cust 1418 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC visit of Multiparameter Analyzer | AMC of Multiparameter Analyzer done. | AMC of Multiparameter Analyzer done. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD23046774764 | 2023-06-27 | 2023-07-03 | Cust 1419 | BARODA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | troubleshooting of Multiparameter Analyzer | Pankaj visited site. Issue resolved. | Pankaj visited site. Issue resolved. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFVD23046814765 | 2023-06-27 | 2023-07-03 | Cust 578 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 1st AMC visit of COD Analyzer at Pearl Division | Pankaj visited site. AMC done. | Pankaj visited site. AMC done. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD23046854766 | 2023-06-27 | 2023-07-03 | Cust 1420 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC visit of Multiparameter Analyzer | kinjal visited site. AMC job done. | kinjal visited site. AMC job done. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD23047544767 | 2023-06-28 | 2023-06-29 | Cust 1421 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data uploaded issue. | We took access of the BB and made some changes in the configuration so now the correct data of Flow is uploading to the server | The issue has been resolved and data is live hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVD23047634768 | 2023-06-28 | 2023-07-04 | Cust 1422 | BARODA | Analytical - CHAKAN | CLOSED | COMMISSIONING | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Commissioning of FMBB for GPCB connectivity | Uday visited site. Commissioning done. | Uday visited site. Commissioning done. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD23051054769 | 2023-07-04 | 2023-07-10 | Cust 578 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 1st AMC of Multiparameter Analyzer | 1st AMC visit done by Pankaj. | 1st AMC visit done by Pankaj. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD23055774770 | 2023-07-13 | 2023-07-14 | Cust 691 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Black Box Hang issue. | Currently, the plant is in shut condition after starting the customer will inform us hence as of now we are closing this complaint | Currently, the plant is in shut condition after starting the customer will inform us hence as of now we are closing this complaint | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVD23058664771 | 2023-07-19 | 2023-07-24 | Cust 555 | BARODA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | Calibration training of Oil Analyser to User | Kinjal visited site and Calibration done. Training given to customer. | Kinjal visited site and Calibration done. Training given to customer. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD23058674772 | 2023-07-19 | 2023-07-24 | Cust 555 | BARODA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | Calibration training for Oil Analyser to user at IOCL dumad site | Kinjal visited site and Calibration done. Training given to customer. | Kinjal visited site and Calibration done. Training given to customer. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD23058694773 | 2023-07-19 | 2023-07-24 | Cust 555 | BARODA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Aqua4Trans | Aqua4Trans(1) | Calibration of DO Analyzer | Kinjal visited site and Calibration done. Training given to customer. | Kinjal visited site and Calibration done. Training given to customer. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD23058714774 | 2023-07-19 | 2023-07-24 | Cust 555 | BARODA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Aqua4Trans | Aqua4Trans(1) | Calibration training for DO analyser at IOCL dumad site | Kinjal visited site and Calibration done. Training given to customer. | Kinjal visited site and Calibration done. Training given to customer. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD23059324775 | 2023-07-20 | 2023-08-07 | Cust 1411 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB and GPCB connectivity needs to be done. | Now the GPCB server issue has been resolved by the GPCB team. | Now the GPCB server issue has been resolved by the GPCB team. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVD23061914776 | 2023-07-25 | 2023-07-31 | Cust 795 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Troubleshooting of Multiparameter Analyzer for Detector counts fluctuation. | Kinjal chauhan visited site. Troubleshooting done. Analyzer working OK. Currently it is under observation. | Kinjal chauhan visited site. Troubleshooting done. Analyzer working OK. Currently it is under observation. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD23061924777 | 2023-07-25 | 2023-07-31 | Cust 1411 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 1st AMC visit of Multiparameter Analyzer at F1 | 1st AMC visit done by Hitendra Parmar. | 1st AMC visit done by Hitendra Parmar. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD23062204778 | 2023-07-26 | 2023-07-31 | Cust 1423 | BARODA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Conductivity Transmitter Aquacon | Conductivity Transmitter Aquamon 2000(CON-TR-2000) | Troubleshooting of Conductivity Transmitter supplied by M/s Aquamech | Kinjal chauhan visited site. Calibration of Sensor done.Sensor and Transmitter both working satisfactory. | Kinjal chauhan visited site. Calibration of Sensor done.Sensor and Transmitter both working satisfactory. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD23062384779 | 2023-07-26 | 2023-07-31 | Cust 1424 | BARODA | Analytical - CHAKAN | CLOSED | COMMISSIONING | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | Commissioning of Oil In Water Analyzer | Commissioning of Oil in Water Analyzer done. | Commissioning of Oil in Water Analyzer done. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD23062394780 | 2023-07-26 | 2023-07-31 | Cust 1411 | BARODA | Analytical - CHAKAN | CLOSED | COMMISSIONING | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Commissioning of FMBB for GPCB & CPCB connectivity | Uday Rajgor visited site. Configuration of FMBB done. | Uday Rajgor visited site. Configuration of FMBB done. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD23062794781 | 2023-07-27 | 2023-07-31 | Cust 1239 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC visit for Multiparameter Analyzer | Kinjal chauhan visited site. AMC visit done. | Kinjal chauhan visited site. AMC visit done. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD23065724782 | 2023-07-31 | 2023-08-07 | Cust 492 | BARODA | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | Conductivity Electrode | CONDUCELL 2UP-PG-120(HT237 600) | Cable is not working, needs to be replaced-Conductivity Transmitter, | cable is not working due to connector. Connector issue resolved locally. Now it is working ok. | connector issue resolved locally. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD23065824783 | 2023-07-31 | 2023-08-21 | Cust 1239 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Display Board Files are not generating. | The Customer IT team allows FTP ports so now the SV files are generating in the local PC. | Issue has been resolved hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVD23066794784 | 2023-08-01 | 2023-08-08 | Cust 1329 | BARODA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | pH transmitter received at Baroda office for testing and checking purpose. | pH transmitter checked with simulator found ok. Same returned to customer. | pH transmitter checked found ok. | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD23067604785 | 2023-08-03 | 2023-08-10 | Cust 578 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC VISIT | AMC visit done | AMC job done | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD23067604785 | 2023-08-03 | 2023-08-10 | Cust 578 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC WORK | AMC visit done | AMC job done | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD23067604785 | 2023-08-03 | 2023-08-10 | Cust 578 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC WORK | AMC visit done | AMC job done | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD23067604785 | 2023-08-03 | 2023-08-10 | Cust 578 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC WORK | AMC visit done | AMC job done | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD23069834786 | 2023-08-07 | 2023-08-14 | Cust 564 | BARODA | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | Conductivity Electrode | CONDUCELL 2DC-PG-120(HT237 610) | faulty reading | Guided telephonic , Electrode Wiring Corrected, Now it is showing correct reading. | Troubleshooting Job done | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD23070904787 | 2023-08-09 | 2023-08-11 | Cust 1425 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | BB Auto reboot patch needs to be update. | We have updated the BB Auto reboot patch and now the data is live. | We have updated the BB Auto reboot patch and now the data is live. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVD23071694788 | 2023-08-10 | 2023-08-17 | Cust 1425 | BARODA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC visit calibration work to be done | Multiparameter Analyser AMC job done. | AMC job done | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD23071704789 | 2023-08-10 | 2023-08-17 | Cust 1411 | BARODA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration of COD TSS and Ph | MPA analyser showing zero reading of COD. So given telephonic support for calibration of MPA analyser. | Calibration Job done through Telephonic support. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD23076184790 | 2023-08-21 | 2023-08-25 | Cust 1411 | BARODA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | GPCB configuration needs to be done. | Here Black Box LAN port-1 is not communicating so the division said to send the Black Box to Chakan hence we are closing this complaint. | Here Black Box LAN port-1 is not communicating so the division said to send the Black Box to Chakan hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVD23077234791 | 2023-08-22 | 2023-08-29 | Cust 1426 | BARODA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | DO Transmitters | VISI pro DO(VISI pro DO) | Routine Calibration Job do analyser | DO analyser calibration job done | calibration job done | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD23077424792 | 2023-08-23 | 2023-08-30 | Cust 1427 | BARODA | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Display of Smartpro(1) | DPI 201 commissioning job | DPI 201 commissioning job work done. | commissioning job done | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD23082714793 | 2023-09-01 | 2023-09-08 | Cust 578 | BARODA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Multiparameter Analyzer AMC work to be done. | AMC visit done by Pankaj. | AMC visit done by Pankaj. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD23083534794 | 2023-09-03 | 2023-09-11 | Cust 1239 | BARODA | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Cleaning pump replacement. | Cleaning pump replaced with spare , Trial taken found ok. | Acid pump replacement job done. | NO | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD23084404795 | 2023-09-05 | 2023-09-08 | Cust 795 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data uploaded issue. | Data is live on the CPCB server. | Data is live on the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVD23093344796 | 2023-09-25 | 2023-09-27 | Cust 1428 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software crashed need to reinstall software. | Successfully did the software reinstallation + local configuration, now data is live on local server. | Successfully did the software reinstallation + local configuration, now data is live on local server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVD23095294797 | 2023-09-28 | 2023-10-05 | Cust 1429 | BARODA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | CALIBRATION OF MULTIPARAMETER ANALYSER JOB TO BE DONE | Multiparameter Analyser Maintenance and calibration job done. Also Three way valve HMI program corrected for opening and closing issue. | calibration job done | NO | Paid | Core | N |  | Analytical - CHAKAN
AFVD23095304798 | 2023-09-28 | 2023-10-05 | Cust 578 | BARODA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | MULTIPARAMETER ANALYSER CALIBRATION JOB TO BE DONE | Multiparameter analyser calibration and maintenance job done | calibration job done | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD23098564799 | 2023-10-02 | 2023-10-10 | Cust 1421 | BARODA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC job to be done | MPA amc job done | AMC job done | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD23098574800 | 2023-10-02 | 2023-10-10 | Cust 1411 | BARODA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC JOB TO BE DONE | Multi parameter analyser amc job done | AMC job done | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD23102374801 | 2023-10-10 | 2023-10-17 | Cust 1054 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | TSS VALUE FLUCTUATING | CX2000 multiparameter analyser commissioning job done. CPCB and MPCB connectivity job done | commissioning job done | NO | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD23104154802 | 2023-10-14 | 2023-10-14 | Cust 1411 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Black Box configuration needs to be done. | we have completed the Black Box configuration and CPCB + GPCB connectivity and now the data is live on the server. | we have completed the Black Box configuration and CPCB + GPCB connectivity and now the data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVD23107624803 | 2023-10-22 | 2023-10-29 | Cust 1417 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Multiparameter analyser AMC job to be done | Multiparameter analyser amc job done | AMC job done | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD23107644804 | 2023-10-22 | 2023-10-29 | Cust 1430 | BARODA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | decreasing pH reading after calibration. | pH meter calibration job done. Observed pH reading deviation in process fluid. | calibration job done | NO | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD23109034805 | 2023-10-25 | 2023-11-03 | Cust 1418 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Multiparameter analyser AMC job to be done | Multiparameter analyser AMC job done | AMC job done | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD23113304806 | 2023-10-31 | 2023-11-02 | Cust 1239 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB and GPCB data uploaded issue | Not found any issue data is live on the server. | Not found any issue data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVD23116554807 | 2023-11-06 | 2023-11-09 | Cust 1054 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB and MPPCB connectivity issue | After bypassing the Isolator the data started uploading to the server. | After bypassing the Isolator the data started uploading to the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVD23117794808 | 2023-11-08 | 2023-11-15 | Cust 1239 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC JOB TO BE DONE | multiparameter analyser amc job done | AMC job done | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD23130064809 | 2023-12-01 | 2023-12-04 | Cust 1413 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB and GPCB data uploaded issue. | e have changed the Internet IP details in the BB configuration and now the data is live on both the server for unit-1 | e have changed the Internet IP details in the BB configuration and now the data is live on both the server for unit-1 | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVD23132734810 | 2023-12-07 | 2024-01-15 | Cust 1431 | BARODA | Steam System Division | CLOSED | APPLICATION ISSUE | CCDS | CCDS - CONDENSATE CONTAMINATION DETECTION SYSTEM - PH BASE(CCDS-P) | pH sensor (FM1104)is showing wrong pH value of Condensate sample | New Sensor arranged (PolylitePro 120)( RF no. 138232530.) | New Sensor arranged (Polilyt Pro 120)( RF no. 138232530.) | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFVD23134344811 | 2023-12-11 | 2023-12-12 | Cust 1428 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data uploading issue | We did some changes in configuration. Data is now live on the Local server. | We did some changes in configuration. Data is now live on the Local server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVD23135274812 | 2023-12-13 | 2023-12-14 | Cust 1428 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data uploading issue. | Data is now live on the local server. | Data is now live on the local server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVD23137854813 | 2023-12-18 | 2024-01-18 | Cust 893 | BARODA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | acing installation / calibration problem with PH & DO at our Indore site, | Guideline given over phone. issue corrected in wiring and problem has been resolved. | Guideline given over phone. issue corrected in wiring and problem has been resolved. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFVD23139494814 | 2023-12-20 | 2023-12-27 | Cust 578 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC job to be done | Multiparameter analyser amc job done. | calibration job done | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD23139504815 | 2023-12-20 | 2023-12-27 | Cust 1420 | BARODA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | ph 2 detector count fluctuating. | Troubleshooting job done, Checked and found ph2 detector is faulty. | calibration job done | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD23139754816 | 2023-12-21 | 2024-01-31 | Cust 1391 | BARODA | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | Ph Electrode | Mecotrode Pt100 VP 120(HT238 437) | one pH electrode was not response , Please arrange for replacement. Meco Trode Pt 100 VP 120. SR NO.: 35295 OA - 5524101538 | As per attachment mail sent to Mr. Devjyoti to raise the RF. | As per attachment mail sent to Mr. Devjyoti to raise the RF. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD23145444817 | 2023-12-30 | 2023-12-30 | Cust 1432 | BARODA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | Calibration of TDS Transmitter after installation of new Conductivity cell | Myself visited the site,Mail copy attached | Myself visited the site,Mail copy attached | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFVD23149484818 | 2024-01-06 | 2024-01-12 | Cust 1411 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC of Multiparameter Analyzer at API2 | AMC visit at API2 done by Hiren Gohel | AMC visit at API2 done by Hiren Gohel | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD23151544819 | 2024-01-10 | 2024-01-12 | Cust 1232 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | ETP flow data is not updating to the CPCB server. | Data from ETP flow is live on the CPCB server. | Data from ETP flow is live on the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVD23151584820 | 2024-01-10 | 2024-01-12 | Cust 1433 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | GPCB Data upload issue. | There is issue from GPCB server end and we have taken up with GPCB team. | There is issue from GPCB server end and we have taken up with GPCB team. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVD23164274821 | 2024-02-01 | 2024-02-07 | Cust 578 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | telephonic support for COD Analyzer reading issue at Ruby plant | Guided customer for zero calibration over phone. | Guided customer for zero calibration over phone. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD23166484822 | 2024-02-06 | 2024-02-09 | Cust 795 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data uploaded issue | Not found any issue and the data is live. | Not found any issue and the data is live. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVD23167254823 | 2024-02-07 | 2024-02-14 | Cust 575 | BARODA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | issue in pH meter | Pankaj bodkhe visited site. Found issue in pH electrode. Suggested to procure new pH electrode. | Pankaj bodkhe visited site. Found issue in pH electrode. Suggested to procure new pH electrode. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD23167444824 | 2024-02-08 | 2024-02-15 | Cust 1213 | BARODA | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | AQUAMAX DOT(1) | Analyzer comissioning. Oa no. 5524104727 Address: JWIL INFRA LIMITED Jaisingh Nagar, Patta Godam Near Bus Stand, Manpur, Dist. Umaria Umaria-Madhya Pradesh 484665 GST Reg No. : 23AABCJ7376A1ZV M- 9471211553 Contact person: YASHWANT SINGH | Guided customer over phone for commissioning. We will schedule visit if required. | Guided customer over phone for commissioning. We will schedule visit if required. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD23169114825 | 2024-02-12 | 2024-02-16 | Cust 1411 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 3rd AMC visit of Multiparameter Analyzer | 3rd AMC visit done by Hiren | 3rd AMC visit done by Hiren | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD23169884826 | 2024-02-13 | 2024-02-16 | Cust 1239 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC visit for Multiparameter Analyzer | AMC visit done by Hiren Gohel | AMC visit done by Hiren Gohel | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD23170974827 | 2024-02-16 | 2024-02-21 | Cust 1434 | BARODA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | pH Transmitter Aquamon | Aquamon 2000 plus(AQ2000 plus) | calibration of pH transmitter | calibration procedure explained over phone | calibration procedure explained over phone. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD23171054828 | 2024-02-16 | 2024-02-21 | Cust 1099 | BARODA | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Commissioning support for pH transmitter and 3 way valve | discussed with customer over phone.Asked for terminal details to guide proper connection of pH meter and 3 way valve. | discussed with customer over phone.Asked for terminal details to guide proper connection of pH meter and 3 way valve. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD23171064829 | 2024-02-16 | 2024-02-23 | Cust 1420 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration support for Multiparameter Analyzer | Guided customer over phone for calibration of Analyzer. | Guided customer over phone for calibration of Analyzer. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD23173644830 | 2024-02-21 | 2024-02-28 | Cust 1435 | BARODA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | telephonic support for MPA calibration | guided customer over phone for calibration of MPA Analyzer. | guided customer over phone for calibration of MPA Analyzer. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD23175504831 | 2024-02-23 | 2024-03-01 | Cust 927 | BARODA | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | COMMISSIONING | Engineer visit schedule on 11th March | Engineer visit schedule on 11th March | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD23182494832 | 2024-03-06 | 2024-03-13 | Cust 1436 | BARODA | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | PCX1 board not working at site after repairing | PCX1 board RF raised , once material reach at site , visit will be planned. | Troubleshooting Job done | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD23182584833 | 2024-03-06 | 2024-03-13 | Cust 564 | BARODA | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | AQUADO Analyzer | Aquamax system(1) | Complaint-conductivity meter is stucked at 0.00 and it is not showing any reading. | Site visit planned from branch on 20.03.2024. | Troubleshooting job pending , visit planned. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD23183334834 | 2024-03-08 | 2024-03-15 | Cust 1232 | BARODA | Analytical - CHAKAN | CLOSED | COMMISSIONING | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Commissioning of FMBB for GPCB connectivity | Hiren Visited site. Configuration of FMBB for connectivity done. | Hiren Visited site. Configuration of FMBB for connectivity done. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD23183414835 | 2024-03-08 | 2024-04-03 | Cust 1100 | BARODA | Steam System Division | CLOSED | QUALITY PROBLEM | PRV - MEASURETROL | FM MEASURETROL(SIZE 100NB) | inlet and outlet pressures are showing zero.mA output from Pressure transmitter is OK.Need to arrange the material urgently and R&D visit is also required. | VISIT DONE, PCB CARDS REPLACED. MOM ATTACHED | VISIT DONE, PCB CARDS REPLACED. MOM ATTACHED | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFVD23183414835 | 2024-03-08 | 2024-04-03 | Cust 1100 | BARODA | Steam System Division | CLOSED | QUALITY PROBLEM | PRV - MEASURETROL | FM MEASURETROL(SIZE 100NB) | inlet and outlet pressures are showing zero.mA output from Pressure transmitter is OK.Need to arrange the material urgently and R&D visit is also required. | VISIT DONE, PCB CARDS REPLACED. MOM ATTACHED | VISIT DONE, PCB CARDS REPLACED. MOM ATTACHED | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFVD23183694836 | 2024-03-10 | 2024-03-15 | Cust 578 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 3rd AMC visit for COD Analyzer | Hiren Gohel visited site. 3rd AMC visit completed. | Hiren Gohel visited site. 3rd AMC visit completed. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD23184084837 | 2024-03-11 | 2024-03-18 | Cust 927 | BARODA | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Commissioning visit for pH and Conductivity transmitter | Uday Rajgor visited site. Commissioning done. | Uday Rajgor visited site. Commissioning done. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD23184094838 | 2024-03-11 | 2024-03-18 | Cust 1337 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 2nd AMC visit for Multiparameter Analyzer | Hiren Gohel visited site. 2nd AMC visit done. | Hiren Gohel visited site. 2nd AMC visit done. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD23184104839 | 2024-03-11 | 2024-03-18 | Cust 1411 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | troubleshooting of Display of Multiparameter Analyzer | Hitendra Parmar visited site. Display need to be replace. | Hitendra Parmar visited site. Display need to be replace. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD23185674840 | 2024-03-14 | 2024-07-25 | Cust 1437 | BARODA | Analytical - CHAKAN | CLOSED | MANUFACTURING ISSUES | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | PH electrodes are unable to fit inside flanges of T piece. | RF material received and Faulty material sent to HO for RMA. | RF material was received and Faulty material sent to HO for RMA. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFVD23185684841 | 2024-03-14 | 2024-03-15 | Cust 1427 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera not updating on CPCB Server. | We took access to the system and camera is live on CPCB Server. | We took access to the system and camera is live on CPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVD23185784842 | 2024-03-14 | 2024-03-19 | Cust 578 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC visit for COD analyzer at Ruby plant | AMC visit for COD Analyzer done. | AMC visit for COD Analyzer done. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD23190834843 | 2024-03-21 | 2024-03-26 | Cust 1438 | BARODA | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Commissioning of pH meter and magnetic flow meter at Deepak Chemtech Dahej | Pankaj Bodkhe and Hiren Gohel visited site. Commissioning of pH and flow meter done. | Pankaj Bodkhe and Hiren Gohel visited site. Commissioning of pH and flow meter done. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD23192014844 | 2024-03-25 | 2024-03-30 | Cust 578 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 4th AMC visit of Multiparameter Analyzer of Saffron plant | Pankaj Bodkhe visited site and 4th AMC visit done | Pankaj Bodkhe visited site and 4th AMC visit done | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD23193084845 | 2024-03-26 | 2024-03-30 | Cust 492 | BARODA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Conductivity Transmitter, Range : 0-10000 ?s/cm, without sensor, Sr. No. 22SCH651 Not showing proper reading. Shows constant value. Calibration also done with standard solution still the issue persists. | Received transmitter yesterday. We will check it on Monday. | Received transmitter yesterday. We will check it on Monday. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD23193484846 | 2024-03-26 | 2024-03-30 | Cust 1427 | BARODA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Calibration of pH transmitter under AMC visit | Kinjal Chauhan visited site. Calibration of pH trnsmitter done. | Kinjal Chauhan visited site. Calibration of pH trnsmitter done. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD23193504847 | 2024-03-26 | 2024-03-30 | Cust 1439 | BARODA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Silica Analyser | Silica 8891(8891) | Telephonic support for Silica Analyzer for Zero reading issue | Guided customer over phone for silica Analyzer. Asked to change measuring period of sample. | Guided customer over phone for silica Analyzer. Asked to change measuring period of sample. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD23193514848 | 2024-03-26 | 2024-03-30 | Cust 1440 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC visit for all field instruments | Uday Rajgor visited site. Identified field instruments. Functional checking done. | Uday Rajgor visited site. Identified field instruments. Functional checking done. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD23193824849 | 2024-03-27 | 2024-03-30 | Cust 1441 | BARODA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | DO2 transmitter not working satisfactory | Hitendra Parmar and Daksh naik visited site. Troubleshooting of DO meter done. Found issue in transmitter. | Hitendra Parmar and Daksh naik visited site. Troubleshooting of DO meter done. Found issue in transmitter. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD23193854850 | 2024-03-27 | 2024-03-30 | Cust 1420 | BARODA | Analytical - CHAKAN | CLOSED | COMMISSIONING | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Commissioning of FMBB for Multiparameter Analyzer for GPCB connectivity | Hiren Gohel visited site. Configuration of FMBB done.GPCB data live now. | Hiren Gohel visited site. Configuration of FMBB done.GPCB data live now. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD23194604851 | 2024-03-28 | 2024-03-30 | Cust 1418 | BARODA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | telephonic support for Calibration of COD Analyzer | telephonic support given for Calibration of Multiparameter Analyzer | telephonic support given for Calibration of Multiparameter Analyzer | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD23194644852 | 2024-03-28 | 2024-03-29 | Cust 1420 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | GPCB connectivity needs to be done. | Today we have completed the GPCB connectivity of your industry and now the Data is live on the server. | Today we have completed the GPCB connectivity of your industry and now the Data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVD23194854853 | 2024-03-28 | 2024-03-31 | Cust 1442 | BARODA | Steam System Division | CLOSED | QUALITY PROBLEM | PRV - MEASURETROL | FM MEASURETROL(SIZE 050NB) | Inlet pressure readings are showing wrong. | VD2319485 - MADURA COATS PVT LTD - MEASURETROL ELECTRONICS DISPATCHED BY JAYKUMAR R&D | VD2319485 - MADURA COATS PVT LTD - MEASURETROL ELECTRONICS DISPATCHED BY JAYKUMAR RnD | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFVD23194914854 | 2024-03-28 | 2024-03-30 | Cust 578 | BARODA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | telephonic support for Multiparameter to reset the configuration | Pankaj Bodkhe guided customer for Analyzer reset. Reconfiguration of the Analyzer done. | Pankaj Bodkhe guided customer for Analyzer reset. Reconfiguration of the Analyzer done. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD23194944855 | 2024-03-28 | 2024-03-30 | Cust 578 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 4th AMC visit for Multiparameter Analyzer at Saffron plant | Pankaj Bodkhe visited site for 4th AMC of Analyzer. | Pankaj Bodkhe visited site for 4th AMC of Analyzer. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD23194984856 | 2024-03-28 | 2024-03-30 | Cust 1443 | BARODA | Analytical - CHAKAN | CLOSED | COMMISSIONING | pH Transmitter Aquamon | Ph Transmitter Aquamon 4000(PH-Trans-4000) | Calibration and commissioning support for pH and ORP transmitter | Hiren Gohel visited site. Calibration of pH & ORP transmitter done. | Hiren Gohel visited site. Calibration of pH & ORP transmitter done. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD23195534857 | 2024-03-29 | 2024-03-30 | Cust 1436 | BARODA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | troubleshooting of Multiparameter Analyzer at Atladra site of VMC | Pankaj Bodkhe visited site for troubleshooting of Multiparameter Analyzer. Asked customer to send base board to Pune for further repairing. | Pankaj Bodkhe visited site for troubleshooting of Multiparameter Analyzer. Asked customer to send base board to Pune for further repairing. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD23195654858 | 2024-03-29 | 2024-03-30 | Cust 1444 | BARODA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Aqua4Trans | Aqua4Trans(1) | telephonic support for chlorineAnalyzer at Toyo Engineering Dahej site | Pankaj Bodkhe guided customer over phone for Calibration of Chlorine Analyzer | Hitendra Parmar and Hitesh jain visited site for emergency visit of antisurge valve | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD23195694859 | 2024-03-29 | 2024-03-30 | Cust 1419 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | telephonic support for Multiparameter Analyzer display reset issue | Pankaj Bodkhe given support for troubleshooting of Multiparameter Analyzer. Concluded that CPU board is faulty. | Pankaj Bodkhe given support for troubleshooting of Multiparameter Analyzer. Concluded that CPU board is faulty. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD23195714860 | 2024-03-29 | 2024-03-30 | Cust 1418 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 1st AMC visit of Multiparameter Analyzer | Pankaj Bodkhe visited site. 1st AMC visit done on 26th January. | Pankaj Bodkhe visited site. 1st AMC visit done on 26th January. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD23195724861 | 2024-03-29 | 2024-03-30 | Cust 1445 | BARODA | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | Commissioning of DO transmitter | Pankaj Bodkhe visited site for Commissioning of DO transmitter | Pankaj Bodkhe visited site for Commissioning of DO transmitter | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD24000474862 | 2024-04-01 | 2024-04-05 | Cust 1411 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | CUSTOMER REPLACED ANALYSER , SO RECONFIGURATION AND CONNECTIVITY TO BE DONE. | Uday Rajgor visited site. Reconfiguration of the Analyzer done. | Uday Rajgor visited site. Reconfiguration of the Analyzer done. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24000494863 | 2024-04-01 | 2024-04-05 | Cust 1417 | BARODA | Analytical - CHAKAN | CLOSED | DIGITAL CONEECTIVITY | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Multiparameter analyser GPCB connectivity to be done. | Hiren Gohel Visited site. GPCB configuration of FMBB done. | Hiren Gohel Visited site. GPCB configuration of FMBB done. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD24000504864 | 2024-04-01 | 2024-04-05 | Cust 1436 | BARODA | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | DISPLAY WITH CPU BOARD NEED TO CHANGE AFTER REPAIRING. | Kinjal Chauhan visited site. PCX2 board and RS485 card found working satisfactory. Issue observed in CPU board. Replaced from another one. Now Analyzer is working satisfactory. | Kinjal Chauhan visited site. PCX2 board and RS485 card found working satisfactory. Issue observed in CPU board. Replaced from another one. Now Analyzer is working satisfactory. | NO | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD24000654865 | 2024-04-02 | 2024-04-05 | Cust 1403 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | GPCB connectivity needs to be done. | we have completed the GPCB connectivity for ETP and now the data is live on the server. | we have completed the GPCB connectivity for ETP and now the data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVD24000964866 | 2024-04-02 | 2024-04-09 | Cust 927 | BARODA | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | telephonic support for calibration of DO meter | guided over phone for Calibration of New DO sensor. | guided over phone for Calibration of New DO sensor. | NO | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD24001414867 | 2024-04-03 | 2024-04-05 | Cust 1446 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | GPCB connectivity needs to be done. | we have completed the GPCB connectivity and now the data is live on the server. | we have completed the GPCB connectivity and now the data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVD24001904868 | 2024-04-04 | 2024-04-09 | Cust 1239 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Data of Multiparameter Analyzer not updating on CPCB server as well as TSS reading not satisfactory | Visited site. Did necessary configuration of FMBB. Now data is live on server. | Hiren Gohel visited site for Configuration of FMBB.GPCB connectivity done. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24002604869 | 2024-04-05 | 2024-04-09 | Cust 1441 | BARODA | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | troubleshooting visit for DO transmitter | Hitendra Parmar visited site. Installed FM stock sensor and transmitter and working satisfactory. | Hitendra Parmar visited site. Installed FM stock sensor and transmitter and working satisfactory. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD24005644870 | 2024-04-12 | 2024-04-12 | Cust 1239 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data uploaded issue. | Now the issue has been resolved and data is live on the server. | Now the issue has been resolved and data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVD24006364871 | 2024-04-15 | 2024-04-21 | Cust 1411 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC visit for Multiparameter Analyzer at API1 | Hiren Gohel visited site for regular AMC of Analyzer on 4th April. | Hiren Gohel visited site for regular AMC of Analyzer on 4th April. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24006374872 | 2024-04-15 | 2024-04-21 | Cust 1337 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Troubleshooting visit for Multiparameter Analyzer | Hiren Gohel visited site on 6th April. Checked Analyzer performance with standard solutions and found satisfactory. Actual COD of sample observed very low. | Hiren Gohel visited site on 6th April. Checked Analyzer performance with standard solutions and found satisfactory. Actual COD of sample observed very low. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24006514873 | 2024-04-15 | 2024-04-18 | Cust 1411 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera connectivity needs to be done. | Now the Camera connectivity has been completed and now the Camera streaming is live on the server. | Now the Camera connectivity has been completed and now the Camera streaming is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVD24007164874 | 2024-04-17 | 2024-04-24 | Cust 1411 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | visit to check connectivity related issue | Pankaj Bodkhe visited site. Reset the Power of FMBB. Issue resolved. | Pankaj Bodkhe visited site. Reset the Power of FMBB. Issue resolved. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24007214875 | 2024-04-17 | 2024-04-24 | Cust 1447 | BARODA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | pH electrode not responding to buffer solutions | Hitendra Parmar visited site. Asked customer to send pH electrode to Pune factory for further analysis. | Hitendra Parmar visited site. Asked customer to send pH electrode to Pune factory for further analysis. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD24010154876 | 2024-04-23 | 2024-04-29 | Cust 1443 | BARODA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Conductivity Transmitter Aquacon | Aquacon 2000(AQCN 2000 p) | Conductivity transmitter not showing value up to 3 digit. | Hitendra Parmar visited site. Checked and observed that Aquacon 2000 Pro has limitation to set decimal points. Asked customer to send transmitter to Pune. | Hitendra Parmar visited site. Checked and observed that Aquacon 2000 Pro has limitation to set decimal points. Asked customer to send transmitter to Pune. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD24010224877 | 2024-04-23 | 2024-04-29 | Cust 1441 | BARODA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | DO Transmitters | optex series(optex series) | DO sensor not showing correct reading | Kinjal Chauhan visited site. Checked and observed that DO sensor showing high reading. Observed issue in DO sensor cap. | Kinjal Chauhan visited site. Checked and observed that DO sensor showing high reading. Observed issue in DO sensor cap. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD24014424878 | 2024-04-30 | 2024-05-07 | Cust 1411 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 4th AMC visit for Multiparameter Analyzer | 4th AMC visit done by Hiren Gohel. | 4th AMC visit done by Hiren Gohel. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24016504879 | 2024-05-03 | 2024-05-03 | Cust 1419 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Re-installation of software.(EnviroConnect) | successfully completed the re-installation of software ( EnviroConnect ) & Data is live on local server. | successfully completed the re-installation of software ( EnviroConnect ) &amp;amp; Data is live on local server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVD24018804880 | 2024-05-08 | 2024-05-10 | Cust 1448 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not updating to CPCB and MPPCB servers. | Data is now live on the MPPCB + CPCB servers. | Data is now live on the MPPCB + CPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVD24021314881 | 2024-05-15 | 2024-05-17 | Cust 1411 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data uploaded issue. | As discussed, we took access of the BB and found that the BB timing is not synchronized with the internet time. | As discussed, we took access of the BB and found that the BB timing is not synchronized with the internet time. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVD24024024882 | 2024-05-21 | 2024-05-28 | Cust 1449 | BARODA | Analytical - CHAKAN | CLOSED | APPLICATION ISSUE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | An issue in readings with the supplied analysers. | Guided customer over phone. Empty pipe detection is to be turned On. If required we will schedule site visit. | Guided customer over phone. Empty pipe detection is to be turned On. If required we will schedule site visit. | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD24024184883 | 2024-05-21 | 2024-05-28 | Cust 1450 | BARODA | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Need support in commissioning of PH sensor and transmitter. Site address - Grasim Cellulosic Division, Vilayat Plot No. 1, GIDC ? Vilayat Taluka Vagra, Gujarat 392012 Contact Person: Prashant Pandey ( +918347014026) | Discussed with customer. Supplied Aqua4trans working satisfactory. There is issue in pH sensor of other make. | Discussed with customer. Supplied Aqua4trans working satisfactory. There is issue in pH sensor of other make. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD24027214884 | 2024-05-27 | 2024-06-03 | Cust 1411 | BARODA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | FM make current output card(1) | Milli amps. not getting for pH reading. | Hiren and Vishal visited site. Analyzer reset done. Checked mA found OK. | Hiren and Vishal visited site. Analyzer reset done. Checked mA found OK. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24027264885 | 2024-05-27 | 2024-05-27 | Cust 1451 | BARODA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Tranmitter Smartpro | Conductivity Electrode Cable(con-cab) | Conductivity reading issue. | calibration job done. | calibration job done. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD24027304886 | 2024-05-27 | 2024-05-27 | Cust 795 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC VISIT DONE | AMC JOB DONE | AMC JOB DONE | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24027324887 | 2024-05-27 | 2024-05-27 | Cust 1054 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC VISIT | CALIBRATION JOB DONE | calibration job done. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24027364888 | 2024-05-27 | 2024-05-27 | Cust 1054 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MIL card of Multiparameter(1) | AMC VISIT DONE | AMC VISIT DONE | AMC JOB DONE | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24028094889 | 2024-05-28 | 2024-07-25 | Cust 1437 | BARODA | Analytical - CHAKAN | CLOSED | MANUFACTURING ISSUES | Ph Electrode | Polilyte PRO VP 120(HT238 417) | pH Sensor PTFE Holder manufacturing error at our end. | RF material received and Faulty material sent to HO for RMA. | RF material received and Faulty material sent to HO for RMA. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFVD24030324890 | 2024-05-31 | 2024-06-04 | Cust 578 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC VISIT, CLEANING MAINTAINANCE DONE | AMC VISIT, CLEANING MAINTAINANCE DONE | AMC JOB DONE & Done calibration | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24030334891 | 2024-05-31 | 2024-06-04 | Cust 578 | BARODA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | TSS READING ISSUE,DONE CLEANING AND CALIBRATION | TSS READING ISSUE,DONE CLEANING AND CALIBRATION,done amc visit | AMC JOB DONE & Done calibration | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24030434892 | 2024-05-31 | 2024-05-31 | Cust 1442 | BARODA | Steam System Division | CLOSED | QUALITY PROBLEM | PRV - MEASURETROL | FM MEASURETROL(SIZE 050NB) | P2 NOT PERFORMING. SHWOING WRONG READING | NEW ELE WITH ZENER DIODE ARE SENT TO BARODA OFFICE - 31-MAY-24 | NEW ELE WITH ZENER DIODE ARE SENT TO BARODA OFFICE - 31-MAY-24 | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFVD24030774893 | 2024-05-31 | 2024-06-04 | Cust 1421 | BARODA | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | Chlorine sensor | CHLORINE SENSOR(1) | sensor replaced and done calibration | sensor replaced and done calibration | sensor changed andcalibration job done. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD24032424894 | 2024-06-04 | 2024-06-07 | Cust 1420 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | New Analyzer commissioning and CPCB + GPCB connectivity needs to be done. | Now the new Analyzer (CX-2000) has been commissioned successfully at your site and we have also modified the configuration as per the same so data is live on both servers. | Now the new Analyzer (CX-2000) has been commissioned successfully at your site and we have also modified the configuration as per the same so data is live on both servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVD24032494895 | 2024-06-04 | 2024-06-11 | Cust 492 | BARODA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Smart-pro 8966 ORP display not working at site from commissioning, Sr. No. 24SPN129. | Asked customer to send faulty transmitter to Baroda Office for further diagnosis. | Asked customer to send faulty transmitter to Baroda Office for further diagnosis. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD24035184896 | 2024-06-10 | 2024-06-17 | Cust 534 | BARODA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Calibration required for PH transmitter. Site address- ITC Fibre innovations LTD, Badiyakhedi industrial area, Plot No.85, Sehore MP- 466001. Contact person Ram Baran ? 8269474856 | Asked customer to share exact wiring connection to verify the issue. We will visit if require. | Asked customer to share exact wiring connection to verify the issue. We will visit if require. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD24035984897 | 2024-06-12 | 2024-06-19 | Cust 1452 | BARODA | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | Aqua4Trans | Aqua4 trans-A4TW033000(1) | This is to inform you Site assets pH & ORP analyzer installed at ACG-Pithampur-Indore site. Having issue ? pH analyzer in calibration & ORP Analyzer not showing result. For that we required OEM technical support. | Guided customer over phone. pH issue resolved. For ORP meter asked customer to arrange mV source. | Guided customer over phone. pH issue resolved. For ORP meter asked customer to arrange mV source. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD24038074898 | 2024-06-17 | 2024-06-17 | Cust 1416 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | troubleshooting of Multiparameter Analyzer | Kinjal Chauhan and Hiren Gohel visited site. UV Lamp flickering continuously. | Kinjal Chauhan and Hiren Gohel visited site. UV Lamp flickering continuously. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24038124899 | 2024-06-17 | 2024-06-17 | Cust 1337 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | troubleshooting of Multiparameter Analyzer for high TSS reading | Pankaj Bodkhe and Hiren Gohel visited site on 16th April. Calibration verified with standard solution and found satisfactory. | Pankaj Bodkhe and Hiren Gohel visited site on 16th April. Calibration verified with standard solution and found satisfactory. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24038154900 | 2024-06-17 | 2024-06-17 | Cust 1239 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | troubleshooting of Multiparmeter Analyzer for high TSS reading | Pankaj and Hiren visited site on 18th April. troubleshooting done. | Pankaj and Hiren visited site on 18th April. troubleshooting done. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24038174901 | 2024-06-17 | 2024-06-17 | Cust 1453 | BARODA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | troubleshooting of Multiparameter Analyzer | Pankaj and Hiren visited site on 22nd April. | Pankaj and Hiren visited site on 22nd April. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFVD24038254902 | 2024-06-17 | 2024-06-17 | Cust 1418 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC visit of MUltiparameter Analyzer | Hiren Gohel visited site on 29th April. AMC visit done. | Hiren Gohel visited site on 29th April. AMC visit done. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24038314903 | 2024-06-17 | 2024-06-17 | Cust 1454 | BARODA | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4 trans-A4TW033000(1) | commissioning support for chlorine sensor with aqua4trans | Pankaj visited site for commissioning support of chlorine sensor at IOCL dumad site on 29th April. | Pankaj visited site for commissioning support of chlorine sensor at IOCL dumad site on 29th April. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD24038474904 | 2024-06-17 | 2024-06-17 | Cust 1239 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC visit for Multiparameter Analyzer | Hiren Gohel visited site on 7th May. AMC done. | Hiren Gohel visited site on 7th May. AMC done. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24039094905 | 2024-06-18 | 2024-06-21 | Cust 1433 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCb and GPCB data uploaded issue. | Now the data is live on both servers. | Now the data is live on both servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVD24043074906 | 2024-06-26 | 2024-06-29 | Cust 1411 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | camera streaming. | Now the Camera streaming is live on the CPCB server. | Now the Camera streaming is live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVD24043844907 | 2024-06-28 | 2024-06-28 | Cust 1442 | BARODA | Steam System Division | CLOSED | QUALITY PROBLEM | PRV - MEASURETROL | FM MEASURETROL(SIZE 050NB) | DURING THE VISIT THERE WAS A OBSERVATION OF DRY-SOLDER. | THE ELECTRONICS HAVE BEEN CHANGED - THE FAILED ONE HAS BEEN BROUGHT BACK TO R&D FOR RCA | THE ELECTRONICS HAVE BEEN CHANGED - THE FAILED ONE HAS BEEN BROUGHT BACK TO R&amp;amp;D FOR RCA | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFVD24047494908 | 2024-07-02 | 2024-07-09 | Cust 1229 | BARODA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(Ph-Trans-Smartpro) | Observed by customer that PH meter value is malfunction & we done calibration of PH metre with stranded solution but PH meter value not stable in display. | Pankaj and vishal visited site. Found issue in sensor. Transmitter found satisfactory. | Pankaj and vishal visited site. Found issue in sensor. Transmitter found satisfactory. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD24049484909 | 2024-07-05 | 2024-07-05 | Cust 1418 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC visit for Multiparameter Analyzer | Hiren and Vishal visited site on 8th May for AMC visit. | Hiren and Vishal visited site on 8th May for AMC visit. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24049524910 | 2024-07-06 | 2024-07-06 | Cust 1426 | BARODA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration of Multiparameter Analyzer and Oxygen Analyzer | Hiren and Pankaj visited site on 9th & 10th May for Calibration visit | Hiren and Pankaj visited site on 9th & 10th May for Calibration visit | NO | Paid | Core | N |  | Analytical - CHAKAN
AFVD24049544911 | 2024-07-06 | 2024-07-06 | Cust 1239 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC visit for Multiparameter Analyzer | Pankaj and Hiren Visited site for AMC of Multiparameter Analyzer on 13th May | Pankaj and Hiren Visited site for AMC of Multiparameter Analyzer on 13th May | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24049564912 | 2024-07-06 | 2024-07-06 | Cust 1425 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | troubleshooting visit for Multiparameter Analyzer | Hiren Visited site on15th May for troubleshooting of Multiparameter Analyzer. Observed issue in UV lamp SMPS. | Hiren Visited site on15th May for troubleshooting of Multiparameter Analyzer. Observed issue in UV lamp SMPS. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24051124913 | 2024-07-09 | 2024-07-09 | Cust 1436 | BARODA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | troubleshooting of Multiparameter Analyzer at Atladra site. | Vishl visited site on 20th may. Display replaced as spares received. | Vishl visited site on 20th may. Display replaced as spares received. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD24051144914 | 2024-07-09 | 2024-07-09 | Cust 1054 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC visit for Multiparameter Analyzer at Rau plant | Pankaj visited site on 23rd may for AMC. | Pankaj visited site on 23rd may for AMC. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24051154915 | 2024-07-09 | 2024-07-09 | Cust 1436 | BARODA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | troubleshooting of Multiparameter Analyzer at Atladra site | Hiren and Vishal visited site on 25th may. Troubleshooting of Display done. | Hiren and Vishal visited site on 25th may. Troubleshooting of Display done. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD24051164916 | 2024-07-09 | 2024-07-09 | Cust 1420 | BARODA | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Commissioning of CX2000 Analyzer | Vishal and Kinjal visited site. Commissioning of CX2000 Analyzer done on 30th May. | Vishal and Kinjal visited site. Commissioning of CX2000 Analyzer done on 30th May. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD24051784917 | 2024-07-11 | 2024-07-11 | Cust 578 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 1st AMC visit of COD Analyzer | Hiren Gohel visited site on 6th June for AMC of Multiparameter Analyzer | Hiren Gohel visited site on 6th June for AMC of Multiparameter Analyzer | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24055724918 | 2024-07-18 | 2024-07-25 | Cust 492 | BARODA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | Smart-pro 8966 ORP display not working, Sr. No. 24SPN129. | Found issue in Electronics module. Meter sent to Pune for further replacement. | Found issue in Electronics module. Meter sent to Pune for further replacement. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD24058594919 | 2024-07-24 | 2024-07-26 | Cust 1417 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | The PH value freeze in GPCB server | Here issue from the hardware side and our site engineer is visiting the site hence for now we are closing this complaint. | Here issue from the hardware side and our site engineer is visiting the site hence for now we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVD24060004920 | 2024-07-26 | 2024-09-05 | Cust 492 | BARODA | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Smartpro-con) | ORDERED - Resistivity Transmitter, Range: 10 to 100 Mega Ohm/cm RECEIVED - RECEIVED RANGE : 0 TO 25 MEGA OHM/CM | Faulty material sent to HO and arranged for replacement. | Faulty material sent to HO and arranged for replacement. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFVD24060594921 | 2024-07-27 | 2024-08-01 | Cust 1439 | BARODA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Silica Analyser | Silica 9210(9210) | Silica analyzer calibration failure | kinjal chavhan visited site,done maintainance,issue resolved. | kinjal chavhan visited site,done maintainance,issue resolved. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD24062144922 | 2024-07-29 | 2024-08-01 | Cust 1455 | BARODA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | CHANGED TUBING,INSTALLED PH ELECTODE AND DONE CALIBRATION | CHANGED TUBING,INSTALLED PH ELECTODE AND DONE CALIBRATION | pankaj visited site,CHANGED TUBING,INSTALLED PH ELECTODE AND DONE CALIBRATION | NO | Paid | Core | N |  | Analytical - CHAKAN
AFVD24062174923 | 2024-07-29 | 2024-08-05 | Cust 1456 | BARODA | Analytical - CHAKAN | CLOSED | COMMISSIONING | Silica Analyser | Silica 8891(8891) | Commissioning of SWAN Silica Analyzer | Kinjal and Uday visited site for commissioning of Silica Analyzer. Commissioning done. | Kinjal and Uday visited site for commissioning of Silica Analyzer. Commissioning done. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD24062254924 | 2024-07-29 | 2024-08-05 | Cust 1457 | BARODA | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4Trans(1) | Commissioning of Portable Cation Conductivity Analyzer at Unit 1 | Kinjal Chauhan visited site. Commissioning of COnductivity transmitter done. | Kinjal Chauhan visited site. Commissioning of COnductivity transmitter done. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD24062334925 | 2024-07-29 | 2024-07-29 | Cust 1411 | BARODA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | troubleshooting of pH connectivity issue in API1 | Hitendra Parmar visited site. pH calibration done. Connectivity issue resolved. | Hitendra Parmar visited site. pH calibration done. Connectivity issue resolved. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD24062484926 | 2024-07-29 | 2024-07-29 | Cust 534 | BARODA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | troubleshooting of pH meter during commissioning | Uday visited site. Customer has not removed protecting cap from pH electrode. Due to this pH not responding. pH transmitter and sensor working satisfactory. | Uday visited site. Customer has not removed protecting cap from pH electrode. Due to this pH not responding. pH transmitter and sensor working satisfactory. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD24063104927 | 2024-07-30 | 2024-08-02 | Cust 1403 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | GPCB data uploded issue for PH parameter. | now the issue has been resolved. | now the issue has been resolved. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVD24067504928 | 2024-08-06 | 2024-08-09 | Cust 1419 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Enviroconnect Software reinstallation | Issue has been resolved and the data is now live on the Local server. | Issue has been resolved and the data is now live on the Local server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVD24067934929 | 2024-08-07 | 2024-08-12 | Cust 1458 | BARODA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | AQUADO Analyzer | AQUAMAX DOT(1) | During our recent visit to the Dahej site, we identified an issue with one of the DO transmitters that we had installed. While the instrument was initially working well, we noticed that the display is no longer showing any values. Upon inspection, we discovered that the fuse was damaged. We replaced the fuse with a new one, but the issue persisted. | Hiren Gohel visited site. Water ingress observed in DPI. Due to water ingress circuit damaged. Remaining 2 Nos DO sensor checked and found satisfactory. | Hiren Gohel visited site. Water ingress observed in DPI. Due to water ingress circuit damaged. Remaining 2 Nos DO sensor checked and found satisfactory. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD24070334930 | 2024-08-12 | 2024-08-19 | Cust 795 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 2nd AMC visit of Multiparameter Analyzer | Pankaj visited site. 2nd AMC visit done. | Pankaj visited site. 2nd AMC visit done. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24072874931 | 2024-08-19 | 2024-08-22 | Cust 1239 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB and GPCB Server. | We took access to your system and now data is live on CPCB and GPCB Server. | We took access to your system and now data is live on CPCB and GPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVD24072904932 | 2024-08-19 | 2024-08-22 | Cust 1448 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB Server. | We took access to the system and now data is updating on CPCB Server. | We took access to the system and now data is updating on CPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVD24075624933 | 2024-08-23 | 2024-09-03 | Cust 1162 | BARODA | Steam System Division | CLOSED | QUALITY PROBLEM | PRV - MEASURETROL | FM MEASURETROL(SIZE 080NB) | PRV is showing P2 pressure transmitter failure error,Electronics is replaced 20 days ago,Again need to replace. | Mail report attached,PRV is working Now | Mail report attached,PRV is working Now | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFVD24076834934 | 2024-08-26 | 2024-08-27 | Cust 1239 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 1st AMC visit of MPA Analyzer | pankaj visited site,done cleaning,calibration | pankaj visited site,done cleaning,calibration | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24077684935 | 2024-08-27 | 2024-08-27 | Cust 1420 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | troubleshooting of CX2000 Multiparameter Analyzer | vishal visited site,done cleaning,calibration | vishal visited site,done cleaning,calibration | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24077744936 | 2024-08-27 | 2024-08-28 | Cust 578 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 2nd AMC visit of MPA Analyzer | HIREN VISITED SITE AND DONE CLEANING AND CALIBRATION | HIREN VISITED SITE AND DONE CLEANING AND CALIBRATION | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24077764937 | 2024-08-27 | 2024-08-30 | Cust 578 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC visit for Multiparameter Analyzer | HIREN VISITED SITE & DONE INSTALLATION AND CALIBRATION | HIREN VISITED SITE & DONE INSTALLATION AND CALIBRATION | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24077824938 | 2024-08-27 | 2024-08-30 | Cust 795 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | calibration and amc visit | PANKAJ VISITED SITE,DONE CLEANING AND CALIBRATION | PANKAJ VISITED SITE,DONE CLEANING AND CALIBRATION | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24077864939 | 2024-08-27 | 2024-08-30 | Cust 1054 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc and calibration visit | PANKAJ VISITED SITE,DONE CLEANING AND CALIBRATION | PANKAJ VISITED SITE,DONE CLEANING AND CALIBRATION | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24077884940 | 2024-08-27 | 2024-08-30 | Cust 1054 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc visit and calibration | PANKAJ VISITED SITE,DONE CLEANING AND CALIBRATION | PANKAJ VISITED SITE,DONE CLEANING AND CALIBRATION | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24078074941 | 2024-08-27 | 2024-08-28 | Cust 1455 | BARODA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Curtsy visit for Multiparameter Analyzer | VISHAL VISITED SITE CHANGE SMPS AND DONE CALIBRATION | VISHAL VISITED SITE CHANGE SMPS AND DONE CALIBRATION | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD24078194942 | 2024-08-27 | 2024-08-29 | Cust 1455 | BARODA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | CALIBRATION JOB | VISHAL VISITED SITE AND DONE CALIBRATION | VISHAL VISITED SITE AND DONE CALIBRATION | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD24079864943 | 2024-08-30 | 2024-09-02 | Cust 1426 | BARODA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | CALIBRATION FAILED ISSUE AND READING ACCURACY ISSUE | pankaj support on phone call to Nikunj and champak patel,Done wiring connection and sensor calibration.resolved ph reading accuracy issue. | pankaj support on phone call to Nikunj and champak patel,Done wiring connection and sensor calibration.resolved ph reading accuracy issue. | NO | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD24082314944 | 2024-09-02 | 2024-09-09 | Cust 1239 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC visit for MPA Analyzer | Vishal visited site on 31/08. AMC done. | Vishal visited site on 31/08. AMC done. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24082544945 | 2024-09-02 | 2024-09-09 | Cust 1459 | BARODA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Tranmitter Smartpro | Smartpro Electronics Assembly(Smartpro-Elc-Assembly) | ACCURACY ISSUE AND CALIBRATION ISSSUE | Pankaj visited site on 31/08. Sensor calibration done. found satisfactory. | Pankaj visited site on 31/08. Sensor calibration done. found satisfactory. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD24083044946 | 2024-09-03 | 2024-09-03 | Cust 946 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera configuration needs to be done. | We took access of the system and complete camera configuration now camera streaming live on CPCB server. | We took access of the system and complete camera configuration now camera streaming live on CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVD24084404947 | 2024-09-06 | 2024-09-09 | Cust 1054 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB and MPPCB data uploaded issue for ETP. | As of now we are closing this complaint and will reopen it once we get the Black Box access. | As of now we are closing this complaint and will reopen it once we get the Black Box access. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVD24085964948 | 2024-09-09 | 2024-09-16 | Cust 578 | BARODA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | Sample Pump(1) | sample pump not working,keypad touch issue | Pankaj visited site,sample pump not working,keypad touch issue thrn done allignment and sample motor changed issue resolved. | Pankaj visited site,sample pump not working,keypad touch issue thrn done allignment and sample motor changed issue resolved. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24086024949 | 2024-09-09 | 2024-09-16 | Cust 1420 | BARODA | Analytical - CHAKAN | CLOSED | APPLICATION ISSUE | MULTIPARAMETER ANALYSER | Photo Detector for Multiparameter(AWA-PD) | vishal visited site for photo detector signal week issue and reading accuracy issue - cx-2000 model | vishal patel site visited.done calibration,alignment and cleaning | vishal patel site visited.done calibration,alignment and cleaning | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD24086154950 | 2024-09-09 | 2024-09-16 | Cust 1099 | BARODA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | hiren visited site for checking PH transmitter and calibration | hiren visited site and checked PH transmitter and calibration | hiren visited site and checked PH transmitter and calibration | NO | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD24086894951 | 2024-09-11 | 2024-09-13 | Cust 1239 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Flow data not updating on CPCB and GPCB Server. | We took access to the system and now flow data is updating on CPCB and GPCB Server. | We took access to the system and now flow data is updating on CPCB and GPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVD24088974952 | 2024-09-16 | 2024-09-22 | Cust 578 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | pankaj visiting site for AMC and cleaning | Pankaj visited site on 10th Sept for AMC visit. AMC visit done. | Pankaj visited site on 10th Sept for AMC visit. AMC visit done. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24089004953 | 2024-09-16 | 2024-09-22 | Cust 1411 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | hiren visiting site for Multiparameter analyzer AMC and installation MPA display | Hiren visited site for AMC on 11/09/2024. | Hiren visited site for AMC on 11/09/2024. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24089454954 | 2024-09-16 | 2024-09-23 | Cust 1329 | BARODA | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | Ph Electrode | Inchtrode N100F (1 NPT with flat pH membrane design, Kynar)(HT238 352) | pH electrode not working satisfactory | Kinjal Chauhan visited site. Guided customer for correct wiring. Now pH sensor working satisfactory. | Kinjal Chauhan visited site. Guided customer for correct wiring. Now pH sensor working satisfactory. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD24090374955 | 2024-09-18 | 2024-09-18 | Cust 1460 | BARODA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | pH Transmitter Aquamon | Ph Transmitter Aquamon 2000(PH2000) | SHOWING ERRATIC READING & CONFIGURATION ISSUE. | Issue Reported : Showing Erratic reading , Configuration issue. we have checked the configuration found Temp set to 1248 , which is corrected and set to 28 Deg C. Zero calibration performed with simulator , Response checked with simulator found ok. Milli amps output verified found ok. Image attached for your reference. We are sending beck the meter today by shree maruti courier , Docket details as per below. shree Maruti courier : Doc No : 24002200371012 | Troubleshooting Job done | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD24090774956 | 2024-09-19 | 2024-09-20 | Cust 795 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on MPCB Server. | As of now we are closing this complaint and will reopen it once we get access. | As of now we are closing this complaint and will reopen it once we get access. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVD24092284957 | 2024-09-23 | 2024-09-25 | Cust 795 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB server. | We took access to the system and found there was customer side internet issue. Now data is live on CPCB Server. | We took access to the system and found there was customer side internet issue. Now data is live on CPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVD24093294958 | 2024-09-25 | 2024-09-27 | Cust 1232 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Camera streaming issue. | As of now we are closing this complaint. We will reopen this once we get access to your system. | As of now we are closing this complaint. We will reopen this once we get access to your system. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVD24093814959 | 2024-09-26 | 2024-09-28 | Cust 1239 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB Server. | We took access to the system and data is live on CPCB Server. | We took access to the system and data is live on CPCB Server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVD24095054960 | 2024-09-28 | 2024-10-05 | Cust 1099 | BARODA | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | Light guide of Multiparameter (1) | sensor PR-3000180114 And PO No-4700218691 customer need install this senor regarding support for supplier. | Guided customer over phone. Issue resolved. | Guided customer over phone. Issue resolved. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD24098124961 | 2024-10-02 | 2024-10-05 | Cust 1162 | BARODA | Forbes Vyncke | CLOSED | COMMISSIONING | Control & Instrumentation | Automatic Blow Down Control(ABCO) | pH measurement system with sample cooler commissioning | mail report attached | mail report attached | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFVD24099574962 | 2024-10-07 | 2024-10-10 | Cust 1232 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | camera streaming issue. | Now the camera is streaming live on the CPCB server. | Now the camera is streaming live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVD24100344963 | 2024-10-08 | 2024-10-11 | Cust 1411 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | The data has not properly transferring for the last 3 days and Data automatically starts at noon12'O clock. | commercial not clear | commercial not clear | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVD24101444964 | 2024-10-10 | 2024-10-15 | Cust 1461 | BARODA | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Smartpro Electronics Assembly(Smartpro-ele) | smartpro 8966 sr no 23SPN1452 and 8969 S no 23STH073 commissioning. | system started ,to verify values under observation | system started ,to verify values | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFVD24101994965 | 2024-10-12 | 2024-10-14 | Cust 578 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Hiren visited site Done AMC visit 18 to 20 Sept | Hiren visited site Done AMC visit | Hiren visited site Done AMC visit | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24102034966 | 2024-10-12 | 2024-10-14 | Cust 1462 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Hiren visited site for MPA and Gas analyzer, done calibration,AMC Visit ( 25 sept) | Hiren visited site for MPA and Gas analyzer, done calibration,AMC Visit | Hiren visited site for MPA and Gas analyzer, done calibration,AMC Visit | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24102074967 | 2024-10-12 | 2024-10-14 | Cust 1417 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | vishal visited site for MPA analyzer AMC visit ,Done AMC visit. ( 26 Sept oct) | vishal visited site for MPA analyzer AMC visit ,Done AMC visit | vishal visited site for MPA analyzer AMC visit ,Done AMC visit | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24102084968 | 2024-10-12 | 2024-10-14 | Cust 1417 | BARODA | Analytical - CHAKAN | CLOSED | APPLICATION ISSUE | MULTIPARAMETER ANALYSER | Peristaltic cleaning pump head for Multiparameter(AWA-PCPH) | vishal visited site for MPA analyzer trouble shooing visit visit ,Changed flow cell with new one. ( 28 Sept oct) | vishal visited site for MPA analyzer trouble shooing visit visit ,Changed flow cell with new one. | vishal visited site for MPA analyzer trouble shooing visit visit ,Changed flow cell with new one. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24102134969 | 2024-10-12 | 2024-10-14 | Cust 1463 | BARODA | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | Aqua4Trans | Aqua4Trans(1) | Installation TDS ,conductivity sensor in Par formulation pithampur industry. pankaj visited site and Done calibration done cleaning.( 26 sept) | pankaj visited site and Done calibration done cleaning.( 26 sept) | pankaj visited site and Done calibration done cleaning. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFVD24102824970 | 2024-10-14 | 2024-10-21 | Cust 1464 | BARODA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | pH transmitter Smartpro 8966 Sr No : 23SPN957 , Showing Constant Value 14 pH , irrespective of input. | Checked and found issue in electronics module. Sent to Pune for further testing. | Checked and found issue in electronics module. Sent to Pune for further testing. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD24105524971 | 2024-10-19 | 2024-10-25 | Cust 1420 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | Xenon Lamp Power supply for Multiparameter(AWA-XLAMPPS) | Hiren & vishal visiting site for photo detector count low, May be issue in xenon lamp power side. | Hiren and Vishal visited site on 19th October. UV lamp SMPS found faulty. Replaced from available spare. Now analyzer working satisfactory. | Hiren and Vishal visited site on 19th October. UV lamp SMPS found faulty. Replaced from available spare. Now analyzer working satisfactory. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD24105584972 | 2024-10-19 | 2024-10-25 | Cust 1416 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Pankaj visiting site for AMC visit ( 16 oct) | Pankaj visited site on 16th October for AMC. AMC visit done. | Pankaj visited site on 16th October for AMC. AMC visit done. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24106994973 | 2024-10-22 | 2024-10-25 | Cust 795 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data uploaded issue. | We have checked and the data is showing live on the CPCB server. | We have checked and the data is showing live on the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVD24112944974 | 2024-11-05 | 2024-11-08 | Cust 1232 | BARODA | DCS - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | MICROCON+ER SERIES | CONTROLLER MODULE (ER-201)(ER-201) | data is not uploading to CPCB and GPCB servers. | We are closing this complaint. We will reopen this once we get access to the Black Box. | We are closing this complaint. We will reopen this once we get access to the Black Box. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVD24116514975 | 2024-11-12 | 2024-11-16 | Cust 1411 | BARODA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Photo detector signal zero, issue in xenon lamp power supply and MUVX card ( Pankaj & vishal visitted Date- 08/11/2024) | Photo detector signal zero,Done site visit and found issue in issue in xenon lamp power supply and MUVX card | Photo detector signal zero,Done site visit and found issue in issue in xenon lamp power supply and MUVX card | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24118544976 | 2024-11-16 | 2024-11-16 | Cust 1461 | BARODA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | TDS SENSOR CALIBRATION. AQUATDS SMARTPro 8969, TDS SENSOR TYPE:FM8311,K-0.1( SR.NO.- 2311P302) | Feed water TDS meter(Smartpro8969) checked and calibrated observed working fine matching with the lab data.You need to connect the cooling water piping to the sample cooler as discussed | VISIT DONE ON 21.10.24. Feed water TDS meter(Smartpro8969) and pH meter(Smartpro8966) are checked and calibrated observed working fine | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFVD24118544976 | 2024-11-16 | 2024-11-16 | Cust 1461 | BARODA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | pH transmitter calibration.Aquamon SMARTPro8966 | Feed water pH meter(Smartpro8966) checked and calibrated observed working fine matching with the lab data.You need to connect the cooling water piping to the sample cooler as discussed | VISIT DONE ON 21.10.24. Feed water TDS meter(Smartpro8969) and pH meter(Smartpro8966) are checked and calibrated observed working fine | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFVD24118604977 | 2024-11-16 | 2024-11-19 | Cust 1426 | BARODA | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | HIREN GOHEL SITE VISITED FOR PH,ORP,CONDUCTITY ANALYZER COMMISSIONING VISIT. 15.11.24 | hIREN VISITED SITE for commissioning PH,conductivity & ORP , found issue in power supply side E-N= 110 VAC. Customer will resolve this issue then M/s FMPL will visit site. | hIREN VISITED SITE for commissioning PH,conductivity & ORP , found issue in power supply side E-N= 110 VAC. Customer will resolve this issue then M/s FMPL will visit site. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD24119444978 | 2024-11-18 | 2024-11-21 | Cust 1239 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data upload issue. | Now hardware side issue is resolved and data is live on CPCB server | Now hardware side issue is resolved and data is live on CPCB server | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVD24119854979 | 2024-11-19 | 2024-11-21 | Cust 578 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | hiren gohel visiting for MAC visit ( 19-11-24) | hiren gohel visiting for MAC visit, Done AMC visit and cleaning. | hiren gohel visiting for MAC visit, Done AMC visit and cleaning. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24119874980 | 2024-11-19 | 2024-11-21 | Cust 1426 | BARODA | Analytical - CHAKAN | CLOSED | COMMISSIONING | pH Transmitter Aquamon | Ph Transmitter Aquamon 4000(PH-Trans-4000) | Vishal patel visiting site for PH,conductivity and ORP commissioning visit. 19-11-2024 | hiren and vishal visited site,done commissioning visit and done calibration using standard solution Ph,conductivity. | vishal visited site,done commissioning visit and done calibration using standard solution Ph,conductivity. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD24120144981 | 2024-11-19 | 2024-11-26 | Cust 1054 | BARODA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Uday rajgor visiting CX_2000 MPA analyzer for photodetector count approximately zero.may be issue in spare side- MUVX card, SMPS or mother board. | Uday rajgor visited site for CX_2000 MPA analyzer for photodetector count approximately zero.changed MUVX card & SMPS 1000v dc,but issue not resolved, issue in main motherboard card. | Uday rajgor visited site for CX_2000 MPA analyzer for photodetector count approximately zero.changed MUVX card & SMPS 1000v dc,but issue not resolved, issue in main motherboard card. | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24120164982 | 2024-11-19 | 2024-11-26 | Cust 795 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Uday rajgor visiting for AMC visit Visit Date- 19-11-2024 | Uday rajgor visited site and done AMC. Visit date 19/11/24 | Uday rajgor visited site and done AMC. Visit date 19/11/24 | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24120964983 | 2024-11-21 | 2024-11-21 | Cust 1054 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Uday Rajgor visiting for AMC Visit. Date-21-11-24 | uday Rajgor visited site and done AMC and cleaning. | uday Rajgor visited site and done AMC and cleaning. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24121014984 | 2024-11-21 | 2024-11-26 | Cust 1420 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | hiren visiting site for changing level switch and flow cell setting. 20/11/24 | hiren visited site and replaced level switch and configuration setting. | hiren visited site and replaced level switch and configuration setting. | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24121064985 | 2024-11-21 | 2024-11-22 | Cust 1433 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data uploaded issue. | commercially not cleared. | commercially not cleared. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVD24124044986 | 2024-11-27 | 2024-11-28 | Cust 1420 | BARODA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | kinjal chavhan & hiren gohel site visit for reading fluctuation issue. Visit date:26/11/24 | kinjal chavhan & hiren gohel site visited site and done configuration setting,changed Model and graph. Visited date:26/11/24 | kinjal chavhan & hiren gohel site visited site and done configuration setting,changed Model and graph. Visited date:26/11/24 | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24124954987 | 2024-11-28 | 2024-12-02 | Cust 1239 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | uday rajgor visit AMC. Visit date;22/11/2024 | uday rajgor visited visit & done AMC. | uday rajgor visited visit & done AMC. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24124964988 | 2024-11-28 | 2024-12-02 | Cust 1427 | BARODA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | kinjal visit for ph meter, mag meter calibration visit. ( visit date:29/11/2024) | kinjal visited site & done f ph meter, mag meter calibration visit. | kinjal visited site & done f ph meter, mag meter calibration visit. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24127814989 | 2024-12-02 | 2024-12-07 | Cust 1239 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Vishal Patel visit site for MPA analyzer AMC | Vishal Patel visited site for MPA analyzer AMC & done cleaning, calibration, AMC visit date. Visited date: 30/11/2024 | Vishal Patel visited site for MPA analyzer AMC & done cleaning, calibration, AMC visit date. Visited date: 30/11/2024 | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24128674990 | 2024-12-03 | 2024-12-07 | Cust 1418 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Hiren gohel will visit site for MPA analyzer AMC visit Date. Date: 04/12/2024 | Hiren gohel visited site for MPA analyzer AMC visit Date & done cleaning, amc & calibration. Date: 04/12/2024 | Hiren gohel visited site for MPA analyzer AMC visit Date & done cleaning, amc & calibration. Date: 04/12/2024 | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24128784991 | 2024-12-03 | 2024-12-06 | Cust 1411 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue. | As of now we are closing this complaint and will reopen it again once we get remote access of the system from customer end. | As of now we are closing this complaint and will reopen it again once we get remote access of the system from customer end. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVD24129674992 | 2024-12-05 | 2024-12-07 | Cust 1433 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB and GPCB data uploaded issue, | After rebooting the BB the data started updating on both sites and we have also updated the latest modifications on the Blackbox | After rebooting the BB the data started updating on both sites and we have also updated the latest modifications on the Blackbox | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVD24130684993 | 2024-12-07 | 2024-12-14 | Cust 578 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Pankaj visit site for MPA analyzer AMC. | Pankaj visited site and done AMC Visit, changed sample pump and done cleaning. Visited date:5th & 6th dec | Pankaj visited site and done AMC Visit, changed sample pump and done cleaning. Visited date:5th & 6th dec | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24130694994 | 2024-12-07 | 2024-12-14 | Cust 1412 | BARODA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Vishal patel visit site for MPA analyzer troubleshooting visit.07/12/2024 | Vishal patel visited site for MPA analyzer troubleshooting visit. done cleaning and found issue in Xenon lamp SMPS & MPA display Visited date:07/12/2024 | Vishal patel visited site for MPA analyzer troubleshooting visit. done cleaning and found issue in Xenon lamp SMPS & MPA display Visited date:07/12/2024 | NO | Paid | Core | N |  | Analytical - CHAKAN
AFVD24132084995 | 2024-12-09 | 2024-12-14 | Cust 578 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | hiren visit site for MPA analyzer AMC. visit date:09/12/2024 | hiren visited site for MPA analyzer AMC and done amc visit. visited date:09/12/2024 | hiren visited site for MPA analyzer AMC and done amc visit. visited date:09/12/2024 | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24132094996 | 2024-12-09 | 2024-12-14 | Cust 578 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Vishal visit Site for MPA analyzer AMC visit | Vishal visited Site for MPA analyzer AMC visit and done amc visit found display fluctuation some time reset display and issue resolved. visited date 09/12/24 | Vishal visited Site for MPA analyzer AMC visit and done amc visit found display fluctuation some time reset display and issue resolved. visited date 09/12/24 | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24133074997 | 2024-12-12 | 2024-12-13 | Cust 795 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | As per the CPCB approval, need to update the contact details on our server. | As per the approval from the CPCB, we have now updated your contact details on the website | As per the approval from the CPCB, we have now updated your contact details on the website | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVD24133604998 | 2024-12-13 | 2024-12-19 | Cust 952 | BARODA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Turbidity Meter | DIS/Q(182) | kindly depute service engineer on an urgent basis to rectify the issue in the Turbidity Meter. | Pankaj Visited Site & found Turbidity sensor not dipped in sample. Removed sensor cleaned and Done installation properly issue resolved. Also checked conductivity sensor, done proper wiring. Reading issue resolved. Visited date: 18/12/24 | Turbidity sensor not dipped in sample. Removed sensor cleaned and Done installation properly issue resolved. Visited date: 18/12/24 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD24134024999 | 2024-12-14 | 2024-12-17 | Cust 1465 | BARODA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Hiren visit site for DO sensor calibration and MPA analyzer calibration. Visit date: 11/12/24 | Hiren visited site for DO analyzer & MPA analyzer calibration. & done cleaning, calibration. Visited date: 11/12/2024 | Hiren visited site for DO analyzer & MPA analyzer calibration. & done cleaning, calibration. Visited date: 11/12/2024 | NO | Paid | Core | N |  | Analytical - CHAKAN
AFVD24134035000 | 2024-12-14 | 2024-12-17 | Cust 578 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Vishal visit site for MPA analyzer trouble shooting visit, MPA display not working. Visit date: 11/12/24 | Vishal visited site for MPA analyzer trouble shooting visit, and found issue in MPA display, MPA display not working. Needs to be send MPA display & motherboard to pune. Visited date: 11/12/24 | Vishal visited site for MPA analyzer trouble shooting visit, and found issue in MPA display, MPA display not working. Needs to be send MPA display & motherboard to pune. Visited date: 11/12/24 | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24134065001 | 2024-12-14 | 2024-12-17 | Cust 1411 | BARODA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Hiren Visit site MPA analyzer troubleshooting visit, issue in m/A output side. visit date: 14/12/24 | Hiren Visited site MPA analyzer troubleshooting visit, and done issue in m/A output side issue . visited date: 14/12/24 | Hiren Visited site MPA analyzer troubleshooting visit, and done issue in m/A output side issue . visited date: 14/12/24 | NO | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD24135465002 | 2024-12-17 | 2024-12-19 | Cust 1054 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Pankaj visit site for MPA analyzer troubleshooting visit, Photo detector showing approximately zero. Visit date 16th & 17th Dec 2024 | Pankaj visited site and changed main mother board PCB & Photo detector issue resolved. Visited date 16th & 17th Dec 2024 | Pankaj visited site and changed main mother board PCB & Photo detecor issue resolved. Visited date 16th & 17th Dec 2024 | YES | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24135625003 | 2024-12-18 | 2024-12-20 | Cust 1422 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Online data is not streaming in GPCB and CPCB sites | Commercially not clear | Commercially not clear | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVD24135645004 | 2024-12-18 | 2024-12-20 | Cust 1239 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | GPCB data uploaded issue. | Not found any issue as the data was live. | Not found any issue as the data was live. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVD24138765005 | 2024-12-23 | 2024-12-26 | Cust 1426 | BARODA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | DO Transmitters | VISI pro DO(VISI pro DO) | Hiren & vishal will visiting site for DO sensor calibration Visit. visit date : 24/12/2024 | Hiren & vishal visited site & done DO sensor calibration Visit. visited date : 24/12/2024 | Hiren & vishal visited site & done DO sensor calibration Visit. visited date : 24/12/2024 | NO | Paid | Core | N |  | Analytical - CHAKAN
AFVD24140945006 | 2024-12-28 | 2025-01-02 | Cust 526 | BARODA | Analytical - CHAKAN | CLOSED | COMMISSIONING | Silica Analyser | Silica 8891(8891) | Uday Rajgor visit site for silica analyzer commissioning & calibration visit. Visit date: 24/12/2024 | Uday Rajgor visited site for silica analyzer, Done commissioning & calibration visit. Visited date: 24/12/2024 to 27/12/2024 | Uday Rajgor visited site for silica analyzer, Done commissioning & calibration visit. Visited date: 24/12/2024 to 27/12/2024 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD24143535007 | 2024-12-31 | 2025-01-02 | Cust 1239 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera Streaming issue on CPCB | Camera IP is not working on the coustomer network so they are checking with the Camera vendor hence as of now we are closing the complain | Camera IP is not working on the coustomer network so they are checking with the Camera vendor hence as of now we are closing the complain | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVD24145235008 | 2025-01-02 | 2025-01-08 | Cust 578 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Vishal patel visit site for MPA analyzer AMC visit. 02/01/2025 | Vishal patel visited site for MPA analyzer AMC visit. Done AMC Visit 02/01/2025 to 04/01/25 | Vishal patel visited site for MPA analyzer AMC visit. Done AMC Visit 02/01/2025 to 04/01/25 | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24150525009 | 2025-01-14 | 2025-01-27 | Cust 1239 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB AND GPCB DATA UPLOAD ISSUE | We have formatted the BB and completed confirmation again because we were facing some memory issues so now the Data is live on the server | We have formatted the BB and completed confirmation again because we were facing some memory issues so now the Data is live on the server | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVD24151645010 | 2025-01-16 | 2025-01-17 | Cust 1232 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera streaming issue on CPCB server. | Camera is now streaming live on the CPCB server. | Camera is now streaming live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVD24153305011 | 2025-01-20 | 2025-03-17 | Cust 492 | BARODA | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | Conductivity Tranmitter Smartpro | Conductivity Electrode Cable(con-cab) | Conductivity sensor is not working. | Checked and found that insulation cap was placed on sensor inneer tip. Due to which measurement does not take place. Removed cap and now working satisfactory. | Checked and found that insulation cap was placed on sensor inneer tip. Due to which measurement does not take place. Removed cap and now working satisfactory. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD24153365012 | 2025-01-20 | 2025-01-26 | Cust 1466 | BARODA | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Polymetron | Conductivity Electrode Cable(cable-9125) | commissioning | Pankaj support on phone call & done installation for commissioning M/s FMPL engineer will visit site. | Pankaj support on phone call & done installation for commissioning M/s FMPL engineer will visit site. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD24153375013 | 2025-01-20 | 2025-01-26 | Cust 1467 | BARODA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Conductivity Transmitter Aquacon | Aquacon Electronics Assembly(CON-ELE-set-4000) | pH conductivity meter is not showing correct reading during commissioning of the same. | Discussed with customer on phone call, FM team will visit in next week. | Discussed with customer on phone call, FM team will visit in next week. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD24153455014 | 2025-01-20 | 2025-03-03 | Cust 927 | BARODA | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | DO Transmitters | optex series(optex series) | Oxyferm DO sensor failed during commissioning | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210132233 LR Dt - 20-FEB-2025 | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210132233 LR Dt - 20-FEB-2025 | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFVD24155025015 | 2025-01-23 | 2025-01-26 | Cust 1416 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | vishal visit site for MPA analyzer AMC Visit | vishal & subhodip visited site & done MPA analyzer AMC Visit | vishal & subhodip visited site & done MPA analyzer AMC Visit | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24155045016 | 2025-01-23 | 2025-01-26 | Cust 1417 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | vishal & subhodip visit site for MPA aalyzer amc visit. | vishal & subhodip visited site & Done MPA aalyzer amc visit. | vishal & subhodip visited site & Done MPA aalyzer amc visit. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24157645017 | 2025-01-27 | 2025-01-31 | Cust 1411 | BARODA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Vishal visit for issue in mA output. High Reading shown on online server. | Vishal visited for issue in mA output. done Ma calibration & High Reading shown on online server issue resolved. | Vishal visited for issue in mA output. done Ma calibration & High Reading shown on online server issue resolved. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24158925018 | 2025-01-29 | 2025-02-01 | Cust 1422 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Camera Streaming and Flow data uploaded issue. | We have updated the camera format from .ogg to web.m and the issue has been resolved. | We have updated the camera format from .ogg to web.m and the issue has been resolved. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVD24159385019 | 2025-01-29 | 2025-01-31 | Cust 1239 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | camera streaming issue. | We have accessed the system, and the issue has now been resolved. The camera streaming is live on the CPCB server. | We have accessed the system, and the issue has now been resolved. The camera streaming is live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVD24161175020 | 2025-01-31 | 2025-02-06 | Cust 578 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | pankaj visit for sample pump suction issue in ETP MPA analyzer. | changed neoprene tube and replaced photodetector assembly and issue resolved. | changed neoprene tube and replaced photodetector assembly and issue resolved. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24161235021 | 2025-01-31 | 2025-02-06 | Cust 526 | BARODA | Analytical - CHAKAN | CLOSED | COMMISSIONING | Silica Analyser | Silica 8891(8891) | Uday rajgor visit for silica analyzer commissioning visit. | Uday rajgor ten days visited site for silica analyzer commissioning visit & done commissioning visit. | Uday rajgor visit for silica analyzer commissioning visit. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD24164245022 | 2025-02-06 | 2025-02-10 | Cust 578 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Uday Rajgor site visit for MPA analyzer AMC visit. | Uday Rajgor site visited and done MPA analyzer AMC visit. | Uday Rajgor site visited and done MPA analyzer AMC visit. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24167405023 | 2025-02-13 | 2025-02-15 | Cust 1411 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data uploaded issue. | we updated a new patch on the Blackbox. Data is now uploading continuously. | we updated a new patch on the Blackbox. Data is now uploading continuously. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVD24167715024 | 2025-02-14 | 2025-02-24 | Cust 1239 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on MPPCB server. | As of now we are closing this complaint there is issue from MPPCB server side. We have taken up with MPPCB Team. | As of now we are closing this complaint there is issue from MPPCB server side. We have taken up with MPPCB Team. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVD24168535025 | 2025-02-17 | 2025-02-18 | Cust 1411 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | vishal site visit for online data connectivity | vishal site visited and done online configuration data connectivity. | vishal site visited and done online configuration data connectivity. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24168575026 | 2025-02-17 | 2025-02-18 | Cust 1411 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | vishal & subhodip site for Multiparameter analyzer troubleshooting visit. | vishal & subhodip site visited for Multiparameter analyzer troubleshooting visit, changed Xenon lamp smps then done calibration. | vishal & subhodip site visited for Multiparameter analyzer troubleshooting visit, changed Xenon lamp smps then done calibration. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24168645027 | 2025-02-17 | 2025-02-18 | Cust 578 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | uday rajgor site vist for MPA analyzer AMC visit | uday rajgor site visited and done MPA analyzer AMC visit | uday rajgor site visited and done MPA analyzer AMC visit | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24168745028 | 2025-02-17 | 2025-02-18 | Cust 578 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | uday rajgor & subhodip site vist for MPA analyzer AMC visit. | uday rajgor & subhodip site visited and done MPA analyzer AMC visit. | uday rajgor & subhodip site visited and done MPA analyzer AMC visit. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24171265029 | 2025-02-24 | 2025-02-27 | Cust 1420 | BARODA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | kinjal chavhan site visit for flow cell disc changing. | kinjal chavhan site visited and changed flow cell disc changed, range issue resolved. | kinjal chavhan site visited and changed flow cell disc changed, range issue resolved. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD24171325030 | 2025-02-24 | 2025-03-01 | Cust 578 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Customer mail for MPA analyzer AM visit. | vishal patel visited site and done calibration MPA analyzer, AMC visit. | vishal patel visited site and done calibration MPA analyzer, AMC visit. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24172055031 | 2025-02-25 | 2025-02-28 | Cust 1239 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera connectivity needs to be done as the customer has changed the Camera. | We have updated the Camera utility from Ogg to webm and now the Camera streaming is live. | We have updated the Camera utility from Ogg to webm and now the Camera streaming is live. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVD24172065032 | 2025-02-25 | 2025-03-04 | Cust 1411 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | FM BB hanging issue. | Based on the customer's confirmations Since last week connectivity is ok and data is also available. hence as of now, we are closing this complaint. | Based on the customer's confirmations Since last week connectivity is ok and data is also available. hence as of now, we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVD24173835033 | 2025-02-27 | 2025-02-28 | Cust 1232 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera streaming issue on CPCB server | Now the camera is streaming live on the CPCB server. | Now the camera is streaming live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVD24178675034 | 2025-03-07 | 2025-03-10 | Cust 1053 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | MPPCB data upload issue. | We have taken the issue with R&D team, they are working on this issue. | We have taken the issue with R&amp;amp;D team, they are working on this issue. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVD24179335035 | 2025-03-09 | 2025-03-14 | Cust 1418 | BARODA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | subhodip das site visit for multiparameter AMC visit. | subhodip das site visited and done multiparameter AMC visit. | subhodip das site visited and done multiparameter AMC visit. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24180625036 | 2025-03-11 | 2025-03-11 | Cust 1422 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera streaming issue. | Camera streaming window was stopped in customers PC, after restarting the same. now camera is streaming live on CPCB server. | Camera streaming window was stopped in customers PC, after restarting the same. now camera is streaming live on CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVD24181085037 | 2025-03-12 | 2025-03-15 | Cust 1054 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | vishal site visit for multiparameter AMC visit | vishal site visited and done multiparameter AMC visit | vishal site visited and done multiparameter AMC visit | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24181095038 | 2025-03-12 | 2025-03-15 | Cust 1054 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | vishal site visit for multiparameter AMC visit | vishal site visited and done multiparameter AMC visit | vishal site visited and done multiparameter AMC visit | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24181105039 | 2025-03-12 | 2025-03-15 | Cust 1239 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | vishal visit site for multiparameter AMC visit. | vishal site visited and done multiparameter AMC visit | vishal site visited and done multiparameter AMC visit | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD24183255040 | 2025-03-18 | 2025-03-20 | Cust 1411 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera streaming issue on CPCB server. | Now the issue has been resolved by installing the Java and camera streaming is live on the CPCB server. | Now the issue has been resolved by installing the Java and camera streaming is live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVD24183985041 | 2025-03-19 | 2025-03-21 | Cust 1419 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Localhost not opening | Now the localhost is working after installing the Java and data is live on the same. | Now the localhost is working after installing the Java and data is live on the same. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVD24184305042 | 2025-03-20 | 2025-03-21 | Cust 1239 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Need to add new Flow and TDS to the MPPCB server. | We will complete the software configuration once hardware issue is resolved. | We will complete the software configuration once hardware issue is resolved. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVD24184845043 | 2025-03-21 | 2025-03-28 | Cust 1468 | BARODA | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Electrode | Ph Electrode(FM1104(4C)) | commissioning complain | M/s FMPL engineer will visit in next week for PH electrode commissioning visit. | M/s FMPL engineer will visit in next week for PH electrode commissioning visit. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD24187535044 | 2025-03-27 | 2025-03-29 | Cust 1239 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on MPPCB server. | We took access to the system and updated the new changes in URL as per MPPCB and now data is live on MPPCB server. | We took access to the system and updated the new changes in URL as per MPPCB and now data is live on MPPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVD24188035045 | 2025-03-28 | 2025-03-31 | Cust 1433 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB and GPCB data uploded issue. | We took access of the Blackbox and observed that the Flow meter data is not communicating with the Blackbox. Hence, this issue needs to be checked from hardware side and our site engineer is coordinating with them. | This issue needs to be checked from the hardware side, and our site engineer is coordinating with them. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVD24188095046 | 2025-03-28 | 2025-03-31 | Cust 1239 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera Streming issue. | We have resolved the issue, but we have we observed that the system's performance is very slow, and the internet connection is also too slow. Due to the internet issue, the camera streaming is not working continuously; hence, we have asked them to provide a proper internet connection to the system for continuous camera streaming. | The issue has been resolved. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVD25000765047 | 2025-04-03 | 2025-04-04 | Cust 1469 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | DATCON | DATCON(DATCON) | Data not updating on CPCB+MPCB server. | We took access to the system and completed the configuration of Flow and TDS to MPPCB server. | We took access to the system and completed the configuration of Flow and TDS to MPPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVD25001165048 | 2025-04-04 | 2025-04-11 | Cust 1470 | BARODA | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | COMMISSIONING COMPLAINT. | Discussed With Mr. Alpesh shravan sir for MPA analyzer commissioning visit. Next week he will complete the installation of panel and arrange power supply etc. Then M/s FMPL engineer will site visit for MPA analyzer commissioning visit. | Discussed With Mr. Alpesh shravan sir for MPA analyzer commissioning visit. Next week he will complete the installation of panel and arrange power supply etc. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD25003405049 | 2025-04-09 | 2025-04-12 | Cust 1469 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | DATCON | DATCON(DATCON) | Data not updating on CPCB server. | We took access to the system and changed the Salve id from hardware end after which data was live on CPCB server. | We took access to the system and changed the Salve id from hardware end after which data was live on CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVD25004525050 | 2025-04-11 | 2025-04-11 | Cust 1453 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | GPCB DATA UPLODE ISSUE. | We took access of the FMBB and found a network issue, and after the same issue has been resolved by the customer's IT team after that the Data is live. | We took access of the FMBB and found a network issue, and after the same issue has been resolved by the customer's IT team after that the Data is live. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVD25006375051 | 2025-04-17 | 2025-04-19 | Cust 1422 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB server and camera not playing on CPCB. | We took access to the system and now data is uploading on CPCB server. For camera issue is from customer end. | We took access to the system and now data is uploading on CPCB server. For camera issue is from customer end. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVD25006505052 | 2025-04-17 | 2025-04-19 | Cust 1469 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB server. | We took access to the system and now data is live on CPCB server. | We took access to the system and now data is live on CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVD25007885053 | 2025-04-21 | 2025-04-26 | Cust 1417 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | pankaj visit the site for AMC visit and Changing flow cell. | pankaj visited the site and done AMC visit and Changed flow cell with new one. | pankaj visited the site and done AMC visit and Changed flow cell with new one. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD25008515054 | 2025-04-23 | 2025-04-26 | Cust 526 | BARODA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Aqua4Trans | Aqua4Trans(1) | Sometime display is shown then after some time the display is automatically turned off. AQUA4TRANS MODEL-A4TW066000 SR.NO-22A4T0635-W066000 | uday rajgor visited the site and checked AQUA4TRANS,tightned all flat cables and card then issue resolved | uday rajgor visited the site and checked AQUA4TRANS,tightned all flat cables and card then issue resolved | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD25008525055 | 2025-04-23 | 2025-04-24 | Cust 1469 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating for STACK 3 and STACK 4 on MPPCB server. | We took access to the system and updated URL changes and now data is live on MPPCB server. | We took access to the system and updated URL changes and now data is live on MPPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVD25008645056 | 2025-04-23 | 2025-04-26 | Cust 1433 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Data Availability issue on GPCB server. | We took access to the FM Blackbox and implemented some modifications. In addition, we have observed that the internet connection is unstable, as it was frequently disconnected during remote access. Hence, we have asked them to ensure a stable and continuous internet connection to facilitate seamless data transmission. Currently, the data is live on the GPCB and CPCB servers. | Currently the data is live. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVD25009295057 | 2025-04-24 | 2025-04-24 | Cust 1239 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | New configuration for flowmeter and TDS connectivity | We took access to the system and completed the TDS and Flowmeter connectivity via blackbox. | We took access to the system and completed the TDS and Flowmeter connectivity via blackbox. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVD25009615058 | 2025-04-25 | 2025-04-28 | Cust 1239 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera re configuration in new PC | We have completed camera configuration on the new PC. Now the camera is streaming live on the CPCB server. | We have completed camera configuration on the new PC. Now the camera is streaming live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVD25009725059 | 2025-04-25 | 2025-05-01 | Cust 1420 | BARODA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | pankaj visit the site for tss reading fluctation and sample pump not suctioned inlet sample issue. | pankaj visited and clened sample pump,sample tank,strainer abd issue resolved. | pankaj visited and clened sample pump,sample tank,strainer abd issue resolved. | NO | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD25010775060 | 2025-04-28 | 2025-04-30 | Cust 1469 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | DATCON | DATCON(DATCON) | Data not updating on MPPCB and CPCB server. | We took access to the system and data is live now on CPCB server. There is internet side issue due to which data is not continuous asked customer to change internet connected to DATCON. | We took access to the system and data is live now on CPCB server. There is internet side issue due to which data is not continuous asked customer to change internet connected to DATCON. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVD25012955061 | 2025-04-30 | 2025-04-30 | Cust 1239 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Data not updating on CPCB server. | There is hardware side issue for Flowmeter. Same communicated to customer and site engineer. | There is hardware side issue for Flowmeter. Same communicated to customer and site engineer. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVD25014765062 | 2025-05-02 | 2025-05-02 | Cust 1471 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on MPPCB server URL need to be updated. | We have updated the URL and data is now live on MPPCB server. | We have updated the URL and data is now live on MPPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVD25016975063 | 2025-05-07 | 2025-05-10 | Cust 1426 | BARODA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | SUBHODIP SITE VISIT FOR MPA ANALYZER AND O2 ANALYZER CALIBRATION VISIT. | SUBHODIP SITE VISITED SITE AND DONE MPA ANALYZER AND O2 ANALYZER CALIBRATION VISIT. | SUBHODIP SITE VISITED SITE AND DONE MPA ANALYZER AND O2 ANALYZER CALIBRATION VISIT. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFVD25017025064 | 2025-05-07 | 2025-05-10 | Cust 554 | BARODA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | DO Transmitters | VISI pro DO(VISI pro DO) | As per the subject purchase order, the DO transmitter readings are fluctuating. End User Name :- Coromandel International Ltd. Ankleshwar, Gujarat Site Person Name :- Ayush More Site Person :- 9284998491 | M/s FMPL enginner Pankaj discussed with Ayush more on phone call,Customer told that four DO sensor working normal in open atmosphere.fluation in sample, suggested checked in stable sample tank or zero check. | M/s FMPL enginner Pankaj discussed with Ayush more on phone call,Customer told that four DO sensor working normal open atmosphere. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD25019265065 | 2025-05-13 | 2025-05-16 | Cust 1239 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB and GPCB data upolded issue. | After restarting the isolator, the data starts updating to the server.Hence, we are closing this complaint. | After restarting the isolator, the data starts updating to the server.Hence, we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVD25019275066 | 2025-05-13 | 2025-05-16 | Cust 1054 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB and MPPCB data upload issue for ETP unit. | We took access of the BB and identified that the internet LAN cable was connected incorrectly, due to this reason the internet was not pinging to the Blackbox.Hence we request that you please ensure the internet cable is connected to the first port of the Blackbox and After correcting the connection, the data starts updating to the server. | the issue is now resolved, and data is live on both sites. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVD25019305067 | 2025-05-13 | 2025-05-20 | Cust 1472 | BARODA | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Polymetron | Conductivity Tranmitter Polymetron 9125(9125) | commissioning complain | Discussed with customer and suggested to arrange low range calibration standard solution. Once low range calibration standard solution will be arranged we will provide support to customer if required. | Discussed with customer and suggested to arrange low range calibration standard solution. Once low range calibration standard solution will be arranged we will provide support to customer if required. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD25020355068 | 2025-05-15 | 2025-05-19 | Cust 1422 | BARODA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | DO Transmitters | VISI pro DO(VISI pro DO) | pankaj visit the site for o2 sensor calibration and troubleshooting visit. | pankaj visited the site for o2 sensor calibration and troubleshooting visit. done one o2 sensor calibration and other two sensor faulty. | pankaj visited the site for o2 sensor calibration and troubleshooting visit. done one o2 sensor calibration and other two sensor faulty. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFVD25020365069 | 2025-05-15 | 2025-05-19 | Cust 1473 | BARODA | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | pankaj visit the site for ph and conductivity calibration. | pankaj visited the site for ph and conductivity calibration and done calibration. | pankaj visited the site for ph and conductivity calibration and done calibration. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD25020375070 | 2025-05-15 | 2025-05-19 | Cust 1416 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | subhodip visit the site for MPA analyzer amc visit | subhodip visited the site for MPA analyzer amc visit and done AMC visit. | subhodip visited the site for MPA analyzer amc visit and done AMC visit. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD25020565071 | 2025-05-16 | 2025-05-16 | Cust 1469 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | DATCON | DATCON(DATCON) | Data is not updating on CPCB server. | We took access to the DATCON and now data is live on CPCB server. | We took access to the DATCON and now data is live on CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVD25021745072 | 2025-05-19 | 2025-05-22 | Cust 1422 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB server. | We took access to the system and observed data is not coming till modscan. Discussed with customer and asked to check from hardware end. | We took access to the system and observed data is not coming till modscan. Discussed with customer and asked to check from hardware end. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVD25023445073 | 2025-05-22 | 2025-05-23 | Cust 1474 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | MPPCB data upload issue. | We took access of the BB and updated it with the Latest version 1.9 and also modify the MPPCB url from HTTP to HTTPS so now the Data is live on both the sites. | We took access of the BB and updated it with the Latest version 1.9 and also modify the MPPCB url from HTTP to HTTPS so now the Data is live on both the sites. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVD25023765074 | 2025-05-23 | 2025-05-26 | Cust 1469 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | DATCON | DATCON(DATCON) | STACK data not updating correct on CPCB software. | We took access to the system and checked data on modscan and same data is reflecting on CPCB server. | We took access to the system and checked data on modscan and same data is reflecting on CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVD25024935075 | 2025-05-26 | 2025-06-05 | Cust 1475 | BARODA | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Conductivity Tranmitter Smartpro | Smartpro Electronics Assembly(Smartpro-Elc-Assembly) | Conductivity Transmitter showing garbage value of Conductivity/TDS around 40000-50000. model:SMARTPro 8967 Ex. It was supplied in DEC-2023, but recently it was taken under commissioning due to project delay. | Replacement material against OA-5526101445(Invoice-5565001625)Dispatched today By Air Vide Mark Express Docket No-210141276 dtd 04.06.2025 | Replacement material against OA-5526101445(Invoice-5565001625)Dispatched today By Air Vide Mark Express Docket No-210141276 dtd 04.06.2025 | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFVD25025545076 | 2025-05-27 | 2025-05-30 | Cust 1469 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | DATCON | DATCON(DATCON) | Wrong data updating for both the STACK 1 and STACK 2 on CPCB server. | We took access to the system and now on both the STACK1 and STACK2 correct data is updating on CPCB server. | We took access to the system and now on both the STACK1 and STACK2 correct data is updating on CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVD25026655077 | 2025-05-28 | 2025-06-02 | Cust 795 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | subhodip visit the site for MPA analyzer visit. | subhodip visited the site for MPA analyzer visit, done MPA analyzer AMC visit. | subhodip visit the site for MPA analyzer visit. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD25026665078 | 2025-05-28 | 2025-06-02 | Cust 1476 | BARODA | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | subhodip site visited for cx-2000 MPA analyzer commissioning visit | subhodip site visited and done cx-2000 MPA analyzer commissioning visit. | subhodip site visited and done cx-2000 MPA analyzer commissioning visit. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD25030105079 | 2025-06-02 | 2025-06-05 | Cust 1239 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on MPPCB Server. | We took access to the system and now data is live on MPPCB server. | We took access to the system and now data is live on MPPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVD25031035080 | 2025-06-04 | 2025-06-07 | Cust 1195 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB+MPPCB server. | Issue has been resolved hence closing this complaint. | Issue has been resolved hence closing this complaint. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVD25031625081 | 2025-06-05 | 2025-06-07 | Cust 1411 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Data not updating on CPCB server. | Commercially not clear. | Commercially not clear. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVD25032065082 | 2025-06-06 | 2025-06-09 | Cust 795 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | FMBB commissioning | we took access of the black box and completed FMBB configuration now data is coming till the putty for TDS parameter. | we took access of the black box and completed FMBB configuration now data is coming till the putty for TDS parameter. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVD25032145083 | 2025-06-06 | 2025-06-11 | Cust 1411 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Subhodip site visit for Multiparameter AMC visit | Subhodip site visited and done Multiparameter AMC visit | Subhodip site visited and done Multiparameter AMC visit | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD25032225084 | 2025-06-07 | 2025-06-11 | Cust 1418 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | subhodip site visit for MPA Analyzer AMC visit | subhodip site visited and done MPA Analyzer AMC visit | subhodip site visited and done MPA Analyzer AMC visit | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVD25032235085 | 2025-06-07 | 2025-06-11 | Cust 1470 | BARODA | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | subhodip site visit for cx-2000 commissioning visit | subhodip site visited and done cx-2000 commissioning visit | subhodip site visited and done cx-2000 commissioning visit | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVD25032665086 | 2025-06-09 | 2025-06-12 | Cust 1239 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | Camera not playing on CPCB server. | We took access to the system and observed that java was uninstalled and after installation camera was live on CPCB server. | We took access to the system and observed that java was uninstalled and after installation camera was live on CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVD25032675087 | 2025-06-09 | 2025-06-11 | Cust 795 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | MPPCB configuration for Flow and TDS. | We took access to the system and completed the MPPCB configuration and now data is live on MPPCB server. | We took access to the system and completed the MPPCB configuration and now data is live on MPPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVD25033015088 | 2025-06-10 | 2025-06-12 | Cust 1411 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera streaming issue on CPCB server. | We have installed the java and now the camera is streaming live on the CPCB server. | We have installed the java and now the camera is streaming live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVD25034445089 | 2025-06-12 | 2025-06-13 | Cust 1239 | BARODA | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | DATA mismatch of software and hardware. | We took access to the system and check putty and server data is matching. | We took access to the system and check putty and server data is matching. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVD25035075090 | 2025-06-14 | 2025-06-17 | Cust 1337 | BARODA | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | subhodip site visit for multiparameter AMC Visit | subhodip site visited and done multiparameter AMC Visit | subhodip site visited and done multiparameter AMC Visit | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH23034065091 | 2023-06-02 | 2023-06-06 | Cust 1477 | VASHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4Trans(1) | commissioning of Aqua4trans and turbidity sensor | commissioning done for all 4 aqua4trans tx and checked found working normal. | commissioning done for all 4 aqua4trans tx and checked found working normal. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH23034905092 | 2023-06-05 | 2023-06-10 | Cust 1478 | VASHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc | amc work done | amc work done | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVH23036935093 | 2023-06-08 | 2023-06-12 | Cust 1478 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MUVX Card of Multiparameter(1) | amc | amc work done | amc work done | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVH23038125094 | 2023-06-12 | 2023-06-15 | Cust 1263 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data upload issue. | We took the access to your system and now the issue has been resolved and data is live on all the servers. | We took the access to your system and now the issue has been resolved and data is live on all the servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH23038415095 | 2023-06-12 | 2023-06-17 | Cust 1444 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | AQUADO Analyzer | OXYFERM(OXYFERM) | wrong reading | found sensor membrane broken , requested to buy new membrane | found sensor membrane broken , requested to buy new membrane | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH23039895096 | 2023-06-14 | 2023-06-15 | Cust 1479 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | data mismatch with actual value. | Now data is matching to the actual value hence we are closing this complaint. | Now data is matching to the actual value hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH23040575097 | 2023-06-15 | 2023-06-15 | Cust 1337 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | localhost is not accessible. | Found apache tomcat service in stopped mode, after restating the same, localhost is accessible. | Found apache tomcat service in stopped mode, after restating the same, localhost is accessible. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH23040885098 | 2023-06-16 | 2023-06-19 | Cust 1480 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data upload issue. | We took the access to the system and now the issue has been resolved and data is uploaded onto the MPCB server. | We took the access to the system and now the issue has been resolved and data is uploaded onto the MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH23040935099 | 2023-06-16 | 2023-06-19 | Cust 1263 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data upload issue | We took the access to the system and did some changes in software configuration end, currently data is available onto the MPCB server. | We took the access to the system and did some changes in software configuration end, currently data is available onto the MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH23041375100 | 2023-06-18 | 2023-06-21 | Cust 692 | VASHI | Analytical - CHAKAN | CLOSED | INSTALLATION/ ERECTION | Chlorine Analyzer-MAC 100 | MAC 100(1) | Installation Guidelines of chlorine analyzer. OPTISYS CL1100+MAC100W+MOUNTING | Visited site and given the installation guidelines of chlorine analyser system. | Visited site and given the installation guidelines of chlorine analyser system. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFVH23042215101 | 2023-06-19 | 2023-06-22 | Cust 1479 | VASHI | Analytical - CHAKAN | CLOSED | SALES PROMOTION | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB + Local data upload issue. | Now data is live on MPCB server, hence we are closing this complaint. | Now data is live on MPCB server, hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH23042265102 | 2023-06-19 | 2023-06-21 | Cust 1263 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data upload issue. | Now the network issue has been resolved and data is live on the MPCB server. | Now the network issue has been resolved and data is live on the MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH23042365103 | 2023-06-19 | 2023-06-23 | Cust 608 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | current output of TSS not coming | visited site and found current output not coming and so suggested to procure new current output. | visited site and found current output not coming and so suggested to procure new current output. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH23042945104 | 2023-06-20 | 2023-06-22 | Cust 1481 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc | done amc and checked found working normal. | done amc and checked found working normal. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVH23044585105 | 2023-06-23 | 2023-06-26 | Cust 1263 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data mismatch with actual value. | For now we are closing this complaint, we will reopen it once the customer turn on the flow meter. | For now we are closing this complaint, we will reopen it once the customer turn on the flow meter. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH23052175106 | 2023-07-06 | 2023-07-09 | Cust 867 | VASHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Silica Analyser | Silica 9210(9210) | awa silica analyser | commissioning of silica analyser done and checked found working normal. | commissioning of silica analyser done and checked found working normal. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH23053635107 | 2023-07-10 | 2023-07-16 | Cust 1482 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | AQUADO Analyzer | Aquamax system(1) | no display | found fuse blown , observed water ingress . collected dpi201 and sent to pune for repairs | found fuse blown , observed water ingress . collected dpi201 and sent to pune for repairs | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH23053655108 | 2023-07-10 | 2023-07-16 | Cust 1478 | VASHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | error in readings | checked sample , found cod is low only , checked with khp found ok. | checked sample , found cod is low only , checked with khp found ok. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVH23054535109 | 2023-07-11 | 2023-07-26 | Cust 1263 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB and MPCB data uploaded issue. | Now data is uploading on CPCB + MPCB server | Now data is uploading on CPCB + MPCB server | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH23055625110 | 2023-07-12 | 2023-07-17 | Cust 1483 | VASHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4Trans(1) | commissioning of DO sensor visiwater DO arc 120. new menbrane to be changed. | commissioning done. | commissioning done. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH23060845111 | 2023-07-24 | 2023-07-27 | Cust 608 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation needs to be done (PC change) | Customer has not change PC name. | Customer has not change PC name. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH23061375112 | 2023-07-25 | 2023-07-31 | Cust 1478 | VASHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | tss reading issue | found xenon power supply card faulty , replaced from customers stock | found xenon power supply card faulty , replaced from customers stock | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH23061455113 | 2023-07-25 | 2023-07-26 | Cust 1479 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Need to change Totalizer range. | We took the access to the system and successfully change the Totalizer range. | We took the access to the system and successfully change the Totalizer range. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH23062745114 | 2023-07-27 | 2023-07-28 | Cust 1263 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue. | data not coming till the local system (checked on MODSCAN) As of now we are closing this complaint. | data not coming till the local system (checked on MODSCAN) As of now we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH23063705115 | 2023-07-28 | 2023-07-31 | Cust 1483 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | ph transmitter not getting calibrate for sensor. | visited site and done calibration on simulator found working normal.but on ph sensor found not responding, so suggested to send ph transmitter to pune for repairing. | visited site and done calibration on simulator found working normal.but on ph sensor found not responding, so suggested to send ph transmitter to pune for repairing. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH23065955116 | 2023-07-31 | 2023-08-03 | Cust 608 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc and current output card need to be change. | done amc and changed current output card found working ok. now other current output card and other spares are required. FMPL will provide offer for the same. | done amc and changed current output card found working ok. now other current output card and other spares are required. FMPL will provide offer for the same. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVH23066325117 | 2023-07-31 | 2023-08-07 | Cust 1484 | VASHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | visit requested to check the status of transmitters installed in DM water | visited site to check and discuss the status of pH and conductivity transmitters . Requested customer to remove the transmitters and keep in lab , we will visit and check . | visited site to check and discuss the status of pH and conductivity transmitters . Requested customer to remove the transmitters and keep in lab , we will visit and check . | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH23067975118 | 2023-08-03 | 2023-08-08 | Cust 1263 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC work | visited site and found ph1 signals were showing 0 , so FMPL will send offer for spares required. | visited site and found ph1 signals were showing 0 , so FMPL will send offer for spares required. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH23067985119 | 2023-08-03 | 2023-08-07 | Cust 1337 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC work | visited site and done amc work and checked found working normal. | visited site and done amc work and checked found working normal. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVH23068045120 | 2023-08-03 | 2023-08-07 | Cust 554 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Smartpro Electronics Assembly(Smartpro-Elc-cbl) | error in pH readings | found pH electrode broken , requested customer replace electrode with temperature range of 80deg c continues application. | found pH electrode broken , requested customer replace electrode with temperature range of 80deg c continues application. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH23070565121 | 2023-08-08 | 2023-08-12 | Cust 671 | VASHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | commissioning | commissioning of 3 analyser done and 1 analyser xenon power supply card need to arrange and same will be arrange. | commissioning of 3 analyser done and 1 analyser xenon power supply card need to arrange and same will be arrange. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH23073925122 | 2023-08-17 | 2023-08-22 | Cust 1485 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc | attended both analysers AMC work done | attended both analysers AMC work done | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVH23075845123 | 2023-08-20 | 2023-08-25 | Cust 1484 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | error in readings | checked smartpro 8966 and 8967 transmitters of DM water plant | checked smartpro 8966 and 8967 transmitters of DM water plant | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH23076835124 | 2023-08-21 | 2023-08-27 | Cust 1486 | VASHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | recommissioning after repairs. | visited site and done recommissioning and checked found working normal. | visited site and done recommissioning and checked found working normal. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH23078635125 | 2023-08-25 | 2023-08-30 | Cust 1478 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc call | amc call done | amc work done | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVH23082795126 | 2023-09-01 | 2023-09-08 | Cust 1487 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Conductivity Transmitter Aquacon | Aquacon 2000(AQCN 2000 p) | startup high current | RF No 138231787 has been approved by GILL VIRENDRA. OA Number 5524800766 generated | RF No 138231787 has been approved by GILL VIRENDRA. OA Number 5524800766 generated | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFVH23083075127 | 2023-09-01 | 2023-09-04 | Cust 1401 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue. | We took the access to the PC and did some changes from configuration end and now data is matching to the CPCB server with actual value. | We took the access to the PC and did some changes from configuration end and now data is matching to the CPCB server with actual value. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH23084205128 | 2023-09-04 | 2023-09-11 | Cust 1482 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | DO Transmitters | VISI pro DO(VISI pro DO) | error in readings | checked Aquamax visiwater two nos , found sensor not getting calibrated. carried both sensors for further checking and if can cover under warranty. | checked Aquamax visiwater two nos , found sensor not getting calibrated. carried both sensors for further checking and if can cover under warranty. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH23085265129 | 2023-09-07 | 2023-09-08 | Cust 1263 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data upload issue. | After restarted enviroconnect RTDAS service data is live on MPCB server. | After restarted enviroconnect RTDAS service data is live on MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH23086625130 | 2023-09-11 | 2023-09-17 | Cust 1488 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | OIL ANALYSER | Current output card (1) | current output not coming. | visited site and checked current output not coming and send to pune for repairs. | visited site and checked current output not coming and send to pune for repairs. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH23087385131 | 2023-09-12 | 2023-09-18 | Cust 1489 | VASHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Actual visit was for aqua2trans field trial meter | checked pH trial transmitter , replaced pH electrode for smartpro | checked pH trial transmitter , replaced pH electrode for smartpro | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH23087425132 | 2023-09-12 | 2023-09-18 | Cust 1484 | VASHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Aqua4Trans | Aqua4 trans-A4TW033000(1) | error in pH readings | done calibration with pH buffer stds | done calibration with pH buffer stds | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH23087745133 | 2023-09-12 | 2023-09-18 | Cust 1490 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc | amc work done | amc work done | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVH23089005134 | 2023-09-15 | 2023-09-18 | Cust 1482 | VASHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | Aquamax system(1) | actual call was to install aqua4trans DO2 transmitter and oxysens sensor for filed trials | trial transmitter is installed . | trial transmitter is installed . | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH23091185135 | 2023-09-20 | 2023-10-25 | Cust 1491 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | AQUADO Analyzer | Visiwater DO sensor(1) | Not responding , showing 0 PPM | Replacement material has been booked through TCI EXPRESS vide LR No - 069637330 LR Dt - 13-OCT-2023 | Replacement material has been booked through TCI EXPRESS vide LR No - 069637330 LR Dt - 13-OCT-2023 | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFVH23092675136 | 2023-09-24 | 2023-09-28 | Cust 1492 | VASHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Calibration work | done calibration work and check found working normal. | done calibration work and check found working normal. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFVH23092685137 | 2023-09-24 | 2023-09-28 | Cust 1481 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Routine maintenance | done routine maintanance work and checked found working normal. | done routine maintanance work and checked found working normal. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVH23094755138 | 2023-09-27 | 2023-09-30 | Cust 1493 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | TOC ANALYSER | FM 101 TOC Analyser(FM 101 TOC) | Spares need to be change and recomissioning of same. | visited site and changed spares and found Co2 detector had problem. FMPL will submit offer for the same. | visited site and changed spares and found Co2 detector had problem. FMPL will submit offer for the same. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH23095665139 | 2023-09-28 | 2023-09-30 | Cust 526 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Electrode | ORP Electrode- Polilyte RX 120(1) | commissioning of new orp sensor. | visited site and found PH and ORP transmitter problem and suggested to send pune for repairs. DO transmitters checked found working normal. | visited site and found PH and ORP transmitter problem and suggested to send pune for repairs. DO transmitters checked found working normal. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH23099345140 | 2023-10-03 | 2023-10-09 | Cust 1493 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | Recommissioming after repairs / spares changed. | done recommissioning and checked found working normal. | done recommissioning and checked found working normal. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH23100705141 | 2023-10-06 | 2023-10-12 | Cust 1263 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation needs to be done. | As of now we are closing this complaint, we will reopen it after getting system access from customers end. | As of now we are closing this complaint, we will reopen it after getting system access from customers end. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH23101955142 | 2023-10-09 | 2023-10-11 | Cust 1337 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation needs to be done. | We took the access to the system and successfully did the software reinstallation + local + MPCB configuration and now data is liv on both server. | We took the access to the system and successfully did the software reinstallation + local + MPCB configuration and now data is liv on both server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH23102075143 | 2023-10-09 | 2023-10-12 | Cust 1337 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | analyser is in breakdown and got reset. | hardware reset done and checked found working normal. | hardware reset done and checked found working normal. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH23102085144 | 2023-10-09 | 2023-10-12 | Cust 1263 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | spares need to be changed to make analyser operational. | visited site and changed the spares and checked found working normal. | visited site and changed the spares and checked found working normal. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH23102105145 | 2023-10-09 | 2023-10-16 | Cust 1494 | VASHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | commissioning | commissioning done and checked found working normal. | commissioning done and checked found working normal. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH23102125146 | 2023-10-09 | 2023-10-15 | Cust 1478 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc | amc done | amc work done | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVH23103605147 | 2023-10-13 | 2023-10-18 | Cust 1491 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | AQUADO Analyzer | Visiwater DO sensor(1) | error in readings | replaced sensor membrane caps for two sensors . | replaced sensor membrane caps for two sensors . | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH23103645148 | 2023-10-13 | 2023-10-13 | Cust 1492 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB Server Connectivity Problem | We took access to the system and data is live on MPCB | We took access to the system and data is live on MPCB | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH23105565149 | 2023-10-17 | 2023-10-20 | Cust 1337 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | mpcb data upload issue. | We did some changes from hardware side and now data is matching on local + MPCB server with actual value. | We did some changes from hardware side and now data is matching on local + MPCB server with actual value. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH23111445150 | 2023-10-30 | 2023-11-05 | Cust 1485 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | error in reading | found issue with calibration of COD parameter , recalibrated with KHP std. | found issue with calibration of COD parameter , recalibrated with KHP std. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVH23112855151 | 2023-10-31 | 2023-11-05 | Cust 1492 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Aqua4Trans | Aqua4Trans(1) | pH reading drifting | Found pH readings ok , kept under observation | Found pH readings ok , kept under observation | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH23115395152 | 2023-11-03 | 2023-11-03 | Cust 1479 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data upload issue. | here is the hardware side issue after correcting the same now data is live on MPCB + Local server | here is the hardware side issue after correcting the same now data is live on MPCB + Local server | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH23116515153 | 2023-11-06 | 2023-11-08 | Cust 1263 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB + MPCB data upload issue. | Here is hardware side issue, after correcting the same data is now live on CPCB + MPCB server. | Here is hardware side issue, after correcting the same data is now live on CPCB + MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH23126565154 | 2023-11-27 | 2023-12-01 | Cust 1495 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Conductivity Transmitter Aquacon | Aquacon 2000(AQCN 2000 p) | conductivity readings are fluctuating | visited site and found condutivity sensor and transmitter problem and so suggested to send to pune . | visited site and found condutivity sensor and transmitter problem and so suggested to send to pune . | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH23127585155 | 2023-11-29 | 2023-12-04 | Cust 1496 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Software installation (UBUNTU). CPCB + MPCB configuration needs to be done. | We took the access to your system and did local + CPCB configuration in the UBUNTU system. and now data is live on the CPCB server. | We took the access to your system and did local + CPCB configuration in the UBUNTU system. and now data is live on the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH23129895156 | 2023-12-01 | 2023-12-07 | Cust 1491 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(Ph-Trans-Smartpro) | no display | repaired pH and orp meters are installed . | repaired pH and orp meters are installed . | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH23131325157 | 2023-12-04 | 2023-12-10 | Cust 1329 | VASHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Aqua4Trans | Aqua4Trans(1) | installation of RTU for aqua4trans trial DO2 transmitter | new RTU is installed , trial unit readings are live | new RTU is installed , trial unit readings are live | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH23131335158 | 2023-12-04 | 2023-12-10 | Cust 1489 | VASHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | pH Transmitter Aquamon | Aquamon 2000 plus(AQ2000 plus) | actual visit is for aqua2trans pH transmitter trial meter | trail unit and smartpro parallel meters checked , report submitted to r&d | trail unit and smartpro parallel meters checked , report submitted to r&d | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH23131705159 | 2023-12-05 | 2023-12-10 | Cust 1479 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | routine maintanance | done routine maintanance work and checked found ok. | done routine maintanance work and checked found ok. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVH23132325160 | 2023-12-07 | 2023-12-12 | Cust 1479 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc work | done amc work and checked found working normal. | done amc work and checked found working normal. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFVH23132335161 | 2023-12-07 | 2023-12-12 | Cust 1479 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc | done amc work and checked found working normal. | done amc work and checked found working normal. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFVH23132345162 | 2023-12-07 | 2023-12-12 | Cust 1479 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc work | done amc work and checked found working normal. | done amc work and checked found working normal. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFVH23134355163 | 2023-12-11 | 2023-12-13 | Cust 1263 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data upload issue. | now the issue has been resolved and totalizer flow value is correctly reflecting on the MPCB server. | now the issue has been resolved and totalizer flow value is correctly reflecting on the MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH23135075164 | 2023-12-13 | 2023-12-15 | Cust 578 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data upload issue. | we took the access to your system and did software reinstallation + local + CPCB + MPCB configuration, now data is live on all the servers. | we took the access to your system and did software reinstallation + local + CPCB + MPCB configuration, now data is live on all the servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH23136305165 | 2023-12-14 | 2023-12-18 | Cust 1481 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc work. | visited site and done amc work and checked found working normal. | visited site and done amc work and checked found working normal. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVH23138365166 | 2023-12-19 | 2023-12-24 | Cust 1490 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc | amc work done | amc work done | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVH23138755167 | 2023-12-19 | 2023-12-22 | Cust 1337 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Display board configuration needs to be done. | We successfully did display board configuration. | We successfully did display board configuration. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH23139195168 | 2023-12-20 | 2023-12-24 | Cust 1478 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc | amc work done | amc work done | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVH23139575169 | 2023-12-20 | 2023-12-23 | Cust 1337 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | routine maintanace | done routine maintanance and checked found working normal. | done routine maintanance and checked found working normal. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVH23139695170 | 2023-12-21 | 2023-12-24 | Cust 1478 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc | amc work done | amc work done | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVH23141385171 | 2023-12-25 | 2024-01-29 | Cust 1491 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | AQUADO Analyzer | AQUAMAX DOT(1) | visiwater PP sensors not working satifactory , need to check at one location with SS visiwater sensor. | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210078413 LR Dt - 12-JAN-2024 | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210078413 LR Dt - 12-JAN-2024 | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFVH23142445172 | 2023-12-26 | 2023-12-31 | Cust 1401 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc | visited site and done amc and checked found working normal. | visited site and done amc and checked found working normal. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVH23144315173 | 2023-12-29 | 2023-12-30 | Cust 1263 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Totalizer data upload issue. | Now data is matching to the actual reading | Now data is matching to the actual reading | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH23147985174 | 2024-01-03 | 2024-01-05 | Cust 1263 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data mismatch issue | We took access to the system and check the data on MODSCAN, and as per calculation s same data is uploading onto the CPCB + MPCB server for flow, hence we are closing this complaint. | We took access to the system and check the data on MODSCAN, and as per calculation s same data is uploading onto the CPCB + MPCB server for flow, hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH23148825175 | 2024-01-05 | 2024-01-10 | Cust 1497 | VASHI | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | Aqua4Trans | Aqua4 trans-A4TW033000(1) | actual visit is for aqua2trans trial unit | checked aqua2trans pH meter | checked aqua2trans pH meter | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH23149975176 | 2024-01-08 | 2024-01-12 | Cust 564 | VASHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | CALIBRATION WORK | visited site and done calibration work found ok. ph sensor , mphdo card and all tubes need to be changed. | visited site and done calibration work found ok. ph sensor , mphdo card and all tubes need to be changed. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFVH23151235177 | 2024-01-09 | 2024-01-11 | Cust 1498 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB connectivity issue | Data is now live on the MPCB server. | Data is now live on the MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH23153055178 | 2024-01-12 | 2024-01-18 | Cust 608 | VASHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | recommissioing after repairs. | visited site and done recommissioning and checked found working normal. | visited site and done recommissioning and checked found working normal. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH23153225179 | 2024-01-14 | 2024-01-19 | Cust 1485 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc | amc work done | amc work done | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVH23153895180 | 2024-01-15 | 2024-01-17 | Cust 608 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data uploading issue. | Data is live on the MPCB server. | Data is live on the MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH23154135181 | 2024-01-15 | 2024-01-18 | Cust 1263 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data upload issue. | As of now we are closing this complaint, and will reopen it after the hardware side issue is resolved. | As of now we are closing this complaint, and will reopen it after the hardware side issue is resolved. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH23154445182 | 2024-01-16 | 2024-01-21 | Cust 1484 | VASHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Aqua4Trans | Aqua4 trans-A4TW033000(1) | TSS meter readings showing high | attended tss,pH and conductivity aqua4trans transmitters | attended tss,pH and conductivity aqua4trans transmitters | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH23154855183 | 2024-01-16 | 2024-01-19 | Cust 1499 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Reinstallation of software. | We took access to the system and successfully completed the installation and reconfiguration of software now data is live on LOCAL+MPCB | We took access to the system and successfully completed the installation and reconfiguration of software now data is live on LOCAL+MPCB | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVH23155095184 | 2024-01-17 | 2024-01-22 | Cust 1500 | VASHI | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | Aqua4Trans | Aqua4Trans(1) | trails | visited site along with selec engineer and amol sathe , required logic is programmed in controller and tested by simulation from flow meter and conductivity transmitter . logic found ok. | visited site along with selec engineer and amol sathe , required logic is programmed in controller and tested by simulation from flow meter and conductivity transmitter . logic found ok. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH23155675185 | 2024-01-18 | 2024-01-23 | Cust 1501 | VASHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | commissioning | visited site and checked ph,orp and conductivity transmitter found ok but sensor is not responding and need to change . | visited site and checked ph,orp and conductivity transmitter found ok but sensor is not responding and need to change . | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH23156105186 | 2024-01-19 | 2024-01-25 | Cust 1329 | VASHI | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | Aqua4Trans | Aqua4Trans(1) | actual visit to check aqua4trans do2 trail unit | checked modbus output of visiwater for connectivity with RTU . provided details to r&d | checked modbus output of visiwater for connectivity with RTU . provided details to r&d | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH23157595187 | 2024-01-22 | 2024-01-29 | Cust 1491 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | AQUADO Analyzer | Visiferm DO sensor(1) | replace visiwater plastic with SS visiwater DO Arc 120 FC10 | new visiwater is installed on tank 2 and kept under observation | new visiwater is installed on tank 2 and kept under observation | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH23158645188 | 2024-01-24 | 2024-01-30 | Cust 547 | VASHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Silica Analyser | Silica 8891(8891) | commissioning for Silica Analyzer Tag No. 21-LE-2101-AT-1606. OA - 5523102408 | as there is no chemicals and regaents avaible onsite so commissioning able to perform, so division will arrange reagents and chemicals onsite and then will attend site for commissioning. | as there is no chemicals and regaents avaible onsite so commissioning able to perform, so division will arrange reagents and chemicals onsite and then will attend site for commissioning. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH23165875189 | 2024-02-05 | 2024-02-15 | Cust 1263 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | DATCON | DATCON(DATCON) | CPCB + MPCB configuration needs to be done in DATCon. | We have completed CPCB + MPCB configuration in DATCon and now data is live on both the servers. | We have completed CPCB + MPCB configuration in DATCon and now data is live on both the servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH23166085190 | 2024-02-05 | 2024-02-10 | Cust 1493 | VASHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | commissioning | commissioning done and checked found working normal. | commissioning done and checked found working normal. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH23167875191 | 2024-02-09 | 2024-02-16 | Cust 1490 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | issue with tss and pH readings | attended analyser , amc work done | attended analyser , amc work done | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVH23172105192 | 2024-02-18 | 2024-02-25 | Cust 1485 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | installed spares received | installed pH electrode , MIL card tested analsyer now ok. | installed pH electrode , MIL card tested analsyer now ok. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH23172665193 | 2024-02-19 | 2024-02-25 | Cust 1481 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | routine maintanance | as customer deny for amc visit, so will visit in march. | as customer deny for amc visit, so will visit in march. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVH23173145194 | 2024-02-20 | 2024-02-23 | Cust 1498 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not updating on the local and MPCB server. | Data is now live on Local and MPCB servers. | Data is now live on Local and MPCB servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH23173385195 | 2024-02-20 | 2024-02-25 | Cust 1502 | VASHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | Aquamax system(1) | commissioning | commissioning done | commissioning done | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH23180025196 | 2024-02-29 | 2024-04-02 | Cust 1503 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Electrode | FM1106(1) | Sensor failed within short period of time for standard application. Sensor was not responding. Site Contact- Mithun Darekar Sr Manager ? Instrument Maintenance Plant Office: Plot K-1, MIDC Industrial Area, Taloja-410208, Dist: Raigad www.dfpcl.com T: +91 22 50684022 | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210091809 LR Dt - 28-MAR-2024 | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210091809 LR Dt - 28-MAR-2024 | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFVH23180025196 | 2024-02-29 | 2024-04-02 | Cust 1503 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Electrode | FM1106(1) | Sensor failed within short period of time for standard application. Sensor was not responding. Site Contact- Mithun Darekar Sr Manager ? Instrument Maintenance Plant Office: Plot K-1, MIDC Industrial Area, Taloja-410208, Dist: Raigad www.dfpcl.com T: +91 22 50684022 | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210091809 LR Dt - 28-MAR-2024 | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210091809 LR Dt - 28-MAR-2024 | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFVH23181475197 | 2024-03-04 | 2024-03-09 | Cust 1493 | VASHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | commissioning | visited site and commissioning done and checked found working normal. | visited site and commissioning done and checked found working normal. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH23184295198 | 2024-03-11 | 2024-03-14 | Cust 1485 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data mismatch issue. | As of now we are closing this complaint, and will reopen it after hardware side issue is solved | As of now we are closing this complaint, and will reopen it after hardware side issue is solved | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH23187195199 | 2024-03-17 | 2024-03-23 | Cust 1478 | VASHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | DATCON | DATCON(DATCON) | actual installation of RTU301 with CX1000 | new RTU panel installation and commissioning is done. | new RTU panel installation and commissioning is done. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH23190225200 | 2024-03-20 | 2024-03-25 | Cust 1401 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc call | amc visit done and found working normal. | amc visit done and found working normal. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVH23190235201 | 2024-03-20 | 2024-03-25 | Cust 1498 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | PH1 filter(1) | ph1 signals low. | visited site and change the flowcell and checked found working satisfactory. | visited site and change the flowcell and checked found working satisfactory. | NO | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH23192095202 | 2024-03-25 | 2024-03-29 | Cust 1337 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc visit | visited site and done amc work. | visited site and done amc work. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFVH23192115203 | 2024-03-25 | 2024-03-29 | Cust 1481 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc visit | done amc work and checked found working normal. | done amc work and checked found working normal. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFVH23192985204 | 2024-03-26 | 2024-03-29 | Cust 1504 | VASHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Transmitter Aquacon | Aquacon 2000(AQCN 2000 p) | Calibration of 7 Nos Aquacon 2000 Pro Models for Water | Calibration done, MOM attached. | Calibration done, MOM attached. | NO | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFVH23195735205 | 2024-03-29 | 2024-03-29 | Cust 1479 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB connectivity needs to be done. | We took the access to your system and completed local + CPCB + MPCB connectivity and now data is live on all the servers. | We took the access to your system and completed local + CPCB + MPCB connectivity and now data is live on all the servers. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH24001025206 | 2024-04-02 | 2024-04-05 | Cust 1481 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue. | we took the access to your system and checked the data on MODSCAN, and as per calculations the same data uploading onto the CPCB server. | we took the access to your system and checked the data on MODSCAN, and as per calculations the same data uploading onto the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH24001265207 | 2024-04-03 | 2024-04-10 | Cust 1490 | VASHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Display card of Aqua4Trans(1) | commsioning | commissioning done | commissioning done | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH24001845208 | 2024-04-04 | 2024-04-10 | Cust 547 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Silica Analyser | Silica 8891(8891) | Problem in the commissioned silica analyser, The Reagent ONE Pump is not coming into operation. | visited site and found reagent 1 motor is not working and need to be changed. | visited site and found reagent 1 motor is not working and need to be changed. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH24002035209 | 2024-04-04 | 2024-04-08 | Cust 1493 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | TOC ANALYSER | Permeation Dryer Tube(001?001) | CO2 detector , PSU PCB with complete leads, permeation tube need to be replace with new spare. | visited site and changed the spares with new and checked found working normal. | visited site and changed the spares with new and checked found working normal. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH24002685210 | 2024-04-05 | 2024-04-08 | Cust 1479 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB connectivity needs to be done | we took access to your system and completed CPCB connectivity and now data is live on CPCB server. | we took access to your system and completed CPCB connectivity and now data is live on CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH24002695211 | 2024-04-05 | 2024-04-08 | Cust 1479 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB connectivity needs to be done. | we took access to your system and completed CPCB connectivity and now data is live on CPCB server. | we took access to your system and completed CPCB connectivity and now data is live on CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH24004465212 | 2024-04-10 | 2024-04-14 | Cust 1481 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | FM make current output card(1) | current output calibration out. PH readings problem | visited site and found current calibration was distrub and same done but it is getting out again and so suggested to change two current output cards. | visited site and found current calibration was distrub and same done but it is getting out again and so suggested to change two current output cards. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH24005465213 | 2024-04-12 | 2024-04-12 | Cust 1481 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB + MPCB data mismatch issue. | We took access to your system, and checked the data on MODSCAN, and as per calculations the same data is uploading onto the CPCB + MPCB + Local server. | We took access to your system, and checked the data on MODSCAN, and as per calculations the same data is uploading onto the CPCB + MPCB + Local server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH24008265214 | 2024-04-19 | 2024-04-19 | Cust 1485 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | data was not uploading. | Data is now live on MPCB server. | Now data is live on MPCB server | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVH24008445215 | 2024-04-19 | 2024-04-23 | Cust 1492 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | 1st amc visit | done amc call. | done amc call. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVH24010635216 | 2024-04-24 | 2024-05-31 | Cust 1479 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data upload issue. | Now data is live on MPCB server for all the parameters | Now data is live on MPCB server for all the parameters | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH24011435217 | 2024-04-25 | 2024-04-29 | Cust 1490 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Chlorine sensor | CHLORINE SENSOR(1) | error in reading | calibrated chlorine and turbidity analyser | calibrated chlorine and turbidity analyser | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH24013645218 | 2024-04-29 | 2024-05-15 | Cust 1493 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | error in reading | visited site and changed the optical filter and PM board and checked found working normal. | visited site and changed the optical filter and PM board and checked found working normal. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH24016735219 | 2024-05-03 | 2024-05-08 | Cust 1489 | VASHI | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | Aqua4Trans | Aqua4Trans(1) | actual call for aqua2trans trail unit | checked calibration of trail unit and smartpro 8966 | checked calibration of trail unit and smartpro 8966 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH24017025220 | 2024-05-05 | 2024-05-11 | Cust 867 | VASHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Silica Analyser | Silica 8891(8891) | spares chamging to upgrade to higher range. | all spares except water reservior is not installed. wipro need to arrange DM water line till analyser. After that FM will visit site. | all spares except water reservior is not installed. wipro need to arrange DM water line till analyser. After that FM will visit site. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH24017255221 | 2024-05-06 | 2024-05-08 | Cust 1505 | VASHI | BHD- Boiler House Division CHAKAN | CLOSED | MANUFACTURING ISSUES | ABCO (Automatic Blowdown Controller ) Skid | ABCO (Automatic Blowdown Controller ) Skid(ABCO) | As discussed, we are facing issues with operation of CCDS. After Conductivity increases above setpoint, relay -1 is energised, but relay-4 (which operates On/Off valves) is not getting energised. Same observed in both CCDS. Final Observation - Jumper Setting Incorrect. Separate Mail is dropped to Divison. | Mail attached. | Mail attached. | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFVH24018365222 | 2024-05-07 | 2024-05-12 | Cust 1506 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | sample pump problem | changed the sample pump motor and pcx2 board and checked found working normal. | changed the sample pump motor and pcx2 board and checked found working normal. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH24018505223 | 2024-05-08 | 2024-05-10 | Cust 617 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | change in permissible range in Local server. | We took access to the system and changed the permissible range in Local server. | We took access to the system and changed the permissible range in Local server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVH24018685224 | 2024-05-08 | 2024-05-13 | Cust 1480 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | mains SMPS failed | recommissioning call | recommissioning done | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH24021115225 | 2024-05-15 | 2024-05-17 | Cust 1263 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB/MPCB data mismatch issue. | We took access to your system, and checked the data on MODSCAN, and as per calculations the same data is uploading onto the CPCB server. | We took access to your system, and checked the data on MODSCAN, and as per calculations the same data is uploading onto the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH24021535226 | 2024-05-15 | 2024-05-20 | Cust 1482 | VASHI | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | AQUADO Analyzer | Aquamax system(1) | actual visit was for demo do2 transmitter at CETP Roha | attended RTU | attended RTU | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH24021595227 | 2024-05-16 | 2024-05-23 | Cust 867 | VASHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Silica Analyser | Silica 9210(9210) | 3S silica analyser recommissioning with spares required to upgrade the range from 1ppm to 25ppm. | Visited site and signals of absorbance zero, so taken detector at chakan factory and check on another analyser and found detector problem. So need to change the detector. | Visited site and signals of absorbance zero, so taken detector at chakan factory and check on another analyser and found detector problem. So need to change the detector. | EXTENDED | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH24021965228 | 2024-05-16 | 2024-05-21 | Cust 1478 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc | amc work done | amc work done | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVH24022005229 | 2024-05-16 | 2024-05-17 | Cust 617 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Configuration of Flow on CPCB Server. | Now data is live on LOCAL+CPCB+MPCB. | Now data is live on LOCAL+CPCB+MPCB. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVH24023015230 | 2024-05-20 | 2024-05-23 | Cust 1480 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation needs to be done. | we took access to your system and completed Software reinstallation + local connectivity. And now data is live on the Local server | we took access to your system and completed Software reinstallation + local connectivity. And now data is live on the Local server | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH24023825231 | 2024-05-20 | 2024-05-25 | Cust 1478 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc | amc | amc work done | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVH24024805232 | 2024-05-22 | 2024-05-27 | Cust 1401 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | LCD of Multiparameter Display(1) | display blank | visited site and changed the display with new one and checked found working normal. | visited site and changed the display with new one and checked found working normal. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH24024815233 | 2024-05-22 | 2024-05-27 | Cust 1481 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | PH Input Card for Multiparameter(AWA-PH) | mphdo card, ph sensor need to be change with new one | visited site and install all the spares and checked found working normal. | visited site and install all the spares and checked found working normal. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH24027535234 | 2024-05-27 | 2024-05-31 | Cust 1494 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | display blank | visited site and checked found display blank and send to pune for further repairs. | visited site and checked found display blank and send to pune for further repairs. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH24027545235 | 2024-05-27 | 2024-06-03 | Cust 1490 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc | amc work done | amc work done | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVH24028935236 | 2024-05-29 | 2024-06-03 | Cust 1507 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | readings not communicating with MPCB server | explained customer , issue is with scan convertor . he will have to check with connectivity service provider | explained customer , issue is with scan convertor . he will have to check with connectivity service provider | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH24029335237 | 2024-05-30 | 2024-05-31 | Cust 617 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Localhost is not accessible. | We took access to the system and started the services and now localhost is accessible. | We took access to the system and started the services and now localhost is accessible. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVH24029805238 | 2024-05-30 | 2024-06-03 | Cust 1479 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc call | done amc call and found working normal. | done amc call and found working normal. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVH24029825239 | 2024-05-30 | 2024-06-03 | Cust 1479 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc call | done amc call and found working normal. | done amc call and found working normal. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVH24032825240 | 2024-06-04 | 2024-06-23 | Cust 1484 | VASHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4 trans-A4TW033000(1) | commissioning of new sensor pH and turbidity | found issue with turbidity transmitter after installing new sensor . Closing this now as division suggested to install new version of software in transmitter . | found issue with turbidity transmitter after installing new sensor . Closing this now as division suggested to install new version of software in transmitter . | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH24034855241 | 2024-06-10 | 2024-06-17 | Cust 1329 | VASHI | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | Aqua4Trans | Aqua4Trans(1) | visit site for demo aqua4trans DO2 transmitter | checked the analyzer performance | checked the analyzer performance | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH24035265242 | 2024-06-10 | 2024-06-11 | Cust 1480 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on CPCB + Local + MPCB. | Updated RTDas version of software, and now local + MPCB data is matching with the actual readings. And now data is live on CPCB + Local + MPCB server. | Updated RTDas version of software, and now local + MPCB data is matching with the actual readings. And now data is live on CPCB + Local + MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVH24036995243 | 2024-06-14 | 2024-06-20 | Cust 1478 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc | issue with TSS reading , found sample tube damaged , replaced the tube . | issue with TSS reading , found sample tube damaged , replaced the tube . | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVH24040395244 | 2024-06-20 | 2024-06-27 | Cust 1490 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Aqua4Trans | 4-20 mA input card of Aqua4trans(1) | not responding to change in chlorine content in sample | replaced input card , now kept under observation | replaced input card , now kept under observation | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH24040775245 | 2024-06-21 | 2024-06-24 | Cust 1337 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc call | done amc call and checked found working normal. | done amc call and checked found working normal. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVH24040795246 | 2024-06-21 | 2024-06-24 | Cust 1263 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | pH Transmitter Aquamon | Aquamon 2000 plus(AQ2000 plus) | calibration | visited site and done calibration work and checked found working normal. | visited site and done calibration work and checked found working normal. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH24048415247 | 2024-07-04 | 2024-07-11 | Cust 1494 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | display issue , visited site for installation of new display | spare sample pump motor is arranged on RF , once received will visit site . | spare sample pump motor is arranged on RF , once received will visit site . | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFVH24048885248 | 2024-07-04 | 2024-07-08 | Cust 1481 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc call | done amc call and chceked found working normal. | done amc call and chceked found working normal. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVH24052595249 | 2024-07-12 | 2024-07-14 | Cust 1508 | VASHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | commissioning | visited site and done commissioning and checked found working normal. | visited site and done commissioning and checked found working normal | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH24054585250 | 2024-07-16 | 2024-07-21 | Cust 1493 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | OIL ANALYSER | PCX2 Power Board(PCX2) | oil in water spares need to be changed. | visited site and changed the spares and found working normal. | visited site and changed the spares and found working normal. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH24054865251 | 2024-07-17 | 2024-07-21 | Cust 1478 | VASHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | isolators installation | installed isolators in CETP analyser | installed isolators in CETP analyser | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH24056935252 | 2024-07-21 | 2024-07-25 | Cust 1263 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc call | visited site and done amc call and checked found working normal. | visited site and done amc call and checked found working normal. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVH24057905253 | 2024-07-22 | 2024-07-28 | Cust 1493 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | SULPHIDE ANALYSER | Pulse pump(695=114=000) | HCL pump not working. | visited site and changed the HCL pump and tygon tubing and checked found analyser found working normal. | visited site and changed the HCL pump and tygon tubing and checked found analyser found working normal. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH24059185254 | 2024-07-25 | 2024-07-28 | Cust 1494 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | sample pump motor not working | new motor is installed . | new motor is installed . | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH24061125255 | 2024-07-28 | 2024-07-30 | Cust 1493 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | routine amc | visited site and done routine amc and checked found working normal. | visited site and done routine amc and checked found working normal. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVH24061155256 | 2024-07-28 | 2024-07-31 | Cust 1401 | VASHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | AQUADO Analyzer | DO2 Arc Sensors(DO2 Arc Sensors) | calibration work | visited site and done calibration work and checked found working normal. | visited site and done calibration work and checked found working normal. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFVH24062155257 | 2024-07-29 | 2024-07-31 | Cust 1479 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB + RSPCB data upload issue. | Now Hardware side issue is resolved and data is live on CPCB + MPCB server. | Now Hardware side issue is resolved and data is live on CPCB + MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH24063185258 | 2024-07-30 | 2024-08-05 | Cust 1484 | VASHI | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | pH Transmitter Aquamon | Aquamon 2000 plus(AQ2000 plus) | actual visit is for aqua2trans trial unit , installation of new card along with r&d and csd | new aqua2trans transmitter is installed | new aqua2trans transmitter is installed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH24064275259 | 2024-07-31 | 2024-08-01 | Cust 1498 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | TEPS server connectivity issue. | As of now we are closing this complaint. we will reopen this once the TEPS server issue is resolved. | As of now we are closing this complaint. we will reopen this once the TEPS server issue is resolved. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVH24066475260 | 2024-08-05 | 2024-08-08 | Cust 1480 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data upload issue. | we took access of your system and now the issue has been resolved and data is live on the MPCB server. | we took access of your system and now the issue has been resolved and data is live on the MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH24066765261 | 2024-08-05 | 2024-08-09 | Cust 1509 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Conductivity Transmitter Aquacon | Aquacon Electronics Assembly(CON-ELE-set-2000) | readings problem. | visited site and done calibration work and checked found working normal. | visited site and done calibration work and checked found working normal. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH24067175262 | 2024-08-06 | 2024-08-12 | Cust 1484 | VASHI | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | Aqua4Trans | Aqua4 trans-A4TW033000(1) | visit planned to check aqua2trans demo meter | visited site to check the performance of aqau2trans trail unit | visited site to check the performance of aqau2trans trail unit | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH24067675263 | 2024-08-07 | 2024-08-09 | Cust 1401 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc call | visited site and done amc call and checked found working normal. | visited site and done amc call and checked found working normal. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVH24067695264 | 2024-08-07 | 2024-08-12 | Cust 1490 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc | amc work done | amc work done | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVH24070375265 | 2024-08-12 | 2024-08-15 | Cust 1492 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | routine maintanance | visited site and done commissioning and checked found working normal. | visited site and done commissioning and checked found working normal. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVH24070385266 | 2024-08-12 | 2024-08-15 | Cust 1492 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC call | visited site and found multiple cards problem and tubing damaged and so suggested to send analyser chakan for repairs. | visited site and found multiple cards problem and tubing damaged and so suggested to send analyser chakan for repairs. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVH24071225267 | 2024-08-14 | 2024-08-19 | Cust 1501 | VASHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | OXYFERM(OXYFERM) | commissioning | commissioning completed | commissioning completed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH24071365268 | 2024-08-15 | 2024-08-18 | Cust 564 | VASHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | commissioning | visited site and done commissioning and checked found working normal. | visited site and done commissioning and checked found working normal. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH24071615269 | 2024-08-15 | 2024-08-18 | Cust 564 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | spares to be changed with new one | visited site and changed spares and checked found ok. | visited site and changed spares and checked found ok. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH24074385270 | 2024-08-21 | 2024-08-22 | Cust 1401 | VASHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | AQUADO Analyzer | Visitrace DO sensor(1) | calibration work | visited site and done calibration work and checked found ok. | visited site and done calibration work and checked found ok. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFVH24074735271 | 2024-08-22 | 2024-08-24 | Cust 608 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation. | As of now we are closing this complaint and will reopen it once we get system access. | As of now we are closing this complaint and will reopen it once we get system access. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH24077605272 | 2024-08-27 | 2024-09-01 | Cust 1484 | VASHI | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | Aqua4Trans | Aqua4 trans-A4TW033000(1) | Actual visit is for aqua2tras demo meter along with R&D team | attended site with R&D team | attended site with R&D team | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH24082995273 | 2024-09-03 | 2024-09-06 | Cust 1510 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB connectivity needs to be done. | we took access to your system and completed MPCB connectivity, and now data is live on MPCB server. | we took access to your system and completed MPCB connectivity, and now data is live on MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH24083025274 | 2024-09-03 | 2024-09-06 | Cust 1263 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB Url not accessible. | Shared URL with customer no issue from server side. | Shared URL with customer no issue from server side. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVH24083465275 | 2024-09-04 | 2024-09-06 | Cust 1499 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | EnviroConnect software re-installation. | we have successfully accessed the system , and re-installed the EnviroConnect software, and configured the system. As discussed, customer will provide hardware connection and access on Monday, 9th September. Hence We are closing this complaint for now and will reopen it once we receive the necessary access. | We are closing this complaint for now and will reopen it once we receive the necessary access. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVH24083995276 | 2024-09-05 | 2024-09-06 | Cust 578 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB Server. | There is internet side issue from customer end. Once the issue is resolved we will re-open the complaint. | There is internet side issue from customer end. Once the issue is resolved we will re-open the complaint. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVH24086395277 | 2024-09-10 | 2024-09-15 | Cust 1484 | VASHI | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | Aqua4Trans | Aqua4 trans-A4TW033000(1) | actual visit is for aqua2trans trail unit along with R&D | visited site with R&D persons yashwant and jaykumar | visited site with R&D persons yashwant and jaykumar | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH24086465278 | 2024-09-10 | 2024-09-10 | Cust 1499 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on Local and MPCB servers | We took access to the system and successfully completed the Local + MPCB Configuration. And now data is live on Local and MPCB servers. | We took access to the system and successfully completed the Local + MPCB Configuration. And now data is live on Local and MPCB servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVH24087395279 | 2024-09-12 | 2024-09-13 | Cust 608 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Software reinstallation needs to be done. | As of now we are closing this complaint and will reopen it once customers PC is ready to be configured. | As of now we are closing this complaint and will reopen it once customers PC is ready to be configured. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH24090065280 | 2024-09-18 | 2024-09-20 | Cust 1498 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data is not uploading to MPCB and Local servers. | we are following for AMC payment since long . | we are following for AMC payment since long . | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVH24090625281 | 2024-09-19 | 2024-09-25 | Cust 1494 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | error in readings | found sample tube chocked , cleaned it , now analyser is working ok. | found sample tube chocked , cleaned it , now analyser is working ok. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH24091015282 | 2024-09-20 | 2024-09-20 | Cust 1337 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | LocalHost + MPCB server data upload issue. | We took the access to your system and found that ApacheTOMCAT service is stopped in your local PC.After restarting the same, now data is live on the Local + MPCB server. | We took the access to your system and found that ApacheTOMCAT service is stopped in your local PC.After restarting the same, now data is live on the Local + MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH24091715283 | 2024-09-22 | 2024-09-28 | Cust 1511 | VASHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Commissioning | this location is not under vashi branch | this location is not under vashi branch | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFVH24091715283 | 2024-09-22 | 2024-09-28 | Cust 1511 | VASHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Commissioning | this location is not under vashi branch | this location is not under vashi branch | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFVH24091715283 | 2024-09-22 | 2024-09-28 | Cust 1511 | VASHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Commissioning | this location is not under vashi branch | this location is not under vashi branch | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFVH24091925284 | 2024-09-23 | 2024-09-26 | Cust 1401 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB connectivity needs to be done. | we took access of the black box and completed the CPCB configuration. Now data is live on the CPCB server for Stack No-2 (Dust) & Stack No-2 (Gas). | we took access of the black box and completed the CPCB configuration. Now data is live on the CPCB server for Stack No-2 (Dust) &amp;amp; Stack No-2 (Gas). | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH24093235285 | 2024-09-25 | 2024-10-01 | Cust 1329 | VASHI | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | Aqua4Trans | Aqua4 trans-A4TW033000(1) | proactive visit for aqua4trans do2 transmitter trail unit | visited to check trial meter | trail meter working ok | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH24094335286 | 2024-09-27 | 2024-09-30 | Cust 1263 | VASHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | routine maintanace and calibration work | done routine maintanance work and found working normal. | done routine maintanance work and found working normal. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFVH24094345287 | 2024-09-27 | 2024-09-30 | Cust 1493 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | sample pump motor and PCX2 board problem | visited site and changed the sameple pump and PCX2 board and checked found working normal. | visited site and changed the sameple pump and PCX2 board and checked found working normal. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH24097805288 | 2024-10-01 | 2024-10-06 | Cust 1489 | VASHI | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | Aqua4Trans | Aqua4 trans-A4TW033000(1) | install new aqua2trans conductivity transmitter for field trials | installation of aqua2trans field trial unit conductivity transmitter done | installation of aqua2trans field trial unit conductivity transmitter done | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH24098215289 | 2024-10-03 | 2024-10-04 | Cust 1496 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Ubuntu software reinstallation + CPCB connectivity needs to be done. | we took access your system and completed the CPCB + MPCB configuration, now data is live on CPCB + MPCB server. | we took access your system and completed the CPCB + MPCB configuration, now data is live on CPCB + MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH24098595290 | 2024-10-03 | 2024-10-08 | Cust 547 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Silica Analyser | Silica 8891(8891) | 3S awa silica analyser reagent 1 motor need to be changed with new one | changed the reagent 1 motor and checked found working normal. | changed the reagent 1 motor and checked found working normal. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH24099035291 | 2024-10-04 | 2024-10-08 | Cust 1481 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | routine maintanance | done amc work and checked found working normal. | done amc work and checked found working normal. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFVH24099265292 | 2024-10-06 | 2024-10-09 | Cust 1478 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc | amc | amc work done | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVH24100335293 | 2024-10-08 | 2024-10-09 | Cust 617 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Localhost is not accessible. | We took access to the system and found services were stopped. Now localhost is accessible. | We took access to the system and found services were stopped. Now localhost is accessible. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVH24100595294 | 2024-10-09 | 2024-10-14 | Cust 867 | VASHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Silica Analyser | Silica 8891(8891) | AWa silica analyser commissioning | visited site and changed the required spares and checked found working normal. | visited site and changed the required spares and checked found working normal. | EXTENDED | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH24100995295 | 2024-10-09 | 2024-10-14 | Cust 617 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data breaking on CPCB Server. | We took access to the system and updated the DAS version and enable the FTP Port. Now data is live on CPCB server. | We took access to the system and updated the DAS version and enable the FTP Port. Now data is live on CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVH24101285296 | 2024-10-10 | 2024-10-11 | Cust 1512 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB connectivity needs to be done. | we took access to your system and completed CPCB configuration, now data is live on CPCB server. | we took access to your system and completed CPCB configuration, now data is live on CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH24103205297 | 2024-10-14 | 2024-10-22 | Cust 1477 | VASHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4 trans-A4TW033000(1) | commissioning | visited site and done commissioning and checked found ok. | visited site and done commissioning and checked found ok. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH24103315298 | 2024-10-15 | 2024-10-21 | Cust 1490 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Chlorine sensor | CHLORINE SENSOR(1) | sensorex sensor no response with change in sample cl2 | replaced sensorex sensor with new FMRCS-05AL 41 | replaced sensorex sensor with new FMRCS-05AL 41 | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH24103405299 | 2024-10-15 | 2024-10-16 | Cust 1337 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Local + MPCB data upload issue. | Data is live on Local + MPCB server. | Data is live on Local + MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH24104515300 | 2024-10-17 | 2024-10-21 | Cust 1498 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data uploading issue | As of now we are closing this complaint. We will reopen this once we get access to your system. | As of now we are closing this complaint. We will reopen this once we get access to your system. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVH24105225301 | 2024-10-18 | 2024-10-25 | Cust 1513 | VASHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | commissioning | transmitter checked in lab , calibration done with 4.0 and 7.0 buffer | transmitter checked in lab , calibration done with 4.0 and 7.0 buffer | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH24106315302 | 2024-10-21 | 2024-10-23 | Cust 1514 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC Visit for Multiparameter Analyser | Attended the AMC Visit, MOM Attached. | Attended the AMC Visit, MOM Attached. | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFVH24108655303 | 2024-10-25 | 2024-10-28 | Cust 1401 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data mismatch issue. | As discussed, we took access to the black box and made some changes from configuration end, now ETP Recycle Flow Meter data is matching with the actual readings and the same is verified by you also. | As discussed, we took access to the black box and made some changes from configuration end, now ETP Recycle Flow Meter data is matching with the actual readings and the same is verified by you also. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH24108735304 | 2024-10-25 | 2024-11-25 | Cust 1494 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | FM make pH card(1) | pH card is not detected | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210119803 LR Dt - 15-NOV-2024 | Replacement material has been booked through MARK EXPRESS PRIVATE LIMITED vide LR No - 210119803 LR Dt - 15-NOV-2024 | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFVH24109125305 | 2024-10-27 | 2024-11-02 | Cust 1515 | VASHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Transmitter Aquacon | Aquacon 2000(AQCN 2000 p) | error in readings | issue with current output , checked did not find issue , some misunderstanding regarding display readings and current output | issue with current output , checked did not find issue , some misunderstanding regarding display readings and current output | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH24113305306 | 2024-11-05 | 2024-11-06 | Cust 1337 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc call | done amc call and found ok. | done amc call and found ok. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVH24113315307 | 2024-11-05 | 2024-11-06 | Cust 1263 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc | visited site and done amc and checked found ok | done amc call and checked found working normal. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVH24113335308 | 2024-11-05 | 2024-11-06 | Cust 1506 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc call | done amc call and checked found working normal. | done amc call and checked found working normal. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVH24115525309 | 2024-11-11 | 2024-12-26 | Cust 1401 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data upload issue. | We took access of the black box and made some changes from configuration end, now data is live on CPCB + MPCB server. | We took access of the black box and made some changes from configuration end, now data is live on CPCB + MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH24117115310 | 2024-11-13 | 2024-11-15 | Cust 617 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data breaking issue on MPCB server. | We took access to the system and now data is live on LOCAL+CPCB+MPCB Server and continuously on all servers. | We took access to the system and now data is live on LOCAL+CPCB+MPCB Server and continuously on all servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVH24118285311 | 2024-11-15 | 2024-11-18 | Cust 1479 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue. | As of now we are cloning this complaint and will reopen it once the hardware side issue is resolved. | As of now we are cloning this complaint and will reopen it once the hardware side issue is resolved. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH24119835312 | 2024-11-19 | 2024-11-21 | Cust 1479 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB + MPCB data upload issue., | Now the hardware side issue is resolved, and data is live on CPCB + MPCB server. | Now the hardware side issue is resolved, and data is live on CPCB + MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH24121455313 | 2024-11-22 | 2024-11-25 | Cust 1516 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | RTDAS version to update | We are pleased to inform you that EnviroConnectDAS has been successfully updated to version 4.6. As discussed during our call, you will provide the hardware connection tomorrow. | We are pleased to inform you that EnviroConnectDAS has been successfully updated to version 4.6. As discussed during our call, you will provide the hardware connection tomorrow. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVH24121685314 | 2024-11-22 | 2024-12-24 | Cust 1493 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | motherboard and sample pump motor not working | Visit Done. Faulty components replaced. Unit is functioning Properly now | Visit Done. Faulty components replaced. Unit is functioning Properly now | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFVH24122885315 | 2024-11-25 | 2024-11-26 | Cust 1517 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | commissioning | visited site and done commissioning of all ph and conductivity transmitter and sensors. only one ph sensor was not responding and so suggested to send to chakan factory for further checking. | visited site and done commissioning of all ph and conductivity transmitter and sensors. only one ph sensor was not responding and so suggested to send to chakan factory for further checking. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH24124395316 | 2024-11-27 | 2024-12-01 | Cust 927 | VASHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | installing spares and make functional | analyser was made functional . Checked with stds of 1000ppm and 2000ppm. | analyser was made functional . Checked with stds of 1000ppm and 2000ppm. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH24125265317 | 2024-11-28 | 2024-12-01 | Cust 1490 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Chlorine sensor | CHLORINE SENSOR(1) | sensorex sensor not responding | checked transmitter , found working ok. | checked transmitter , found working ok. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH24128025318 | 2024-12-02 | 2024-12-06 | Cust 1481 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB MPCB data break issue. | now data is continuously live on CPCB/MPCB server and data availability is above 95%. | now data is continuously live on CPCB/MPCB server and data availability is above 95%. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH24129135319 | 2024-12-04 | 2024-12-09 | Cust 1480 | VASHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MIL card of Multiparameter(1) | install new MIL card | installation of new MIL Card done and checked found working normal. | installation of new MIL Card done and checked found working normal. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH24129145320 | 2024-12-04 | 2024-12-09 | Cust 1478 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc | visited site and done amc call and checked found working normal. | visited site and done amc call and checked found working normal. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVH24129905321 | 2024-12-05 | 2024-12-09 | Cust 1490 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc | done amc call and checked found working normal. | done amc call and checked found working normal. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVH24130005322 | 2024-12-05 | 2024-12-08 | Cust 499 | VASHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | commissioning | done commissioning of DO transmitter and found ok and for ph transmitter found electronics problem and suggested to send to Pune for repairs. | done commissioning of DO transmitter and found ok and for ph transmitter found electronics problem and suggested to send to Pune for repairs. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH24131935323 | 2024-12-09 | 2024-12-12 | Cust 608 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | software reinstallation needs to be done. | As of now we are closing this complaint and will open it once the Firewall/Antivirus is removed from your PC. | As of now we are closing this complaint and will open it once the Firewall/Antivirus is removed from your PC. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH24133805324 | 2024-12-13 | 2024-12-13 | Cust 1337 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB & Local host parameters data upload issue. | We had taken access today and we have resolved the issue of local server access. It was found that the Apache Tomcat was not running in mode hence the software not working.Now data is live in the local + MPCB server | Now data is live in the local + MPCB server | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH24135765325 | 2024-12-18 | 2024-12-24 | Cust 1517 | VASHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | commissioning | checked all instruments . | commissioning done | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH24136105326 | 2024-12-18 | 2024-12-22 | Cust 610 | VASHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Calibration | Visited site and done calibration of 6 no.s ph transmitter and check found working normal. | Done calibration work and found working normal. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFVH24138025327 | 2024-12-22 | 2024-12-27 | Cust 1503 | VASHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4Trans(1) | Commissioning of aqua4trans (ph , conductivity) and chemitech nitrate analyser | commissioning done and checked found ok. | commissioning done and checked found ok. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH24138925328 | 2024-12-24 | 2024-12-24 | Cust 1263 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Correct pH data not updating on CPCB and MPCB Server. | We took access to the system and completed the installation and configuration of software now data is live. | We took access to the system and completed the installation and configuration of software now data is live. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVH24139595329 | 2024-12-25 | 2024-12-31 | Cust 1494 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | display blank | found issue with 12V power supply . Will check if meter is under warranty so will share offer . | found issue with 12V power supply . Will check if meter is under warranty so will share offer . | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH24140405330 | 2024-12-27 | 2025-02-19 | Cust 1512 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | MPCB connectivity needs to be done. | We have completed MPCB connectivity for this site, now data is coming till the MPCB server. | We have completed MPCB connectivity for this site, now data is coming till the MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH24140925331 | 2024-12-28 | 2024-12-28 | Cust 1337 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Localhost not accessible. | We took access to the system and found services were stopped we have started and now localhost is accessible and data is live on server. | We took access to the system and found services were stopped we have started and now localhost is accessible and data is live on server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVH24141065332 | 2024-12-28 | 2024-12-30 | Cust 617 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on MPCB server | We have accessed the system, and the data for the Coke Oven Plant 2_BOD Plant_outlet recycling is now live on the MPCB server. | We have accessed the system, and the data for the Coke Oven Plant 2_BOD Plant_outlet recycling is now live on the MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVH24141905333 | 2024-12-29 | 2024-12-31 | Cust 1263 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | routine amc | Commissioning done and checked found ok | Commissioning done and checked found ok | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVH24145715334 | 2025-01-02 | 2025-01-06 | Cust 610 | VASHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | ORP transmitter(1) | Calibration work | visited site and done calibration work and found ok. | visited site and done calibration work and found ok. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFVH24150925335 | 2025-01-15 | 2025-01-21 | Cust 1480 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc | amc work done | amc work done | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVH24151375336 | 2025-01-15 | 2025-01-20 | Cust 1493 | VASHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | commissioning | visited site and done commissioning and checked found working normal. | visited site and done commissioning and checked found working normal. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH24151385337 | 2025-01-15 | 2025-01-20 | Cust 1481 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc call | done amc and checked found working normal. | done amc and checked found working normal. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVH24153415338 | 2025-01-20 | 2025-01-26 | Cust 1518 | VASHI | Steam System Division | CLOSED | QUALITY PROBLEM | TRAP - MONITORING SYSTEM | FMTMS(FMTMS) | DIP Switch Setting - As a standard, DIP Switch setting for S2 should be ON ON ON OFF. But onsite it was seen that the setting was ON OFF ON OFF. Proper setting needs to be done in factory and then dispatched to site. | Visit made. Mail attached. | Visit made. Mail attached. | YES | Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFVH24154435339 | 2025-01-22 | 2025-01-24 | Cust 1337 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Our Local host & MPCB server is not working | As of now we are closing this complaint and will reopen it again once we get remote access of the system from coutomer end. | As of now we are closing this complaint and will reopen it again once we get remote access of the system from coutomer end. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH24154515340 | 2025-01-22 | 2025-01-26 | Cust 1477 | VASHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Turbidity Meter | Optex turbidity sensor(1) | commissioning | Visited site and done commissioning and checked found working normal. | Visited site and done commissioning and checked found working normal. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH24154575341 | 2025-01-22 | 2025-01-29 | Cust 1493 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Aqua4Trans | Communication card of Aqua4Trans(1) | showing zero readings on modscan. | RF for aqua4trans has been intitated and rf number is 138255425, when material receives will coordinate with customer and install the same. | RF for aqua4trans has been intitated and rf number is 138255425, when material receives will coordinate with customer and install the same. | EXTENDED | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH24154665342 | 2025-01-22 | 2025-01-26 | Cust 1337 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc work | Visited site and done amc and check found working normal. | Visited site and done amc and check found working normal. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVH24154945343 | 2025-01-23 | 2025-01-28 | Cust 1519 | VASHI | Steam System Division | CLOSED | MANUFACTURING ISSUES | PRV - MEASURETROL | FM MEASURETROL(SIZE 050NB) | During commissioning on 22-Jan, it was found that bluetooth PCB LED is not glowing. Device not detected in mobile app | Reference RF OA Number 5525109022 has been booked for M/s DEEPAK NOVOCHEM SPECIALITIES LTD. Item: SPARE-050MTROL-Q1EFS41MI = Qty01No | Reference RF OA Number 5525109022 has been booked for M/s DEEPAK NOVOCHEM SPECIALITIES LTD. Item: SPARE-050MTROL-Q1EFS41MI = Qty01No | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFVH24157225344 | 2025-01-27 | 2025-01-31 | Cust 1514 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC visit for multiparameter analyser | Visit made. MOM attached. | Visit made. MOM attached. | NO | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFVH24157525345 | 2025-01-27 | 2025-02-03 | Cust 1494 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | error in readings | cleaned the flow cell , tubings , recalibrated | cleaned the flow cell , tubings , recalibrated | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH24158105346 | 2025-01-28 | 2025-02-03 | Cust 1263 | VASHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | Flow cell for Multiparameter(AWA-FLOWCELL) | calibration work | calibration done | calibration done | NO | Paid | Core | N |  | Analytical - CHAKAN
AFVH24158985347 | 2025-01-29 | 2025-01-31 | Cust 617 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on MPCB server. | Now the data is live on the MPCB server. | Now the data is live on the MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVH24163665348 | 2025-02-05 | 2025-02-10 | Cust 1490 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc | amc done | amc work done | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVH24164445349 | 2025-02-07 | 2025-02-10 | Cust 1478 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc | amc done | amc work done | NO | Paid | Core | N |  | Analytical - CHAKAN
AFVH24164895350 | 2025-02-08 | 2025-02-14 | Cust 610 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | SS flow through chamber need to be change with PP chamber for all conductivity sensors. 4no.s total | visited site and changed the FTC and checked found ok. | visited site and changed the FTC and checked found ok. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH24164905351 | 2025-02-08 | 2025-02-14 | Cust 1493 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | SULPHIDE ANALYSER | SULPHIDE 8810 Single Channel ANALYSER(8810 Sulphide) | light too low error. awa sulphide analyser | visited site and found multiple problems in analyser, so suggested to send to pune for repairs. | visited site and found multiple problems in analyser, so suggested to send to pune for repairs. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH24166465352 | 2025-02-12 | 2025-02-14 | Cust 1337 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Local + MPCB data upload issue. | We took access to your system and now the local server is accessible, and data is live on the Local server. | We took access to your system and now the local server is accessible, and data is live on the Local server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH24166585353 | 2025-02-12 | 2025-02-14 | Cust 617 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on MPCB server. | We took access to your system and now data is live on MPCB server. | We took access to your system and now data is live on MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVH24168145354 | 2025-02-16 | 2025-02-21 | Cust 564 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Smartpro-con) | readings problem | visited site and done calibration of ph transmitters and for conductivity transmitter found problem and so suggested to send to chakan factory for repairs. | visited site and done calibration of ph transmitters and for conductivity transmitter found problem and so suggested to send to chakan factory for repairs. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH24170625355 | 2025-02-21 | 2025-02-24 | Cust 1496 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | PM data not updating on CPCB server. | Commercially not cleared. | Commercially not cleared. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFVH24170695356 | 2025-02-21 | 2025-02-24 | Cust 617 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Localhost is not accessible. | We have updated the license file after which data is live on LOCAL+CPCB+MPCB server. | We have updated the license file after which data is live on LOCAL+CPCB+MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVH24171845357 | 2025-02-24 | 2025-03-02 | Cust 1520 | VASHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Electrode | FM1106(1) | not responding to pH buffers | electrode not responding to pH buffer . requested customer to send it to pune | electrode not responding to pH buffer . requested customer to send it to pune | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH24172695358 | 2025-02-26 | 2025-02-28 | Cust 1481 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue. | As of now we are closing this complaint and will reopen it once we get system access. | As of now we are closing this complaint and will reopen it once we get system access. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH24174655359 | 2025-02-28 | 2025-03-04 | Cust 608 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc call | visited site and done amc call and checked found ok. | visited site and done amc call and checked found ok. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVH24177585360 | 2025-03-05 | 2025-03-06 | Cust 1337 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on LOCAL+MPCB server. | We took access to your system and found services were stopped and restarting the services data is live on CPCB server. | We took access to your system and found services were stopped and restarting the services data is live on CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVH24177755361 | 2025-03-05 | 2025-03-07 | Cust 1401 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Camera configuration needs to be done. | we took access to your system and updated the configuration of the camera, now the camera is streaming live on CPCB server. | we took access to your system and updated the configuration of the camera, now the camera is streaming live on CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH24178265362 | 2025-03-06 | 2025-03-07 | Cust 1479 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data upload issue. | we took access to your system and now the issue has been resolved, and data is live on the MPCB server. | we took access to your system and now the issue has been resolved, and data is live on the MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH24178905363 | 2025-03-07 | 2025-03-12 | Cust 1492 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | Flow cell for Multiparameter(AWA-FLOWCELL) | amc | done amc work and checked found ok. | done amc work and checked found ok. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVH24180885364 | 2025-03-11 | 2025-03-14 | Cust 1512 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | CPCB data mismatch issue. | we took access of the black box and made some changes from configuration end, now Flow data is matching with the actual readings | we took access of the black box and made some changes from configuration end, now Flow data is matching with the actual readings | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH24181135365 | 2025-03-12 | 2025-03-16 | Cust 1337 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc work | visited site and done amc work and checked found working normal. | visited site and done amc work and checked found working normal. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVH24181155366 | 2025-03-12 | 2025-03-16 | Cust 1510 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc work | visited site and done amc work and checked found working normal. | visited site and done amc work and checked found working normal. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVH24181165367 | 2025-03-12 | 2025-03-16 | Cust 1263 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc work | visited site and done amc work and checked found working normal. | visited site and done amc work and checked found working normal. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVH24182655368 | 2025-03-17 | 2025-03-20 | Cust 1481 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data break issue. | login Issue resolved. | login Issue resolved. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH24183325369 | 2025-03-18 | 2025-03-21 | Cust 1263 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB server is not accessible. | We have share the link with the customer and MPCB is accessible. | We have share the link with the customer and MPCB is accessible. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVH24183465370 | 2025-03-18 | 2025-03-23 | Cust 1483 | VASHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | commissioning | visited site and done commissioning and checked found ok. | visited site and done commissioning and checked found ok. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH24184365371 | 2025-03-20 | 2025-03-21 | Cust 1337 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on LOCAL+MPCB server. | As of now we are closing this complaint and we will re-open once the system is accessible. | As of now we are closing this complaint and we will re-open once the system is accessible. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVH24185585372 | 2025-03-24 | 2025-03-27 | Cust 1337 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on LOCAL+MPCB server. | As of now we are closing tis complaint we will reopen once the PC name is changed from customer end | As of now we are closing tis complaint we will reopen once the PC name is changed from customer end | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVH24186535373 | 2025-03-25 | 2025-03-28 | Cust 1481 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data upload issue. | we took access to your system and updated the EnviroconnectDAS and Patches till Patch55. data has been continuously uploading to the MPCB server. | we took access to your system and updated the EnviroconnectDAS and Patches till Patch55. data has been continuously uploading to the MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH24187465374 | 2025-03-27 | 2025-05-09 | Cust 1521 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | pH Transmitter Aquamon | Ph Transmitter Aquamon 2000(PH2000) | Issue is with transmitter , it is continuously booting. Following is the Sr and details. Sr No-C4EACB230486 Range- 0 to 10. Vikas Jadhav Director +918369709799 Sprintek Technologies Pvt Ltd. W-14, MIDC Phase 2, Dombivli (E), Thane-421204 | spare transmitter is arranged on RF | spare transmitter is arranged on RF | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFVH24188245375 | 2025-03-28 | 2025-03-29 | Cust 1337 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Configuration and installation of software. | We took access to the system and completed the installation and configuration of software. Now data is live on LOCAL+MPCB server. | We took access to the system and completed the installation and configuration of software. Now data is live on LOCAL+MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVH24190485376 | 2025-03-31 | 2025-03-31 | Cust 1514 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC Visit for Multiparameter | Visit made. MOM attached. | Visit made. MOM attached. | NO | Non Warranty (Free) | Process | N |  | Analytical - CHAKAN
AFVH25000485377 | 2025-04-02 | 2025-04-04 | Cust 1479 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data mismatch issue | As of now we are closing this complaint. We will reopen this once the issue is resolved from the hardware side. We have checked the data on modscan and observed that the same readings of flow are reflected to the CPCB server. | As of now we are closing this complaint. We will reopen this once the issue is resolved from the hardware side. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVH25001875378 | 2025-04-06 | 2025-04-13 | Cust 1513 | VASHI | Analytical - CHAKAN | CLOSED | PRO-ACTIVE VISIT | AQUADO Analyzer | OXYSENS(OXYSENS) | installation of new sensor to aqua2trans filed trail unit. R&D work | new aqua2trans is installed , new smartpro 8968 is installed . RTU commissioning is pending | new aqua2trans is installed , new smartpro 8968 is installed . RTU commissioning is pending | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH25001985379 | 2025-04-07 | 2025-04-09 | Cust 1479 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data mismatch issue | Now the issue has been resolved and data is live on the CPCB server. | Now the issue has been resolved and data is live on the CPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVH25002125380 | 2025-04-07 | 2025-04-09 | Cust 617 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | localhost not accessible. | localhost is accessible now. Services were stopped due to which data was not live. | localhost is accessible now. Services were stopped due to which data was not live. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVH25002555381 | 2025-04-07 | 2025-04-13 | Cust 1263 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | Display with cpu board | change the display with cpu card and power supply card and checked found ok. | change the display with cpu card and power supply card and checked found ok. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH25003825382 | 2025-04-10 | 2025-04-14 | Cust 617 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Totalizer value is not correct on CPCB server. | We took access to the system and found totalizer values on MODSCAN are high. Due to which on CPCB values are getting high asked customer to check data from their end. | We took access to the system and found totalizer values on MODSCAN are high. Due to which on CPCB values are getting high asked customer to check data from their end. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVH25006105383 | 2025-04-16 | 2025-04-18 | Cust 1481 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data Availability issue on the MPCB server for the ETP station. | Here we have updated the Delayed data jar file and now the Data availability is 99 % on the MPCB server. | Here we have updated the Delayed data jar file and now the Data availability is 99 % on the MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH25006435384 | 2025-04-17 | 2025-04-23 | Cust 1522 | VASHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | commissioning | commissioning done | commissioning done | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH25006685385 | 2025-04-18 | 2025-04-24 | Cust 1481 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | AMC call | Vardhaman engineer Mr. suraj attend the site for amc and complete the same. | Vardhaman engineer Mr. suraj attend the site for amc and complete the same. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVH25007715386 | 2025-04-21 | 2025-04-24 | Cust 617 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on RTDMS server. | We took access to the system and now data is live on RTDMS server. | We took access to the system and now data is live on RTDMS server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVH25008685387 | 2025-04-23 | 2025-05-01 | Cust 1523 | VASHI | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Electrode | Ph Electrode(FM1104(4C)) | Replacement of PH sensor Contact Person: Mr. Ganesh Thite - 9892880434 / 8097308666 Address: SWAMINI ENGINEERING FLAT NO.102, BUILDING NO. 102 SAMRUDDHI EVERGREEN, PODDAR BADLAPUR (E) MUMBAI 421503 Maharashtra India | new electrode is issued on goodwill | new electrode is issued on goodwill | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFVH25009375388 | 2025-04-24 | 2025-04-28 | Cust 1520 | VASHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | install new aqua2trans pH transmitter | installed new aqua2trans pH transmitter | installed new aqua2trans pH transmitter | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH25011425389 | 2025-04-28 | 2025-05-03 | Cust 1479 | VASHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc | done amc work and checked found working normal. | done amc work and checked found working normal. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVH25011585390 | 2025-04-29 | 2025-04-30 | Cust 617 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data not updating on CPCB+MPCB server. | We took access to the system and now data is updating on CPCB+MPCB server. | We took access to the system and now data is updating on CPCB+MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVH25015765391 | 2025-05-05 | 2025-05-06 | Cust 1479 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data upload issue. | we took access to your system and now the issue has been resolved and data is live on MPCB server for MPL. | we took access to your system and now the issue has been resolved and data is live on MPCB server for MPL. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH25015775392 | 2025-05-05 | 2025-05-08 | Cust 1498 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | local host is not opening and date and time are showing incorrect on CPCB server. | Now the correct date and time is updating to the CPCB server. Localhost is also accessible. | Now the correct date and time is updating to the CPCB server. Localhost is also accessible. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVH25016755393 | 2025-05-07 | 2025-05-12 | Cust 1490 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc | amc work done | amc work done | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVH25016785394 | 2025-05-07 | 2025-05-19 | Cust 617 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Totalizer value mismatch. | We took access to the system and updated the patch 64 and added calculations on local server after which correct Daily Totalizer data is updating on both LOCAL+MPCB server. | We took access to the system and updated the patch 64 and added calculations on local server after which correct Daily Totalizer data is updating on both LOCAL+MPCB server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVH25018185395 | 2025-05-10 | 2025-05-14 | Cust 1479 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc | amc done by suraj and working ok. | amc done by suraj and working ok. | NO | Periodic Maintenance (AMC) | Core | N |  | Analytical - CHAKAN
AFVH25018195396 | 2025-05-10 | 2025-05-14 | Cust 1479 | VASHI | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | amc | amc done by suraj and working ok. | amc done by suraj and working ok. | NO | Paid | Core | N |  | Analytical - CHAKAN
AFVH25018215397 | 2025-05-10 | 2025-05-14 | Cust 1492 | VASHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | CX2000 MPA commissioning | commissioning done and checked found working normal. | commissioning done and checked found working normal. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH25019935398 | 2025-05-14 | 2025-05-19 | Cust 653 | VASHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4 trans-A4TW033000(1) | commissioning | visited site and done commissioning and checked found ok. | visited site and done commissioning and checked found ok. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH25019945399 | 2025-05-14 | 2025-05-19 | Cust 534 | VASHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | commissioning | visited site and done commissioning and checked found ok. | done comissioning and checked found ok. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH25019945399 | 2025-05-14 | 2025-05-19 | Cust 534 | VASHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Aqua4Trans | Aqua4 trans-A4TW033000(1) | commissioning | done comissioning and checked found ok. | done comissioning and checked found ok. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH25023535400 | 2025-05-22 | 2025-05-27 | Cust 1520 | VASHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | commissioning | new transmitter is installed , checked calibration with pH stds of 4 and 7 . | new transmitter is installed , checked calibration with pH stds of 4 and 7 . | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH25023845401 | 2025-05-23 | 2025-05-26 | Cust 1480 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB + MPCB data upload issue. | we took access to the system and observed that the Communication for ETP was changed (From COM4 to COM5), so we have made changes in our server for the same now data is live on CPCB + MPCB server. | we took access to the system and observed that the Communication for ETP was changed (From COM4 to COM5), so we have made changes in our server for the same now data is live on CPCB + MPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH25024925402 | 2025-05-26 | 2025-05-29 | Cust 1512 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | MPCB data upload issue. | As of now we are closing this complaint and will reopen it once we get system access. | As of now we are closing this complaint and will reopen it once we get system access. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH25028165403 | 2025-05-30 | 2025-06-03 | Cust 1493 | VASHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | SULPHIDE ANALYSER | SULPHIDE 8810 Single Channel ANALYSER(8810 Sulphide) | recommissioning after repairs. awa SULPHIDE ANALYSER | recommissioning done of sulphide analyser and working normal. | recommissioning done of sulphide analyser and working normal. | EXTENDED | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH25030125404 | 2025-06-02 | 2025-06-05 | Cust 1512 | VASHI | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | FM BLACK BOX | FM BLACK BOX(FM BLACK BOX) | MPCB data mismatch issue. | we took access of the black box and made some changes in FMBB, now data is live on MPCB sever and matching with the actual readings and same is verified by customer also. | we took access of the black box and made some changes in FMBB, now data is live on MPCB sever and matching with the actual readings and same is verified by customer also. | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFVH25030215405 | 2025-06-02 | 2025-06-08 | Cust 1490 | VASHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(Ph-Trans-Smartpro) | actual transmitter is aqua2trans trail unit | installation of aqua2trans , FM1104 dip pH assembly and RTU completed . RTU programing done . | installation of aqua2trans , FM1104 dip pH assembly and RTU completed . RTU programing done . | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH25032295406 | 2025-06-07 | 2025-06-12 | Cust 1479 | VASHI | Analytical - CHAKAN | CLOSED | COMMISSIONING | Digital Product | EVERSENSE PC() | energy meter data on modscan checking and connectivity. | suraj has visited site and checked readings on modscan found ok for unit 2 and unit 10. | suraj has visited site and checked readings of energy meters on modscan found ok for unit 2 and unit 10. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH25032495407 | 2025-06-09 | 2025-06-16 | Cust 1513 | VASHI | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | AQUADO Analyzer | OXYFERM(OXYFERM) | calibration of aqua2trans trial unit | attended aqua2trans do transmitter with Sameer from r&d | attended aqua2trans do transmitter with Sameer from r&d | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVH25035605408 | 2025-06-16 |  | Cust 1480 | VASHI | Analytical - CHAKAN | ASSIGNED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | MPCB data upload issue. |  |  | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVH25035615409 | 2025-06-16 |  | Cust 1481 | VASHI | Analytical - CHAKAN | ASSIGNED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB + MPCB data upload issue. |  |  | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVP23033615410 | 2023-06-01 | 2023-06-01 | Cust 1524 | VIZAG | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Smartpro-con) | The conductivity transmitter is showing incorrect readings, needs support | Telephonic support has been provided and the issue is addressed. | Telephonic support has been provided and the issue is addressed. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP23036275411 | 2023-06-07 | 2023-06-09 | Cust 1525 | VIZAG | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB and APPCB data uploaded issue. | The issue has been resolved, and we have observed that in your system some other application was using the 8080 port that's why our application was not able to access that port and now we made changes in our application & set it to 8081 port so now the data is live on the server. | Issue has been resolved and data is live on the server hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVP23045315412 | 2023-06-24 | 2023-06-27 | Cust 1525 | VIZAG | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | APPCB data upload issue. | increased the java memory so now data is live on the APPCB server. | increased the java memory so now data is live on the APPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVP23050505413 | 2023-07-03 | 2023-07-04 | Cust 1526 | VIZAG | Codel Products - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB and APPCB data uploaded issue. | Here the issue is from the APPCB server end and now the issue resolved by the APPCB team and the data is live for CPCB the AIPL team resolved the issue and the data is live. | The issue has been resolved and data is live on both servers hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVP23055895414 | 2023-07-13 | 2023-07-14 | Cust 1527 | VIZAG | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | APPCB data upload issue. | As of now we are closing this complaint, we will reopen it after getting system access from customers side | As of now we are closing this complaint, we will reopen it after getting system access from customers side | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVP23060865415 | 2023-07-24 | 2023-07-25 | Cust 1527 | VIZAG | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB + APPCB data upload issue. | Data is live on all the server, hence we are closing this complaint. | Data is live on all the server, hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVP23062715416 | 2023-07-27 | 2023-07-31 | Cust 1525 | VIZAG | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | APPCB data upload issue. | there is issue from customers network end (APPCB server url and FTP port block in network)after allowing url and FTP port data is now live on APPCB server. | there is issue from customers network end (APPCB server url and FTP port block in network)after allowing url and FTP port data is now live on APPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVP23070085417 | 2023-08-07 | 2023-08-13 | Cust 1528 | VIZAG | Analytical - CHAKAN | CLOSED | COMMISSIONING | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | The oil in water Analyzers are to be set right and calibrated. | The site is concluded and attached the Minutes copy. | The site is concluded and attached the Minutes copy. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP23070275418 | 2023-08-08 | 2023-08-11 | Cust 1525 | VIZAG | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | APPCB data upload issue. | As of now we are closing this complaint, we will reopen this after the server url and FTP port is allowed from your end. | As of now we are closing this complaint, we will reopen this after the server url and FTP port is allowed from your end. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVP23078135419 | 2023-08-24 | 2023-08-31 | Cust 1529 | VIZAG | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Electrode Cable(Ph-Cab) | The new pH transmitter is to be calibrated at site. | The pH transmitter is calibrated at site. | The pH transmitter is calibrated at site. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP23078245420 | 2023-08-24 | 2023-08-25 | Cust 1527 | VIZAG | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | APPCB data upload issue. | As discussed, please share the access to your system once the hardware issue is resolved from your end. | As discussed, please share the access to your system once the hardware issue is resolved from your end. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVP23078735421 | 2023-08-25 | 2023-08-28 | Cust 1525 | VIZAG | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | APPCB data upload issue. | As of now we are closing this complaint and we will reopen it after getting a system access from customers end. | As of now we are closing this complaint and we will reopen it after getting a system access from customers end. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVP23080785422 | 2023-08-28 | 2023-09-02 | Cust 1493 | VIZAG | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Tranmitter Smartpro | Conductivity Electrode Cable(con-cab) | The Conductivity values are high needs checking & calibration at site | The Analyzers are calibrated and taken into line. Minutes copy is attached. | The Analyzers are calibrated and taken into line. Minutes copy is attached. | NO | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP23081085423 | 2023-08-29 | 2023-09-01 | Cust 1527 | VIZAG | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB + APPCB data upload issue. | As of now we are closing this complaint, we will reopen this after getting system access from customers end. | As of now we are closing this complaint, we will reopen this after getting system access from customers end. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVP23081685424 | 2023-08-30 | 2023-09-01 | Cust 1525 | VIZAG | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | APPCB data upload issue. | In customers network APPCB server url block, we convey same to the customer as of now we are closing this complaint. | In customers network APPCB server url block, we convey same to the customer as of now we are closing this complaint. | YES | No Trip | Core | Y | Product / equipment failure impacting safety of equipment, process or people. | Analytical - CHAKAN
AFVP23084225425 | 2023-09-04 | 2023-09-10 | Cust 1529 | VIZAG | Analytical - CHAKAN | CLOSED | COMMISSIONING | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | Recommissioning of the oil in water Analyzer with replacement of filter | The optical filter is replaced and set right the Analyzer. | The optical filter is replaced and set right the Analyzer. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP23085415426 | 2023-09-07 | 2023-09-08 | Cust 1527 | VIZAG | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | CPCB data upload issue. | We took the access to the system and did some changes from configuration end and now CPCB data is matching to the actual reading. | We took the access to the system and did some changes from configuration end and now CPCB data is matching to the actual reading. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVP23089475427 | 2023-09-16 | 2023-09-23 | Cust 1529 | VIZAG | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | The Oil in water Analyser is not working needs to be checked. | The site is attended & concluded. Minutes copy is attached. | The site is attended & concluded. Minutes copy is attached. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP23089485428 | 2023-09-16 | 2023-09-23 | Cust 604 | VIZAG | Analytical - CHAKAN | CLOSED | COMMISSIONING | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | Replacement of display and set right the Analyzer | The display is replaced by Amol sathe and the Analyzer is working ok. | The display is replaced by Amol sathe and the Analyzer is working ok. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP23099655429 | 2023-10-04 | 2023-10-09 | Cust 972 | VIZAG | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | for ph transmitter virification with simulator | calibration carried out and informed to customer on the results | calibration carried out and informed to customer on the results | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFVP23105845430 | 2023-10-17 | 2023-10-17 | Cust 1529 | VIZAG | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | Ph Tranmitter Smartpro | Display of Smartpro(1) | ph value showing wrong | calibration done | analyzer detector pcb replaced | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP23105845430 | 2023-10-17 | 2023-10-17 | Cust 1529 | VIZAG | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | OIL ANALYSER | PM Board(PM-BD) | photo -x- detector board replacement | Detector pcb replaced & checked done | analyzer detector pcb replaced | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP23105855431 | 2023-10-17 | 2023-10-20 | Cust 1525 | VIZAG | Analytical - CHAKAN | CLOSED | COMMISSIONING | Data logger | DLP 200(1) | New VOC Analyzer hookup to PC | The VOC Analyzer data is hooked up to PC.Minutes copy is attached. | The VOC Analyzer data is hooked up to PC. | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP23106515432 | 2023-10-18 | 2023-10-23 | Cust 1000 | VIZAG | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | MULTIPARAMETER ANALYSER(Multipar- 01) | for troubleshooting | analyzer physically verified and found abnormal and poweredup with temperory power supply but it is not poweredup and informed customer to send for repair | analyzer physically verified and found abnormal and poweredup with temperory power supply but it is not poweredup and informed customer to send for repair | NO | No Trip | Process | N |  | Analytical - CHAKAN
AFVP23106975433 | 2023-10-19 | 2023-10-20 | Cust 1525 | VIZAG | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Additon of new device on CPCB Server. | We successfully add a VOC meter in local configuration and now data is live on local server. | We successfully add a VOC meter in local configuration and now data is live on local server. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVP23115405434 | 2023-11-03 | 2023-11-03 | Cust 1525 | VIZAG | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | APPCB data upload issue. | Now data is live on APPCB server, hence we are closing this complaint. | Now data is live on APPCB server, hence we are closing this complaint. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVP23117975435 | 2023-11-08 | 2023-11-10 | Cust 1524 | VIZAG | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | Chlorine Analyzer-MAC 100 | MAC 100(1) | Chlorine analyzer checking | calibration done | value mis match | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP23120455436 | 2023-11-16 | 2023-11-18 | Cust 1525 | VIZAG | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | APPCB data upload issue. | now data is live on APPCB server. | now data is live on APPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVP23125825437 | 2023-11-27 | 2024-01-18 | Cust 1530 | VIZAG | Steam System Division | CLOSED | QUALITY PROBLEM | PRS | PRS - SIZE 80X125 - WITH ARCA-POSITIONER(PRS80X125ARCA-POSITIONER) | measure trol control valve inlet PT connector damaged | Material arranged from R&D | Material arranged from R&amp;amp;D | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFVP23127435438 | 2023-11-29 | 2023-11-30 | Cust 1529 | VIZAG | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Electrode Cable(Ph-Cab) | Flow through chamber replaced | Flow through camber replaced | pH calibration | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP23127445439 | 2023-11-29 | 2023-12-01 | Cust 1531 | VIZAG | Analytical - CHAKAN | CLOSED | AMC | MULTIPARAMETER ANALYSER | LCD of Multiparameter Display(1) | calibration | calibration | AMC | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP23131535440 | 2023-12-05 | 2023-12-07 | Cust 1525 | VIZAG | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | APPCB data upload issue. | After restart Enviroconnect DAS services, now data is live on APPCB server. | After restart Enviroconnect DAS services, now data is live on APPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVP23132545441 | 2023-12-07 | 2023-12-14 | Cust 1532 | VIZAG | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | VISI pro DO(VISI pro DO) | for do commissioning | aqua do system commissioning complered | aqua do system commissioning complered | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFVP23139525442 | 2023-12-20 | 2023-12-22 | Cust 1529 | VIZAG | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(Ph-Trans-Smartpro) | training | Training on Analyzers is provided to the HPCL representatives at site | Training of Analyzers is provided to the HPCL representatives at site | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP23141295443 | 2023-12-25 | 2024-01-31 | Cust 1530 | VIZAG | Steam System Division | CLOSED | PRODUCT FAILURE | PRS | PRS - SIZE 100X125 - WITH ARCA-POSITIONER(PRS100X125ARCA-POSITIONER) | new measuretrol prs station not working | The electronics is changed | The electronics is changed | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFVP23144605444 | 2023-12-29 | 2024-01-01 | Cust 1533 | VIZAG | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Electrode | FM1106(1) | electrode not responding | pH electrode not responding | pH electrode priblem | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP23145615445 | 2023-12-31 | 2024-01-18 | Cust 1530 | VIZAG | Steam System Division | CLOSED | QUALITY PROBLEM | PRS | PRS - SIZE 100X150 - WITH ARCA-POSITIONER(PRS100X150ARCA-POSITIONER) | measuretrol prv not working as per the setpoints | prv inspected and found no 24v dc output to the pressure transmitters from the card itself | prv inspected and found no 24v dc output to the pressure transmitters from the card itself | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFVP23148225446 | 2024-01-04 | 2024-01-18 | Cust 1530 | VIZAG | Steam System Division | CLOSED | PRODUCT FAILURE | PRS | PRS - SIZE 100X150 - WITH ARCA-POSITIONER(PRS100X150ARCA-POSITIONER) | measuretrol internal cards replacement` | visited site and replaced the cards with new and all inter connecting cable also replaced with new and taken valve into line and kept under observation | visited site and replaced the cards with new and all inter connecting cable also replaced with new and taken valve into line and kept under observation | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFVP23152555447 | 2024-01-11 | 2024-01-13 | Cust 1534 | VIZAG | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | The commissioning assistance is required for the supplied Analyzers. | Customer not responded & gate pass not granted | Calibration | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP23152555447 | 2024-01-11 | 2024-01-13 | Cust 1534 | VIZAG | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | The commissioning assistance is required for the supplied Analyzers. | Customer not responded to calls & no gate pass | Calibration | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP23153955448 | 2024-01-15 | 2024-01-15 | Cust 1530 | VIZAG | Steam System Division | CLOSED | QUALITY PROBLEM | PRS | PRS - SIZE 80X125 - WITH ARCA-POSITIONER(PRS80X125ARCA-POSITIONER) | MEASURETROL FMPRS - Cable lengths to be re-defined as the current cable length is not sufficient. JB not supplied | Mr Ramesh Tarra arrange for a JB locally and address the issue | Mr Ramesh Tarra arrange for a JB locally and address the issue | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFVP23153995449 | 2024-01-15 | 2024-01-15 | Cust 1530 | VIZAG | Steam System Division | CLOSED | QUALITY PROBLEM | PRS | PRS - SIZE 80X125 - WITH ARCA-POSITIONER(PRS80X125ARCA-POSITIONER) | MEASURETROL FMPRS - 3) Quality of the connectors is not reliable. The replacement was arranged from R&D. 4) Connector should be screwed-on with gland instead of crimp connectors (reference of Codel / Krohne given). | Spares sent from R&D Stock and same replaced by Ramesh Tarra at site. Inputs taken by QC for RCA and CAPA | Spares sent from R&amp;amp;D Stock and same replaced by Ramesh Tarra at site. Inputs taken by QC for RCA and CAPA | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFVP23154025450 | 2024-01-15 | 2024-01-15 | Cust 1530 | VIZAG | Steam System Division | CLOSED | QUALITY PROBLEM | PRS | PRS - SIZE 80X125 - WITH ARCA-POSITIONER(PRS80X125ARCA-POSITIONER) | MEASURETROL FMPRS - The way the cabling is laid inside the actuator could be made better. | The site inputs are taken up with MFG and QC | The site inputs are taken up with MFG and QC | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFVP23154045451 | 2024-01-15 | 2024-01-18 | Cust 1530 | VIZAG | Steam System Division | CLOSED | QUALITY PROBLEM | PRS | PRS - SIZE 80X125 - WITH ARCA-POSITIONER(PRS80X125ARCA-POSITIONER) | MEASURETROL FMPRS - The range indication on the app always showing limited range even if we stand next to the valve | The input has been taken up by Reliability, MFG and QC Teams and will revert with RCA and CAPA. Site is commissioned | The input has been taken up by Reliability, MFG and QC Teams and will revert with RCA and CAPA. Site is commissioned | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFVP23154065452 | 2024-01-15 | 2024-01-15 | Cust 1530 | VIZAG | Steam System Division | CLOSED | QUALITY PROBLEM | PRS | PRS - SIZE 80X125 - WITH ARCA-POSITIONER(PRS80X125ARCA-POSITIONER) | MEASURETROL FMPRS - 25th December 2023 - Found that there is no output getting generating from the card itself for the both P1 and P2. Jay from R&D visited the site with new Card and replaced the same on 03rd January 2024 | MEASURETROL FMPRS - 25th December 2023 - Found that there is no output getting generating from the card itself for the both P1 and P2. Jay from R&D visited the site with new Card and replaced the same on 03rd January 2024 | MEASURETROL FMPRS - Found that there is no output getting generating from the card itself for the both P1 and P2. Jay from RND visited the site with new Card and replaced the same on 03rd January 2024 | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFVP23154085453 | 2024-01-15 | 2024-01-15 | Cust 1530 | VIZAG | Steam System Division | CLOSED | QUALITY PROBLEM | PRS | PRS - SIZE 80X125 - WITH ARCA-POSITIONER(PRS80X125ARCA-POSITIONER) | MEASURETROL FMPRS - Cable Quality and Reliable Conduits for cable and its routing. | The inputs are given to Reliability, MFG and QC | The inputs are given to Reliability, MFG and QC | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFVP23157655454 | 2024-01-22 | 2024-01-27 | Cust 1534 | VIZAG | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Aqua4Trans | Aqua4 trans-A4TW033000(1) | Calibration | Calibration done | Calibration done | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP23157655454 | 2024-01-22 | 2024-01-27 | Cust 1534 | VIZAG | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Aqua4Trans | Aqua4 trans-A4TW033000(1) | Calibration | Calibration done | Calibration done | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP23157655454 | 2024-01-22 | 2024-01-27 | Cust 1534 | VIZAG | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Aqua4Trans | Aqua4 trans-A4TW033000(1) | Calibration | Calibration done | Calibration done | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP23157655454 | 2024-01-22 | 2024-01-27 | Cust 1534 | VIZAG | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Aqua4Trans | Aqua4 trans-A4TW033000(1) | Calibration | Calibration done | Calibration done | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP23157655454 | 2024-01-22 | 2024-01-27 | Cust 1534 | VIZAG | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Aqua4Trans | Aqua4 trans-A4TW033000(1) | Calibration | Calibration done | Calibration done | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP23157655454 | 2024-01-22 | 2024-01-27 | Cust 1534 | VIZAG | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Aqua4Trans | Aqua4 trans-A4TW033000(1) | Calibration | Calibration done | Calibration done | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP23169125455 | 2024-02-12 | 2024-02-15 | Cust 1527 | VIZAG | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | APPCB data upload issue. | We took the access to your system and observed that JAVA file is installed in your local PC. After uninstalling the same now local host server is accessible and data is live on both the server. | We took the access to your system and observed that JAVA file is installed in your local PC. After uninstalling the same now local host server is accessible and data is live on both the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVP23171165456 | 2024-02-16 | 2024-02-19 | Cust 1525 | VIZAG | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | APPCB data upload issue. | We took access to your system and now issue has been resolved and data is live on APPCB server | We took access to your system and now issue has been resolved and data is live on APPCB server | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVP23171175457 | 2024-02-16 | 2024-02-16 | Cust 497 | VIZAG | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | APPCB data upload issue. | We took the access to your system and now data is live on the APPCB server. | We took the access to your system and now data is live on the APPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVP23172775458 | 2024-02-20 | 2024-02-27 | Cust 1493 | VIZAG | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | OIL ANALYSER | OIL AWA ANALYSER(AWA Oil) | Display | Troubleshooting | Troubleshooting | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP23185385459 | 2024-03-13 | 2024-03-20 | Cust 1531 | VIZAG | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | MULTIPARAMETER ANALYSER | pH input card-FM of Multiparameter(1) | calibration and troubleshooting | Troubleshooting | Troubleshooting | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP23188665460 | 2024-03-19 | 2024-03-22 | Cust 1527 | VIZAG | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | APPCB data upload issue. | Now data is live on APPCB server. | Now data is live on APPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVP23192135461 | 2024-03-25 | 2024-04-01 | Cust 1535 | VIZAG | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | OIL ANALYSER | PCX1 Microprocessor with Dispay Board(PCX1) | Cpu fault | cpu problem sent for repair | sent to repair | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP24003335462 | 2024-04-08 | 2024-04-10 | Cust 613 | VIZAG | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | APPCB data issue | data is live on the APPCB server | data is live on the APPCB server | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVP24017555463 | 2024-05-06 | 2024-05-13 | Cust 1524 | VIZAG | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Chlorine sensor | Chlorine sensor-Krohne(1) | calibration | Calibration | Calibration | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP24028515464 | 2024-05-28 | 2024-05-31 | Cust 635 | VIZAG | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | Envirosense Software for CPCB(ESENSE1) | Camera Streaming issue on CPCB server. | Now the Camera streaming is live on the CPCB server. | Now the Camera streaming is live on the CPCB server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVP24028845465 | 2024-05-29 | 2024-06-03 | Cust 1535 | VIZAG | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Turbidity Meter | Optex(FMSC-P3) | Value constant | troubleshooting | troubleshooting | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP24047535466 | 2024-07-02 | 2024-07-08 | Cust 1359 | VIZAG | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Electrode | Ph Electrode(FM1104(4C)) | Commissioning | Commissiong | commissioning done | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP24066965467 | 2024-08-05 | 2024-08-11 | Cust 650 | VIZAG | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | Ph Tranmitter Smartpro | Ph Electrode Cable(Ph-Cab) | 2 nos ph tx issue | Online support | Troubleshooting | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP24066965467 | 2024-08-05 | 2024-08-11 | Cust 650 | VIZAG | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | Ph Tranmitter Smartpro | Ph Electrode Cable(Ph-Cab) | 2 nos ph tx issue | Online support | Troubleshooting | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP24075835468 | 2024-08-23 | 2024-08-28 | Cust 990 | VIZAG | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | Commissiong | Reagents not proper and ph sensor not working | Troubleshooting | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP24075835468 | 2024-08-23 | 2024-08-28 | Cust 990 | VIZAG | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | Commissiong | Pdo sensor response slow | Troubleshooting | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP24083305469 | 2024-09-03 | 2024-09-08 | Cust 1493 | VIZAG | Analytical - CHAKAN | CLOSED | COMMISSIONING | TOC ANALYSER | FM 101 TOC Analyser(FM 101 TOC) | Opc toc commissioning | Commissioning | New complaint will be logged | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP24083305469 | 2024-09-03 | 2024-09-08 | Cust 1493 | VIZAG | Analytical - CHAKAN | CLOSED | COMMISSIONING | TOC ANALYSER | FM 101 TOC Analyser(FM 101 TOC) | Opc toc commissioning | Commissioning | New complaint will be logged | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP24083305469 | 2024-09-03 | 2024-09-08 | Cust 1493 | VIZAG | Analytical - CHAKAN | CLOSED | COMMISSIONING | TOC ANALYSER | FM 101 TOC Analyser(FM 101 TOC) | Opc toc commissioning | New complaint will be raised | New complaint will be logged | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP24084375470 | 2024-09-06 | 2024-09-09 | Cust 1525 | VIZAG | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on CPCB | Internal problem is not resolved from customer side ,so we are closing this complaint for now. and we will reopen this complaint once we get access. | Internal problem is not resolved from customer side ,so we are closing this complaint for now. and we will reopen this complaint once we get access. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVP24086865471 | 2024-09-11 | 2024-09-11 | Cust 1525 | VIZAG | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | Data upload issue on CPCB server | We took access to the system and now the data is live on CPCB + RTDMS servers. | We took access to the system and now the data is live on CPCB + RTDMS servers. | YES | No Trip | Core | Y |  | Analytical - CHAKAN
AFVP24087185472 | 2024-09-12 | 2024-09-17 | Cust 1493 | VIZAG | Analytical - CHAKAN | CLOSED | COMMISSIONING | TOC ANALYSER | Protoc 300 24V Pump Motor/gearbox(17018) | Commissioning | Commissioning done | Commissioning completed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP24087195473 | 2024-09-12 | 2024-09-17 | Cust 650 | VIZAG | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Tranmitter Smartpro | Ph Electrode Cable(Ph-Cab) | Pcb issue | Troubleshooting | Troubleshooting | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP24091245474 | 2024-09-20 | 2024-09-25 | Cust 1536 | VIZAG | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Ph Electrode | Easycontrol ORP 120(HT238 523) | ph and conductivity & orp | Conductivity & orp sensor issue quotation guven | Troubleshooting | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP24091915475 | 2024-09-23 | 2024-09-28 | Cust 1019 | VIZAG | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | AQUADO Analyzer | OXYFERM(OXYFERM) | Values constant | Troubleshooting | Troubleshooting | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP24103995476 | 2024-10-16 | 2024-10-21 | Cust 1535 | VIZAG | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | OIL ANALYSER | Peristaltic cleaning pump motor(PMP-Clean-MT) | Not working | tubbing's and filters need to be replace | troubleshooting | YES | Non Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP24105235477 | 2024-10-18 | 2024-10-24 | Cust 650 | VIZAG | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | VISI pro DO(VISI pro DO) | Commission done by online support | Connection & value | Commissioning completed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP24109095478 | 2024-10-26 | 2024-11-02 | Cust 1537 | VIZAG | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Electrode | FM1104(1) | Commissioning | Commissioning | Commissioning completed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP24109095478 | 2024-10-26 | 2024-11-02 | Cust 1537 | VIZAG | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Electrode | FM1104(1) | Commissioning | Commissioning not done because of no flow through chamber | Commissioning completed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP24110305479 | 2024-10-28 | 2024-11-02 | Cust 1493 | VIZAG | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | COMMISSING | Commissioning done | Commissioning completed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP24110305479 | 2024-10-28 | 2024-11-02 | Cust 1493 | VIZAG | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Electrode Cable(Cond-Elec-cbl) | COMMSSING | Commissioning done | Commissioning completed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP24110305479 | 2024-10-28 | 2024-11-02 | Cust 1493 | VIZAG | Analytical - CHAKAN | CLOSED | COMMISSIONING | TOC ANALYSER | Protoc 300 24V Pump Motor/gearbox(17018) | commissioning | Not due to vu lamp issue | Commissioning completed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP24113275480 | 2024-11-05 | 2024-11-11 | Cust 1538 | VIZAG | Analytical - CHAKAN | CLOSED | COMMISSIONING | AQUADO Analyzer | OXYFERM(OXYFERM) | Issue in tx | Commissioning purpose | Commissioning completed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP24114445481 | 2024-11-08 | 2024-11-13 | Cust 1537 | VIZAG | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | Conductivity Tranmitter Smartpro | Conductivity Electrode Cable(Cond-Elec-cbl) | Flow through chamber need supply | rf raised | Rf need to raised | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFVP24114445481 | 2024-11-08 | 2024-11-13 | Cust 1537 | VIZAG | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | Conductivity Tranmitter Smartpro | Conductivity Electrode Cable(Cond-Elec-cbl) | Flow through chamber need supply | rf raised | Rf need to raised | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFVP24114445481 | 2024-11-08 | 2024-11-13 | Cust 1537 | VIZAG | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Need to supply 2 nos pH flow through chamber | rf raised | Rf need to raised | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFVP24114445481 | 2024-11-08 | 2024-11-13 | Cust 1537 | VIZAG | Analytical - CHAKAN | CLOSED | QUALITY PROBLEM | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Need to supply 2 nos pH flow through chamber | rf raised | Rf need to raised | YES | No Trip | Core | N |  | Analytical - CHAKAN
AFVP24120095482 | 2024-11-19 | 2024-11-24 | Cust 1432 | VIZAG | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Transmitter Aquacon | Conductivity Transmitter Aquamon 2000(CON-TR-2000) | calibration | calibration | Troubleshooting | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP24155195483 | 2025-01-23 | 2025-01-27 | Cust 1539 | VIZAG | Analytical - CHAKAN | CLOSED | COMMISSIONING | pH Transmitter Aquamon | Aquamon 2000 plus(AQ2000 plus) | Ph tx commissioning | online support | Online support given | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP24155195483 | 2025-01-23 | 2025-01-27 | Cust 1539 | VIZAG | Analytical - CHAKAN | CLOSED | COMMISSIONING | pH Transmitter Aquamon | Aquamon 2000 plus(AQ2000 plus) | Ph tx commissioning | commissioning done through online support | Online support given | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP24155215484 | 2025-01-23 | 2025-01-27 | Cust 1540 | VIZAG | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | ORP transmitter(1) | Commissioning | online support | Online support given | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP24159175485 | 2025-01-29 | 2025-02-28 | Cust 1000 | VIZAG | Steam System Division | CLOSED | QUALITY PROBLEM | PRV - MEASURETROL | FM MEASURETROL(SIZE 050NB) | New measurtrol reached to site. after commissioning one week prs got issue. visited the site. and checked and found that 24 v power supply diode is damaged. after discussion with division jithesh sir found that internal electronic cards and power cable is damaged. Kindly send the electronics and power cable as early as possible. | Mr Ramesh Tarra visited - Further to my visit to the site, electronics and power cable were replaced with a new in measuretrol valve. All required configuration and calibrations carried out and finally the valve was taken into operation and found working satisfactorily as per the setpoints. | Mr Ramesh Tarra visited -Electronics and power cable were replaced with a new in measuretrol valve. The valve was taken into operation and found working satisfactorily as per the setpoints. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFVP24165145486 | 2025-02-09 | 2025-02-14 | Cust 1541 | VIZAG | Analytical - CHAKAN | CLOSED | CALIBRATION JOB | Conductivity Transmitter Aquacon | Conductivity Transmitter Aquamon 2000(COND-Trans Aqu-2000) | Sensor wires connection | Calibration and troubleshooting | Troubleshooting | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP24182375487 | 2025-03-16 | 2025-03-22 | Cust 1542 | VIZAG | FM Valves - CHAKAN | CLOSED | COMMISSIONING | CHECK VALVE | BODY CHECK VALVE 25NB (1")(SCV25BD) | Dcv troubleshooting at s4 | Dcv troubleshooting | Troubleshooting | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP24182595488 | 2025-03-17 | 2025-03-24 | Cust 1543 | VIZAG | Analytical - CHAKAN | CLOSED | COMMISSIONING | Conductivity Tranmitter Smartpro | Conductivity Tranmitter Smartpro(Cond-Trans-Smartpro) | new ccds reached to site. needs to visit the site for commissioinng. | commissioning done. mom signed off. | commissioning done. mom signed off. | YES | No Trip | Process | N |  | Analytical - CHAKAN
AFVP24190255489 | 2025-03-31 | 2025-04-06 | Cust 1544 | VIZAG | Analytical - CHAKAN | CLOSED | PRODUCT FAILURE | Ph Electrode | FM1104(1) | The new pH Electrodes are not working and needs support. | pH sensor issue | Troubleshooting | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP25004535490 | 2025-04-11 | 2025-04-11 | Cust 1525 | VIZAG | Analytical - CHAKAN | CLOSED | DATA LOGGING/ CPCB SOFTWARE | SCADA Software | AIPL Software for CPCB(AIPL1) | APPCB DATA UPLODED ISSUE. | We took access and updated the Delay data output.jar file, so now the Data is live on the server. | We took access and updated the Delay data output.jar file, so now the Data is live on the server. | YES | No Trip | Core | Y | Failure of CPCB / Regulatory products / systems. | Analytical - CHAKAN
AFVP25009625491 | 2025-04-25 | 2025-05-02 | Cust 1019 | VIZAG | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | AQUADO Analyzer | OXYFERM(OXYFERM) | Tx issue | Sensor issue | Troubleshooting | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP25009625491 | 2025-04-25 | 2025-05-02 | Cust 1019 | VIZAG | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | AQUADO Analyzer | OXYFERM(OXYFERM) | Tx issue | Commission | Troubleshooting | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP25009625491 | 2025-04-25 | 2025-05-02 | Cust 1019 | VIZAG | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | AQUADO Analyzer | OXYFERM(OXYFERM) | Tx issue | Commissioninh | Troubleshooting | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP25009625491 | 2025-04-25 | 2025-05-02 | Cust 1019 | VIZAG | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | AQUADO Analyzer | OXYFERM(OXYFERM) | Tx issue | Commission | Troubleshooting | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP25024975492 | 2025-05-26 | 2025-05-31 | Cust 1545 | VIZAG | Analytical - CHAKAN | CLOSED | SYSTEM PERFORMANCE | OIL ANALYSER | UV Lamp box(1) | 4-20mA issue | Out of 4 oiw 3 of them having issue with 4-20mA card | Troubleshooting | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP25027075493 | 2025-05-29 | 2025-06-03 | Cust 993 | VIZAG | Analytical - CHAKAN | CLOSED | COMMISSIONING | DO Transmitters | DO Tranmitter Smartpro(8968 Smartpro DO) | Issue with housing connection | HOUSING NEED CHANGE RF Raised | AMC | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN
AFVP25034475494 | 2025-06-12 | 2025-06-17 | Cust 1000 | VIZAG | Analytical - CHAKAN | CLOSED | COMMISSIONING | Ph Tranmitter Smartpro | Ph Tranmitter Smartpro(8966 Smartpro PH) | Commissiong | Commissioning done | Commissioning completed | YES | Warranty (Free) | Core | N |  | Analytical - CHAKAN